## Summary20230816
GHWP 33 Members | QMS
East Asia (16) | Manufacturer | Distributor
Brunei Darussalam | No QMS requirements for medical devices and in-vitro diagnostic products (incl. software)
Cambodia | 1. No QMS requirements. 2. For registration need original Certificate for Foreign Government (CFG) or Free Sales Certificate (FSC) from the Country of Origin, Good Manufacturing Practice (GMP) or ISO13485 (ISO9001) certificate from the Country of Origin
China | GMP | GSP
Chinese Taipei | GMP | GDP
Hong Kong SAR, China | - | GDP
Indonesia | ISO13485 | CDAKB
Japan | JQMS | QMS Ordinance
Korea | GMP | GDP
Laos PDR | 1. No QMS requirements.
Malaysia | ISO13485 | GDPMD
Mongolia | 1. No QMS requirements. 2. ISO of Manufacturer, CE and quality certificates of medical device shall be required before import.
Myanmar | 1. No QMS requirements. 2. For all devices in scope a registration with MoH is required prior to import. Legalized Certificate for Foreign Government (CFG) or Free Sales Certificate (FSC) from the Country of Origin Good Manufacturing Practice (GMP) or ISO13485 (ISO9001) certificate from the Country of Origin
Philippines | GMP | GDP
Thailand | GMP | GDP
Vietnam | GMP | -
Singapore | ISO13485 | GDPMDS
GHWP 33 Members | QMS
West Asia (17) | Manufacturer | Distributor
Chile | ISO13485 | -
India | ISO13485 | -
Jordan | GMP | GDP
Kazakhstan | ISO13485 | -
Kingdom of Bahrain | Not found
Kingdom of Saudi Arabia | ISO13485 | GDP
Kyrgyz Republic | ISO13485 | -
Pakistan | ISO13485 | -
Kenya | GMP | GDP
South Africa | GMP | -
State of Kuwait | ISO13485
Sultanate of Oman | ISO13485 or ISO9001 (for low class medical devices)
Tanzania | ISO13485
United Arab Emirates | ISO13485
USA | QSR | -
Yemen | ISO13485
Zimbabwe | No regulations

## ISO13485 & JQMS
The New QMS Oridance (MHLW Ministerial Ordinance No. 169 in 2004: Ordinance on Standards for Manufacturing Control and Quality Control of Medical Devices and In Vitro Diagnostic Reagents. Consolidated Version of 26 March 2021) is the basis for the quality management system for MAH and manufacturers (only specified in Chapter 6 of the New QMS ordinance). This ordinance is the basis for GMP/QMS assessments for product certification in Japan. It should be noted that QMS certification is required for each medical device or each product family. New QMS ordinance was revised to align with ISO13485:2016 in March 26, 2021. The second chapter of the ordinance was revised to align with ISO13485:2016. The transition period is 3 years. Hence, MAH etc. and Registered Manufacturing Sites must comply with the revised ordinance by March 25, 2024.
Article 4.1 specifies that Class 1 medical devices are exempted from the requirements of design and development, Article 30 to Article 36-2. Article 4.2 and 4.3 specifies the rule of exclusion and non-application of the requirements. These articles are identical to the description of ISO 13485:2016 clause 1, paragraph 4 and 5.
ISO 13485:2016 | MHLW MO No.169, Chapter 2 | Note for understanding the requirements of MHLW MO No.169 Chapter 2, as amended in 2021
Clause 1 Scope | Section 1 General Rules
Clause 1, paragraph 4-5 | Article 4 | Article 4.1 specifies that Class 1 medical devices are exempted from the requirements of design and development, Article 30 to Article 36-2. Article 4.2 and 4.3 specifies the rule of exclusion and non-application of the requirements. These articles are identical to the description of ISO 13485:2016 clause 1, paragraph 4 and 5.
Clause 4 Quality management system | Section 2 Quality management system
Clause 4.1.1 | Article 5-1 | Roles undertaken by the organization are Marketing Authorization Holder provided by Article 23-2.1 of PMD Act, Registered Manufacturing Site provided by Article 23-2-3.1 and 23-2-4.1 of PMD Act, Seller of pharmaceutical products provided by Article 24.1 of PMD Act, Seller and Leaser of specially-controlled medical devices provided by Article 39.1 of PMD Act, Repairer of medical devices provided by Article 40-2.1 of PMD Act, or Seller and Leaser of controlled medical devices provided by Article 39-3.1 of PMD Act.
Clause 4.1.2 | Article 5-2
Clause 4.1.3 | Article 5-3
Clause 4.1.4 | Article 5-4
Clause 4.1.5 | Article 5-5
Clause 4.1.6 | Article 5-6
Clause 4.2.1 | Article 6
Clause 4.2.2 | Article 7-1
Clause 4.2.3 | Article 7-2
Clause 4.2.4 | Article 8 | The retention period of obsolete documents required by the ordinance is specified by Article 67 of MHLW MO No.169.
Clause 4.2.5 | Article 9 | The record retention period required by the ordinance is specified by Article 68 of MHLW MO No.169.
Clause 5 Management responsibility | Section 3 Management responsibility
Clause 5.1 | Article 10
Clause 5.2 | Article 11
Clause 5.3 | Article 12
Clause 5.4.1 | Article 13
Clause 5.4.2 | Article 14
Clause 5.5.1 | Article 15
Clause 5.5.2 | Article 16
Clause 5.5.3 | Article 17
Clause 5.6.1 | Article 18
Clause 5.6.2 | Article 19 | The organization is not required to input “reporting to regulatory authorities”, the item specified in ISO 13485:2016 5.6.2 c), to management review, when the organization is the person operating the registered manufacturing site.
Clause 5.6.3 | Article 20
Clause 6 Resource Management | Section 4 Resource Management
Clause 6.1 | Article 21
Clause 6.2, paragraph 1 and 2 | Article 22
Clause 6.2, paragraph 3 | Article 23
Clause 6.3 | Article 24
Clause 6.4.1 | Article 25-1
Clause 6.4.2 | Article 25-2
Clause 7 Product realization | Section 5 Product realization
Clause 7.1 | Article 26
Clause 7.2.1 | Article 27
Clause 7.2.2 | Article 28
Clause 7.2.3 | Article 29
Clause 7.3.1 and 7.3.2 | Article 30
Clause 7.3.3 | Article 31
Clause 7.3.4 | Article 32
Clause 7.3.5 | Article 33
Clause 7.3.6 | Article 34
Clause 7.3.7 | Article 35-1 | Clinical evaluations and/or evaluation of performance of the medical devices are required to be implemented as part of design and development validation, in the case that the medical device is designated by 23-2-5.3 or 23-2-9.4 of PMD Act.
Clause 7.3.8 | Article 35-2
Clause 7.3.9 | Article 36-1
Clause 7.3.10 | Article 36-2
Clause 7.4.1 | Article 37
Clause 7.4.2 | Article 38
Clause 7.4.3 | Article 39
Clause 7.5.1 | Article 40
Clause 7.5.2 | Article 41
Clause 7.5.3 | Article 42
Clause 7.5.4 | Article 43
Clause 7.5.5 | Article 44
Clause 7.5.6 | Article 45
Clause 7.5.7 | Article 46
Clause 7.5.8 | Article 47
Clause 7.5.9.1 | Article 48
Clause 7.5.9.2 | Article 49 | The requirements of Article 49.2 and Article 49.3, which are identical to the requirements of ISO 13485:2016 7.5.9.2 paragraph 2 and 3, are not applied, when the organization is the person operating the registered manufacturing site.
Clause 7.5.10 | Article 51
Clause 7.5.11 | Article 52
Clause 7.6 | Article 53
Clause 8 Measurement, analysis and improvement | Section 6 Measurement, analysis and improvement
Clause 8.1 | Article 54
Clause 8.2.1 | Article 55-1
Clause 8.2.2 | Article 55-2 | This article is identical to the requirement of ISO 13485:2016 8.2.2. However, it should be noted that the organization is required to determine the need to notify the information to the Marketing Authorization Holder instead of the regulatory authorities, when the organization is the person operating the registered manufacturing site.
Clause 8.2.3 | Article 55-3 | This article is identical to the requirement of ISO 13485:2016 8.2.3. However, it should be noted that the organization is required to notify the information to the Marketing Authorization Holder instead of the regulatory authorities, when the organization is the person operating the registered manufacturing site. Record of the notification shall also be maintained.
Clause 8.2.4 | Article 56
Clause 8.2.5 | Article 57
Clause 8.2.6, paragraph 1-3 | Article 58
Clause 8.2.6, paragraph 1-4 | Article 59
Clause 8.3.1 | Article 60-1
Clause 8.3.2 | Article 60-2
Clause 8.3.3 | Article 60-3
Clause 8.3.4 | Article 60-4
Clause 8.4 | Article 61
Clause 8.5.1 | Article 62
Clause 8.5.2 | Article 63
Clause 8.5.3 | Article 64
| MHLW MO No.169, Chapter 3-6 | Remarks
| Article 65 – 72-3 | Additional requirements for manufacturing control and quality control of medical devices, etc. Storage for the documents and records
| Article 73 – 79 | Manufacturing control and quality control of biological origin products, Add the contents of the Regulations for Buildings and Equipment
| Article 80 – 81 | Manufacturing control and quality control of radioactive in-vitro diagnostics The Regulations for Buildings and Equipment
| Article 81-2～81-2-6 | Manufacturing control and quality control of Single-use device (SUD )
| Article 82 - 84 | QMS for import QMS on outsourcing contractor
Sale to Consumers. Medical Devices
医療機器の販売業及び貸与業の取扱いについて The consumer including the hospital and medical institution can purchase medical devices from the retailer who has a business license for sales medical devices in Japan. In order to perform retail and rental business with medical devices, a notification for sales and rental business for Controlled Medical Devices is required for each sales office, and a retail and a rental business license are required for Highly Controlled Medical Devices. In the case of Internet commerce or Mail-order, the distributor/sales office must comply with the same regulation (PMD Act) as ordinary sales business requirements 医療機器の分割販売について Divided sales are allowed only for medical devices that are unlikely to be deteriorated or be harmful to health/hygiene after unpacking. Other cautions for divided sales of medical devices are explained in this document.
Post-market requirement:
MAHs should comply with Good Vigilance Practice (GVP) ordinance (MHLW Ordinance No. 135 dated November 24, 2017). GVP sets requirements regarding the organization, procedures for handling safety information.

## ISO13485 & QSR820
ISO13485 | QSR820 | Comments
On February 23, 2022, the FDA issued a proposed rule to amend the device current good manufacturing practice (CGMP) requirements of the Quality System (QS) Regulation to align more closely with the international consensus standard for Quality Management Systems for medical devices used by many other regulatory authorities around the world. The FDA proposes to do so primarily by incorporating by reference the 2016 edition of the international standard specific for medical device quality management systems set by the International Organization for Standardization (ISO), ISO 13485. Through this rulemaking, the FDA is also proposing additional requirements that help connect and align ISO 13485 with existing requirements in the Federal Food, Drug, and Cosmetic Act and its implementing regulations, and make conforming edits to 21 CFR Part 4 to clarify the device CGMP requirements for combination products. There are a lot of similarities between the requirements of the QS regulation and the clauses of ISO 13485:2016. The substance of the ISO 13485 requirements and the activities and actions required for compliance are substantively the same as under the current Part 820. The most noticeable difference between the QS regulation and the standard is the risk management requirements integrated throughout the aspects of the quality management system in ISO 13485. This differs from 21 CFR 820, in that the only risk specific requirement in the QS regulation is listed in § 820.30(g), as it relates to risk analysis as a part of design validation.
Clause 1. Scope, Clause 4. Quality Management System | Subpart A General Provisions | Requirements substantively similar.
Clause 4. Quality Management System, Clause 5. Management Responsibility, Clause 6. Resource Management, Clause 8. Measurement, Analysis, and Improvement | Subpart B Quality System Requirements | Requirements substantively similar.
Clause 7. Product Realization | Subpart C Design Controls | Requirements substantively similar.
Clause 4. Quality Management System | Subpart D Document Controls | Differences addressed in 820.35
Clause 7. Product Realization | Subpart E Purchasing Controls | Requirements substantively similar.
Clause 7. Product Realization | Subpart F Identification and Traceability | Requirements substantively similar.
Clause 4. Quality Management System, Clause 6. Resource Management, Clause 7. Product Realization | Subpart G Production and Process Controls | Requirements substantively similar.
Clause 7. Product Realization, Clause 8. Measurement, Analysis, and Improvement | Subpart H Acceptance Activities | Requirements substantively similar.
Clause 8. Measurement, Analysis, and Improvement | Subpart I Nonconforming Product | Requirements substantively similar.
Clause 8. Measurement, Analysis, and Improvement | Subpart J Corrective and Preventive Action | Requirements substantively similar.
Clause 7. Product Realization | Subpart K Labeling and Packaging Control | Differences addressed in 820.45
Clause 7. Product Realization | Subpart L Handling, Storage, Distribution, and Installation | Requirements substantively similar.
Clause 4. Quality Management System | Subpart M Records | Differences addressed in 820.35
Clause 7. Product Realization | Subpart N Servicing | Differences addressed in 820.35
Clause 7. Product Realization, Clause 8. Measurement, Analysis, and Improvement | Subpart O Statistical Techniques | Requirements substantively similar.

## Philippines & ISO13485
| ISO 13485:2016 | Good Storage and Distribution Practices for Medical Devices | Difference
Definition Medical device (ISO 3.11 ) | instrument, apparatus, implement, machine, appliance, implant, reagent for in vitro use, software, material or other similar or related article, intended by the manufacturer to be used, alone or in combination, for human beings, for one or more of the specific medical purpose(s) of: — diagnosis, prevention, monitoring, treatment or alleviation of disease; — diagnosis, monitoring, treatment, alleviation of or compensation for an injury; — investigation, replacement, modification, or support of the anatomy or of a physiological process; — supporting or sustaining life; — control of conception; — disinfection of medical devices; — providing information by means of in vitro examination of specimens derived from the human body; | Medical Device – any instrument, apparatus, implement, machine, appliance, implant, in-vitro reagent or calibrator, software, material, or other similar or related article intended by the manufacturer to be used alone, or in combination, for human beings for one or more of the specific purpose(s) of diagnosis, prevention, monitoring, treatment or alleviation of disease; diagnosis, monitoring, treatment, alleviation of, or compensation for an injury; investigation, replacement, modification, or support of the anatomy or of a physiological process; supporting or sustaining life; preventing infection; control of conception; disinfection of medical devices; and providing information for medical or diagnostic purposes by means of in-vitro examination of specimens derived from the human body. This device does not achieve its primary intended action in or on the human body by pharmacological, immunological or metabolic means but which may be assisted in intended function by such means.
Control of records (ISO 4.2.5 ) | Records shall be maintained to provide evidence of conformity to requirements and of the effective operation of the quality management system. The organization shall document procedures to define the controls needed for the identification, storage, security and integrity, retrieval, retention time and disposition of records. The organization shall define and implement methods for protecting confidential health information contained in records in accordance with the applicable regulatory requirements. Records shall remain legible, readily identifiable and retrievable. Changes to a record shall remain identifiable. The organization shall retain the records for at least the lifetime of the medical device as defined by the organization, or as specified by applicable regulatory requirements, but not less than two years from the medical device release by the organization. | 1. Records shall be kept for each purchase and sale of medical device showing the following details specified below: 2. Records shall be kept for as long as the medical device product is registered with the FDA. 3. Records shall show traceability of the origin. 4. Proof of delivery transactions shall be kept. 5. There shall be an orderly and secure system of filing up to date invoices from suppliers and buyer. 6. Records shall be easily retrievable, stored and protected against unauthorizedbmodification, damage, and/or loss of documentation. 7. The establishment shall keep records of laws, policies and latest issuances relativebto the regulation of medical devices. 8. If the records are generated and kept in electronic form, there shall be backups to prevent any accidental data loss.
|  | a. For manufacturer i. Date of purchase ii. Name of medical device iii. Model number (if applicable) iv. Batch/Lot number (if applicable) v. Expiration date (if applicable) vi. Specification of product (if applicable) vii. Quantity distributed/sold viii. Distribution records ix. CPR/CMDR/CMDN of the product x. Records of full documentation of the registration of the product submitted during registration/notification | Specific requirement on record
|  | b. Distributor (exporter/importer/wholesaler) i. Date of purchase ii. Name of medical device iii. Model number (if applicable) iv. Batch/Lot number (if applicable) v. Expiration date (if applicable) vi. Specification of product (if applicable) vii. Quantity received and sold viii. Name of licensed supplier ix. Copies of valid LTO of licensed supplier x. Notarized MOA/Contract with supplier /distribution agreement (if applicable) xi. Importation and distribution records xii. Copies of CPR/CMDR/CMDN of the product xiii. ISO certificate for the foreign supplier (manufacturer) or GMP certification for local manufacturers xiv. Records of full documentation of the registration of the product submitted during registration/notification | Specific requirement on record
STANDARD OPERATING PROCEDURES (ISO 4.2.4) |  | The establishments shall have standard operating procedures for the following activities:
7.4.1 | The organization shall document procedures (see 4.2.4) to ensure that purchased product conforms to specified purchasing information. | 1. Procurement of products
7.5.1 | As appropriate, production controls shall include but are not limited to: f) implementation of product release, delivery and post-delivery activities. | 2. Release of products | SOP on product release
6.3 | The organization shall document the requirements for the infrastructure needed to achieve conformity to product requirements, prevent product mix-up and ensure orderly handling of product. Infrastructure includes, as appropriate: c) supporting services (such as transport, communication, or information systems). | 3. Transportation / shipment of products including investigation in cases of failure to comply with storage requirement during transit (if applicable) | Storage requirement during transit
7.6 | The organization shall document procedures to ensure that monitoring and measurement can be carried out and are carried out in a manner that is consistent with the monitoring and measurement requirements. | 4. Maintenance of vehicle and equipment
7.5.11 | mThe organization shall document procedures for preserving the conformity of product to requirements during processing, storage, handling, and distribution. | 5. Storage of stocks including temperature monitoring and retention of monitoring records (if applicable) | Temperature
8.2.2 | The organization shall document procedures for timely complaint handling in accordance with applicable regulatory requirements. | 6. Handling of complaints - a distinction should be made between complaints about a product or its packaging and those relating to distribution
7.5.8 | The organization shall document procedures to ensure that medical devices returned to the organization are identified and distinguished from conforming product. | 7. Handling of return of medical devices.
|  | 8. Handling of product recall including storage and final disposition of recalled products. | Recall requirement
|  | 9. Destruction of medical device (expired, damaged, unusable stocks) | Destruction
|  | 10. Handling of medical device products at the port/s of entry (if importer) | Importer
11. Self-inspection to monitor implementation and compliance with the principles of this guideline and, if necessary, corrective and preventive measures.
PERSONNEL Responsibility and Authority (5.5.1) | Top management shall ensure that responsibilities and authorities are defined, documented and communicated within the organization. Top management shall document the interrelation of all personnel who manage, perform and verify work affecting quality and shall ensure the independence and authority necessary to perform these tasks. | Establishment shall have an organizational chart and documented job descriptions, including key responsibilities and authorities.
Training and Competence 6.2.2 | Personnel performing work affecting product quality shall be competent on the basis of appropriate education, training, skills and experience. | Key personnel in charge including those in the warehousing operations shall possess the necessary competence in terms of education, training, skills and experience. Training may include proper handling, use and operation of the device, as applicable, warehouse management and inventory.
Cleanliness 6.4.2 | As appropriate, the organization shall plan and document arrangements for the control of contaminated or potentially contaminated product in order to prevent contamination of the work environment, personnel, or product. | a. Premises shall be kept clean. b. A written sanitation program shall be available indicating the frequency of cleaning and the methods to be used to clean the premises and storage areas. c. There shall be no smoking, eating and drinking in areas where medical devices are stored and handled. | Detailed requirements for PH
Vermin Control 6.4.2 |  | a. The storage area shall be designed and equipped to prevent entry of vermin and other animals into the area b. There shall be a written vermin control program to prevent vermin infestation. Appropriate records shall be kept for this purpose. | Detailed requirements on Vermin
Receipt of stocks 7.4.3 | The organization shall establish and implement the inspection or other activities necessary for ensuring that purchased product meets specified purchasing requirements. | a. The receiving area shall be adequate and shall have protection for deliveries from bad weather during unloading. b. It shall be separate from the storage area. c. Each incoming delivery shall be checked to ensure that the correct product is delivered from the correct supplier. This may include checking of the purchase order, containers, label description, batch number, expiry date, product and quantity. d. The delivery shall be examined for uniformity of the containers and, if necessary, shall be subdivided according to the supplier’s batch number should the delivery comprise more than one batch. e. Each container shall be carefully checked for possible contamination, tampering and damage. Any suspect containers or, if necessary, the entire delivery shall be quarantined for further investigation. f. Medical devices subject to specific storage measures shall be immediately identified and stored in accordance with the specified instruction(s). g. Measures shall be taken to ensure that rejected medical devices cannot be used. They shall be segregated and securely stored while awaiting destruction or return to the supplier. | Receiving requirements
Storage 7.5.11 | The organization shall document procedures for preserving the conformity of product to requirements during processing, storage, handling, and distribution. Preservation shall apply to the constituent parts of a medical device. The organization shall protect product from alteration, contamination or damage when exposed to expected conditions and hazards during processing, storage, handling, and distribution by: a) designing and constructing suitable packaging and shipping containers; b) documenting requirements for special conditions needed if packaging alone cannot provide preservation. | a. Precautions shall be taken to prevent unauthorized persons from entering storage areas. b. Storage areas shall be of sufficient capacity to allow the orderly storage of the different categories of medical devices. c. There shall be adequate lighting and ventilation. d. Storage areas shall be maintained within acceptable and specified temperature limits. The storage conditions for medical devices shall be in accordance with their labelling. e. Where special storage conditions are required on the label (e.g. temperature, relative humidity), these shall be provided, controlled, monitored and recorded. Temperature of the storage areas shall be measured at suitable predetermined intervals to show the maximum and minimum temperatures for the day and recorded. The equipment used for monitoring shall be calibrated and be suitable for their intended use. Calibration shall be traceable to applicable international/national standard and the calibration records shall be maintained. f. Medical devices shall be stored off the floor and suitably spaced to allow cleaning and inspection. If pallets are used, such shall be well maintained and kept clean at all times. g. Medical devices shall be handled and stored in such a manner as to prevent contamination, mix-ups and cross-contamination. h. Appropriate and suitable storage conditions shall be provided for hazardous, sensitive and dangerous materials such as combustible liquids and solids, pressurized gases, highly toxic and radioactive substances. i. There shall be designated areas for quarantined, saleable stock, expired, rejected/damaged, recalled and returned medical devices. Such medical devices shall be properly labeled. j. A list of medical devices shall be maintained and the medical devices properly identified.
Delivery to Customers |  | a. Medical products shall be transported in accordance with the conditions stated on the labels. There shall be no risk to the quality of the medical devices during transport and delivery. b. Delivery vehicles shall be suitable for their purpose, with sufficient space and appropriately equipped to protect the products. c. When the medical device is intended solely for professional use, such medical device shall be sold to the indicated type of user or to customers with medical prescription only. | Requirement on shipment
INSTALLATIONS AND SERVICING 7.5.3 | The organization shall document requirements for medical device installation and acceptance criteria for verification of installation, as appropriate. If the agreed customer requirements allow installation of the medical device to be performed by an external party other than the organization or its supplier, the organization shall provide documented requirements for medical device installation and verification of installation. | 1. Where installation and servicing of a medical equipment are required, these shall be conducted by competent or trained personnel. 2. Installation, inspection, any required testing and/or servicing shall be performed in accordance with the manufacturer’s instructions and procedures. 3. The establishment shall maintain records of inspection and any test results to demonstrate proper installation of the medical equipment. Records of servicing shall likewise be maintained.
MEDICAL DEVICE COMPLAINTS 8.2.2 | These procedures shall include at a minimum requirements and responsibilities for: a) receiving and recording information; b) evaluating information to determine if the feedback constitutes a complaint; c) investigating complaints; d) determining the need to report the information to the appropriate regulatory authorities; e) handling of complaint-related product; f) determining the need to initiate corrections or corrective actions. If any complaint is not investigated, justification shall be documented. Any correction or corrective action resulting from the complaint handling process shall be documented. If an investigation determines activities outside the organization contributed to the complaint, relevant information shall be exchanged between the organization and the external party involved. | 1. The establishment shall have a record for each individual complaint. The records on complaints may include the following information: a. the medical device brand name, CPR/CMDR/CMDN number, model/catalogue number or bar code, control/serial/ lot number and any other means of identification of the medical device; b. the name(s) and address(es) of the supplier; c. records pertaining to the problem investigation. 2. The record shall show that the complaints received are investigated and followed through, and that corrective actions are taken to prevent repeated complaints and, where a decision is made to recall the product, the details of the recall. 3. There shall be a designated personnel to handle complaints. 4. The written standard operating procedures for handling of complaints shall identify the designated personnel and describe his/her functions and responsibilities. The procedure shall contain the following: a. determination of whether there is a health hazard associated with the medical device; b. determination of whether the medical device fails to conform to any claim made by the dealer relating to its effectiveness, benefits, performance characteristics or safety; c. determination of whether the medical device fails to meet any regulatory requirements d. determination of the most appropriate preventive/corrective action; e. justification when no action is taken, for example, in the case of receiving an unfounded or invalid complaint. 5. Complaint records maintained with respect to a medical device should be retained for a period of five (5) years on top of the projected useful life of the medical device as determined by the product owner. For example, if the projected useful life of the medical device is one year, the complaint records should be kept for six years.
RETURN OF MEDICAL DEVICES |  | 1. There shall be standard operating procedures for handling of return of medical devices. 2. There shall be a record for returned medical device products. 3. All returned medical devices shall be kept separate from saleable stock until a decision has been reached regarding their disposal. 4. Returned medical devices shall be re-evaluated and can be returned to saleable stock provided the products have met the following criteria as applicable: a. the medical devices are in their original unopened containers and in good condition; b. it is known that the medical devices have been stored and handled under proper conditions; c. the remaining shelf life period is acceptable; and d. the medical devices have been examined and assessed by appropriate personnel. The assessment shall take into account the nature of the medical device, any special storage conditions required, and the time that has elapsed since it was distributed. Special attention shall be given to thermolabile medical devices. Advice shall be sought from the product owner as | Return of medical device
DISPOSAL OF MEDICAL DEVICES |  | 1. Medical devices that are expired, contaminated, damaged or defective or found to be counterfeit shall be disposed of properly in accordance with applicable government rules and regulations 2. The establishment shall ensure that the medical devices for disposal will not be sold. 3. The disposal shall be properly recorded and such records shall be maintained by the establishment. | Disposal of medical device
RECALLED, COUNTERFEIT, ADULTERATED, TAMPERED, BANNED, 7 PHASED-OUT, UNUSABLE MEDICAL DEVICES |  | 1. Any recalled, counterfeit, adulterated, tampered, banned, phased-out or unusable medical devices shall be physically segregated and stored separately from other medical devices. For recalled products, the establishment shall comply with FDA Circular No. 2016-012 entitled “Guidelines for Product Recall”. 2. These medical devices shall be labeled as “Not for Sale” for easy identification and guidance of the personnel. 3. There shall be a designated personnel to handle recalled, counterfeit, adulterated, tampered, banned, phased-out or unusable medical devices. 4. The establishment shall inform the Food and Drug Administration – Center for Device Regulation, Radiation Health, and Research, the supplier (manufacturer/trader/distributor) or the product owner immediately regarding the recalled, counterfeit, banned, adulterated or tampered medical device product. | Recall of medical device

## ISO13485&IndiaGMP
From Stryker 总的来说，印度的GMP要求基本就是ISO13485直接采纳来的，它法规的条款框架几乎跟标准一样。其中差别最大的部分，个人总结： 1. ISO标准会比印度法规更广泛一点，比如标准不仅针对生产厂商，其中组织的概念可能包括制造商、授权代表、进口商或经销商。但是印度法规明确就只是生产商。 2. 印度GMP里没有标准中对于投诉处置的这部分内容。 3. 标准中有不少关于文件记录的要求，印度法规没有一一对应，但是我认为是印度法规把这一要求集中在了一个条款中。 所以文件中你可以看到我mark了蓝色的部分，是ISO标准中有的内容在印度法规没有的条款。我mark红色的部分，是两者存在差异的部分。
ISO 13485: 2016 |  | India
ISO 13485: 2016 |  | Medical Device Rules, 2017 Fifth Schedule [See rule 20(3), 20(5), 20(8), 22(i)] Quality Management System for medical devices and in vitro diagnostic medical devices | 对应中文翻译
本标准为需要证实自身有能力提供持续满足顾客要求和适用的法规要求的医疗器械和相关服务的组织规定了质量管理体系要求。这类组织能涉及医疗器械生命周期的一个或多个阶段，包括医疗器械的设计和开发、生产、贮存和流通、安装或服务，以及相关活动（例如技术支持）的设计和开发或提供。本标准也能用于向这类组织提供产品（包括与质量管理体系相关的服务）的供方或外部方。 | This International Standard specifies requirements for a quality management system where an organization needs to demonstrate its ability to provide medical devices and related services that consistently meet customer and applicable regulatory requirements. Such organizations can be involved in one or more stages of the life-cycle including design and development, production, storage and distribution, installation, or servicing of a medical device and design and development or provision of associated activities (e.g. technical support). This International Standard can also be used by suppliers or external parties that provide product including quality management system-related services to such organizations. | 1. General Requirements: 1.1. This schedule specifies requirements for a quality management system that shall be used by the manufacturer for the design and development, manufacture, packaging, labelling, testing, installation and servicing of medical devices and in vitro diagnostic medical devices. If the manufacturer does not carry out design and development activity, the same shall be recorded in the quality management system. The manufacturer shall maintain conformity with this Schedule to reflect the exclusions. | 1. 一般要求: 1.1. 本附表规定了制造商在医疗器械和体外诊断医疗器械的设计和开发、制造、包装、标签、测试、安装和服务中应采用的质量管理体系要求。如果制造商不进行设计和开发活动，则应将其记录在质量管理体系中。制造商应保持与本附表的一致性，以反映排除情况。
除非明确规定，本标准的要求适用于各种规模和类型的组织。本标准中应用于医疗器械的要求同样适用于组织提供的相关服务。对于本标准所要求的适用于组织但不是由组织实施的过程，在质量管理体系中组织通过监视、维护和控制这些过程对其负有责任。 | Requirements of this International Standard are applicable to organizationsregardless of their size and regardless of their type except where explicitlystated. Wherever requirements are specified as applying to medical devices,the requirements apply equally to associated services as supplied by theorganization.The processes required by this International Standard that are applicableto the organization, but are not performed by the organization, are theresponsibility of the organization and are accounted for in the organization’squality management system by monitoring, maintaining, and controllingthe processes. | 1.3. The processes required by this Schedule, which are applicable to the medical device and in-vitro diagnostic medical device, but which are not performed by the manufacturer are the responsibility of the manufacturer and are accounted for in the manufacturer’s quality management system. | 1.3. 本附表所要求的适用于医疗器械和体外诊断医疗器械，但不是由制造商执行的工艺，由制造商负责，并在制造商的质量管理体系中进行解释。
如果适用的法规要求允许对设计和开发控制进行删减，则能作为在质量管理体系中将其删减的理由。若这些法规要求能提供其他方法，这些方法要在质量管理体系中予以说明。组织有责任确保在符合本标准的声明中明确对设计和开发控制的任何删减。 | If applicable regulatory requirements permit exclusions of design and development controls, this can be used as a justification for their exclusion from the quality management system. These regulatory requirements can provide alternative approaches that are to be addressed in the quality management system. It is the responsibility of the organization to ensure that claims of conformity with this International Standard reflect any exclusion of design and development controls. | 1.4 If a manufacturer engages in only some operations subject to the requirements of this part, and not in others, that manufacturer need only to comply with those requirements which are applicable to the operations in which it is engaged. | 1.4 如果制造商只从事受本部分要求约束的某些操作，而不从事其他操作，则该制造商只需遵守适用于其从事的操作的那些要求。
本标准第6、7或8章中的任何要求，如果因组织开展的活动或质量管理体系所涉及的医疗器械的特点而不适用时，组织不需要在其质量管理体系中包含这样的要求。对于经确定不适用的任何条款，组织按照4.2.2的要求记录其理由。 | If any requirement in Clauses 6, 7 or 8 of this International Standard is not applicable due to the activities undertaken by the organization or the nature of the medical device for which the quality management system is applied, the organization does not need to include such a requirement in its quality management system. For any clause that is determined to be not applicable, the organization records the justification as described in 4.2.2. | 1.2 If any requirement in paragraph 7 (product realisation) of this Schedule is not applicable due to the nature of the medical device and in vitro diagnostic medical devices for which the quality management system is applied, the manufacturer does not need to include such a requirement in its quality management system. | 1.2 如果由于应用质量管理体系的医疗器械和体外诊断医疗器械的性质，本附表第7段(产品实现)中的任何要求不适用，制造商不需要在其质量管理体系中包括该要求。
/ | / | 1.5. It is emphasized that the quality management system requirements specified in this Schedule are in addition to complementary to technical requirements for products. | 1.5. 需要强调的是，本附表规定的质量管理体系要求是对产品技术要求的补充。
/ | / | 1.6. Manufacturers of components or parts of finished devices and in vitro diagnostic medical devices are encouraged to use appropriate provisions of this schedule as guidance. | 1.6. 鼓励成品器械和体外诊断医疗器械的组件或部件制造商使用本附表的适当规定作为指导。
/ | / | 2. Applicability: The provisions of this Schedule shall be applicable to manufacturers of finished devices, in vitro diagnostic medical devices, mechanical contraceptives (condoms, intrauterine devices, tubal rings), surgical dressings, surgical bandages, surgical staplers, surgical sutures and ligatures, blood and blood components collection bags with or without anticoagulants. | 2. 适用性: 本附表的条文适用于成品器械、体外诊断医疗器械、物理避孕药具(避孕套、宫内节育器、输卵管环)、外科敷料、外科绷带、外科吻合器、外科缝合线及结扎线、含或不含抗凝血剂的血液及血液成分收集袋的制造商。
4 质量管理体系 | 4 Quality management system | 4 Quality management system. | 4 质量管理体系
4.1 总要求 | 4.1 General requirements | 4.1 General | 4.1 总要求
4.1.1 组织应按照本标准的要求和适用的法规要求将质量管理体系形成文件并保持其有效性。 | 4.1.1 The organization shall document a quality management system and maintain its effectiveness in accordance with the requirements of this International Standard and applicable regulatory requirements. | The manufacturer shall establish, document, implement and maintain a quality management system and maintain its effectiveness in accordance with the requirements of this schedule. | 制造商应按照本附表的要求建立、形成文件、实施和保持质量管理体系，并保持其有效性。
组织应按照本标准或者适用的法规要求建立、实施和保持需要形成的文件的所有要求、程序、活动或安排。 | The organization shall establish, implement and maintain any requirement, procedure, activity or arrangement required to be documented by this International Standard or applicable regulatory requirements. | / | /
组织应将其在适用的法规要求下所承担的一个或多个角色形成文件。 | The organization shall document the role(s) undertaken by the organization under the applicable regulatory requirements. | / | /
注：组织所承担的角色可能包括制造商、授权代表、进口商或经销商。 | NOTE Roles undertaken by the organization can include manufacturer, authorized representative, importer or distributor. | / | /
4.1.2 组织应： | 4.1.2 The organization shall: | The manufacturer shall; | 制造商应
a） 考虑组织承担的角色来确定质量管理体系所需的过程及这些过程在整个组织中的应用； | a) determine the processes needed for the quality management system and the application of these processes throughout the organization taking into account the roles undertaken by the organization; | (a) identify the processes needed for the quality management system and their application throughout the organisation; | (a)识别质量管理体系所需的过程及其在整个组织中的应用;
b） 应用基于风险的方法控制质量管理体系所需的适当过程 | b) apply a risk based approach to the control of the appropriate processes needed for the quality management system; | / | /
c） 确定这些过程的顺序和相互作用。 | c) determine the sequence and interaction of these processes. | (b) determine the sequence and interaction of these processes; | (b)确定这些过程的顺序和相互作用;
4.1.3 对于每个质量管理体系过程，组织应： | 4.1.3 For each quality management system process, the organization shall: | / | /
a）确定所需的准则和方法，以确保这些过程的有效运行和控制。 | a) determine criteria and methods needed to ensure that both the operation and control of these processes are effective; | (c) determine criteria and methods needed to ensure that both the operation and control of these processes are effective; | (c)确定所需的标准和方法，以确保这些过程的操作和控制都是有效的;
b）确保可获得必要的资源和信息，以支持这些过程的运行和对这些过程的监视 | b) ensure the availability of resources and information necessary to support the operation and monitoring of these processes; | (d) ensure the availability of resources and information necessary to support the operation and monitoring of these processes; | (d)确保提供必要的资源和资料，以支持这些进程的运作和监测;
c）实施必要的措施，以实现这些过程策划的结果并保持这些过程的有效性 | c) implement actions necessary to achieve planned results and maintain the effectiveness of these processes; | (f) implement actions necessary to achieve planned results and maintain the effectiveness of these processes | (f)实施必要的行动，以实现所策划的结果并保持这些过程的有效性
d）监视、测量（适当时）和分析这些过程 | d) monitor, measure as appropriate, and analyse these processes; | (e) monitor, measure and analyse these processes; and | (e)监测、测量和分析这些过程;和
e）建立和保持所需的记录以证实符合本标准并满足适用的法规要求（见4.2.5） | e) establish and maintain records needed to demonstrate conformance to this International Standard and compliance with applicable regulatory requirements (see 4.2.5). | / | /
4.1.4 组织应按照本标准要求和适用的法规要求管理这些质量管理体系过程。更改这些过程应： | 4.1.4 The organization shall manage these quality management system processes in accordance with the requirements of this International Standard and applicable regulatory requirements. Changes to be made to these processes shall be: | / | /
a）评价过程更改对质量管理体系的影响 | a) evaluated for their impact on the quality management system; | / | /
b）评价过程更改对该质量管理体系中所生产的医疗器械的影响 | b) evaluated for their impact on the medical devices produced under this quality management system; | / | /
c）按照本标准的要求和适用的法规要求进行控制 | c) controlled in accordance with the requirements of this International Standard and applicable regulatory requirements. | / | /
4.1.5 若组织选择将影响产品符合要求的任何过程外包，组织应监视这类过程并确保对其控制。组织应保留外包过程符合本标准要求、顾客要求和适用的法规要求的责任。控制应与所涉及的风险和外部方满足7.4中要求的能力相适应。控制应包括书面质量协议。 | 4.1.5 When the organization chooses to outsource any process that affects product conformity to requirements, it shall monitor and ensure control over such processes. The organization shall retain responsibility of conformity to this International Standard and to customer and applicable regulatory requirements for outsourced processes. The controls shall be proportionate to the risk involved and the ability of the external party to meet the requirements in accordance with 7.4. The controls shall include written quality agreements. | These processes shall be managed by the manufacturer in accordance with the requirements of this Schedule. Where a manufacturer chooses to outsource any process that affects product conformity with requirements, the manufacturer shall ensure control over such processes. Control of such outsourced processes shall be identified within the quality management system. NOTE: Processes needed for the quality management system referred to above shall include processes for management activities, provision of resources, product realization and measurement. | 这些过程应由制造商按照本附表的要求进行管理。如果制造商选择外包任何影响产品符合性的过程，制造商应确保对这些过程进行控制。应在质量管理体系内确定对此类外包过程的控制。 注:上述质量管理体系所需的过程应包括管理活动、资源提供、产品实现和测量的过程。
4.1.6 组织应将用于质量管理体系的计算机软件应用的确认程序形成文件。在软件首次使用前应对软件应用进行确认，适当时，软件或其应用更改后也应对软件应用进行确认。 | 4.1.6 The organization shall document procedures for the validation of the application of computer software used in the quality management system. Such software applications shall be validated prior to initial use and, as appropriate, after changes to such software or its application. | / | /
与软件确认和再确认有关的特定方法和活动应与软件使用有关的风险相适应。 | The specific approach and activities associated with software validation and revalidation shall be proportionate to the risk associated with the use of the software. | / | /
应保留这些活动的记录（见4.2.5）。 | Records of such activities shall be maintained (see 4.2.5). | / | /
4.2 文件要求 | 4.2 Documentation requirements | 4.2 Documentation requirements | 4.2 文件要求
4.2.1 总则 | 4.2.1 General | 4.2.1 General | 4.2.1 总则
质量管理体系文件（见4.2.4）应包括： | The quality management system documentation (see 4.2.4) shall include: | The quality management system documentation shall include;- | 质量管理体系文件应包括:-
a） 形成文件的质量方针和质量目标； | a) documented statements of a quality policy and quality objectives; | (a) documented statements of a quality policy and quality objectives; | (a)质量方针和质量目标的文件化声明;
b） 质量手册； | b) a quality manual; | (b) a quality manual; | (b)质量手册;
c） 本标准所要求的形成文件的程序和记录； | c) documented procedures and records required by this International Standard; | (c) documented procedures required by this schedule; | (c)本附表要求的文件化程序;
d） 组织确定的为确保其过程有效策划、运行和控制所需的文件，包括记录； | d) documents, including records, determined by the organization to be necessary to ensure the effective planning, operation, and control of its processes; | (d) documents needed by the manufacturer to ensure the effective planning, operation and control of its processes; | (d)制造商为确保其过程的有效策划、运行和控制所需的文件;
e） 适用的法规要求规定的其他文件。 | e) other documentation specified by applicable regulatory requirements. | (e) records required by this Schedule, and | (e)本附表所要求的记录;以及
/ | / | where this Schedule specifies that a requirement, procedure, activity or special arrangement be “documented”, it shall, in addition, be implemented and maintained. For each type of medical device or in vitro diagnostic medical devices, the manufacturer shall establish and maintain a file either containing or identifying documents defining product specifications and quality management system requirements. These documents shall define the complete manufacturing process and, if applicable, installation. The manufacture shall prepare documentation for device or in vitro diagnostic medical devices in a form of a Device Master File containing specific information as referred to in Fourth Schedule. Data may be recorded by electronic data processing systems or other reliable means, but documents and record relating to the system in use shall also be available in a hard copy to facilitate checking of the accuracy of the records. Wherever documentation is handled by electronic data processing methods, authorized persons shall enter or modify data in the computer. There shall be record of changes and deletions. Access shall be restricted by ‘passwords’ or other means and the result of entry of critical data shall be independently checked. Batch records electronically stored shall be protected by a suitable back-up. During the period of retention, all relevant data shall be readily available. | 当本附表规定要求、程序、活动或特殊安排“形成文件”时，还应予以实施和保持。对于每种类型的医疗器械和体外诊断医疗器械，生产企业应当建立并保持包含或识别定义产品规格和质量管理体系要求的文件。这些文件应定义完整的制造过程，如果适用，还应定义安装。 制造商须为器械或体外诊断医疗器械以器械主文件的形式准备文件，其中包含附表4所提述的特定信息。数据可以通过电子数据处理系统或其他可靠的方式记录，但与所使用的系统有关的文件和记录也应以硬拷贝的形式提供，以方便检查记录的准确性。凡采用电子数据处理方式处理文件的，被授权人应当在计算机中输入或者修改数据。变更和删除应当有记录。应通过“密码”或其他方式限制访问，并应独立检查关键数据的输入结果。电子存储的批记录应当有适当的备份保护。在保存期间，所有相关数据应随时可用。
4.2.2 质量手册 | 4.2.2 Quality manual | 4.2.2 Quality manual | 4.2.2 质量手册
组织应编制质量手册，质量手册包括： | The organization shall document a quality manual that includes: | The manufacturer shall establish and maintain a quality manual that includes:- | 制造商应建立并保持质量手册，其中包括
a） 质量管理体系的范围，包括任何删减不适用的详细说明和理由； | a) the scope of the quality management system, including details of and justification for any exclusion or non-application; | (a) the scope of the quality management system, including details of and justification for any exclusion or nonapplication or both; | (a)质量管理体系的范围，包括任何删减或不适用的细节和理由，或两者兼而有之;
b） 质量管理体系的形成文件的程序或对其引用； | b) the documented procedures for the quality management system, or reference to them; | (b) the documented procedures established for the quality management system, or reference to them; and | (b)为质量管理体系制定的形成文件的程序或对其的参考;和
c） 质量管理体系过程之间的相互作用的表述。 | c) a description of the interaction between the processes of the quality management system. | (c) a description of the interaction between the processes of the quality management system. | (c)质量管理体系过程之间相互作用的描述。
质量手册应概述质量管理体系的文件结构。 | The quality manual shall outline the structure of the documentation used in the quality management system. | The quality manual shall outline the structure of the documentation used in the quality management system. | 质量手册应概述质量管理体系所用文件的结构。
4.2.3 医疗器械文档 | 4.2.3 Medical device file
组织应为每个医疗器械类型或医疗器械族建立并保持一个或多个文档，文档包含或引用形成的文件以证明符合本标准要求和适用的法规要求。 | For each medical device type or medical device family, the organization shall establish and maintain one or more files either containing or referencing documents generated to demonstrate conformity to the requirement of this International Standard and compliance with applicable regulatory requirements. | The manufacturer shall prepare documentation in a form of a Plant Master File containing specific information about the facilities, personnel and other details as prescribed in Fourth. | 制造商应以工厂主文档的形式准备文件，其中包含第4条规定的有关设施、人员和其他细节的具体信息。
文档的内容应包括但不限于： | The content of the file(s) shall include, but is not limited to: | / | /
a） 医疗器械的概述、预期用途/预期目的和标记，包括所有使用说明； | a) general description of the medical device, intended use/purpose, and labelling, including any instructions for use; | / | /
b） 产品规范； | b) specifications for product; | / | /
c） 制造、包装、贮存、处置和流通的规范或程序； | c) specifications or procedures for manufacturing, packaging, storage, handling and distribution; | / | /
d） 测量和监视程序； | d) procedures for measuring and monitoring; | / | /
e） 适当时，安装要求； | e) as appropriate, requirements for installation; | / | /
f） 适当时，服务程序。 | f) as appropriate, procedures for servicing. | / | /
4.2.4 文件控制 | 4.2.4 Control of documents | 4.2.3 Control of documents.- | 4.2.3文件控制
质量管理体系所要求的文件应予控制。记录是一种特殊类型的文件，应依据4.2.5的要求进行控制。 | Documents required by the quality management system shall be controlled. Records are a special type of document and shall be controlled according to the requirements given in 4.2.5. | Documents required by the quality management system shall be controlled. Records are a special type of document and shall be controlled according to the requirements given in the control of records. Documents shall be approved, signed and dated by the appropriate and the authorised person. A documented procedure shall be established to define the controls needed.-shall be controlled according to the requirements given in the control of records. Documents shall be approved, signed and dated by the appropriate and the authorised person. | 应控制质量管理体系所要求的文件。记录是一种特殊类型的文件，应按记录控制的要求进行控制。文件应由适当的和被授权的人批准、签署和注明日期。应建立形成文件的程序，以规定所需的控制应按记录控制中的要求进行控制。文件应由适当的和被授权的人批准、签署和注明日期。
形成文件的程序应规定以下方面所需的控制： | A documented procedure shall define the controls needed to: | A documented procedure shall be established to define the controls needed.- | 应建立形成文件的程序，以规定所需的控制
a） 为使文件充分和适宜，文件发布前得到评审和批准； | a) review and approve documents for adequacy prior to issue; | (a) to review and approve documents for adequacy prior to issue; | (a)在文件发布前审查和批准文件的充分性;
b） 必要时对文件进行评审与更新，并再次批准； | b) review, update as necessary and re-approve documents; | (b) to review and update as necessary and re-approve documents; | (b)必要时审查和更新文件，并重新批准文件;
c） 确保文件的现行修订状态和更改得到识别； | c) ensure that the current revision status of and changes to documents are identified; | (c) to ensure that changes and the current revision status of documents are identified; | (c)确保识别文件的变更和当前修订状态;
d） 确保在使用处可获得适用文件的有关版本； | d) ensure that relevant versions of applicable documents are available at points of use; | (d) to ensure that relevant versions of applicable documents are available at points of use; | (d)确保在使用点可获得适用文件的相关版本;
e） 确保文件保持清晰、易于识别； | e) ensure that documents remain legible and readily identifiable; | (e) to ensure that documents remain legible and readily identifiable; | (e)确保文件保持易读和易于识别;
f） 确保组织所确定的策划和运行质量管理体系所需的外来文件得到识别，并控制其分发； | f) ensure that documents of external origin, determined by the organization to be necessary for the planning and operation of the quality management system, are identified and their distribution controlled; | (f) to ensure that documents of external origin are identified and their distribution controlled; and | (f)确保外来文件得到识别并控制其分发;和
g） 防止文件的损坏或丢失； | g) prevent deterioration or loss of documents; | / | /
h） 防止作废文件的非预期使用，对这些文件进行适当的标识。 | h) prevent the unintended use of obsolete documents and apply suitable identification to them. | (g) to prevent the unintended use of obsolete documents, and to apply suitable identification to them if they are retained for any purpose. | (g)防止对作废文件的意外使用，并对作废文件进行适当的标识为任何目的保留。
组织应确保文件的更改得到原审批部门或指定的其他审批部门的评审和批准，被指定的审批部门应能获取用于做出决定的相关背景资料。 | The organization shall ensure that changes to documents are reviewed and approved either by the original approving function or another designated function that has access to pertinent background information upon which to base its decisions. | Changes to document shall be reviewed and approved. Change records shall be maintained which will include a description of the change, identification of the affected documents, the signature of the approving individual, the approval date, and when the change becomes effective. The manufacturer shall ensure that changes to documents are reviewed and approved either by the original approving functionary or another designated functionary which has access to pertinent background information upon which to base its decisions. | 对文件的变更应进行评审和批准。应保持变更记录，包括变更的描述、受影响文件的识别、批准人员的签字、批准日期和变更生效的时间。 制造商应确保对文件的更改由原批准人员或其他指定的人员进行审查和批准，该人员有权获得相关的背景信息，以作为其决策的基础。
对于至少应保存一份的作废文件，组织应规定其保存期限。此期限应确保至少在组织所规定的医疗器械寿命期内，可得到这些医疗器械的制造和试验的文件，而且还应不少于记录（见4.2.5）或适用的法规要求所规定的保存期限。 | The organization shall define the period for which at least one copy of obsolete documents shall be retained. This period shall ensure that documents to which medical devices have been manufactured and tested are available for at least the lifetime of the medical device as defined by the organization, but not less than the retention period of any resulting record (see 4.2.5), or as specified by applicable regulatory requirements. | The manufacturer shall define the period for which at least one copy of obsolete controlled documents shall be retained. This period shall ensure that documents to which medical devices or in vitro diagnostic medical devices have been manufactured and tested are retained for at least one year after the date of expiry of the medical device or in vitro diagnostic medical devices as defined by the manufacturer. | 制造商应规定至少保留一份作废控制文件的期限。 该期限应确保制造和测试医疗器械或体外诊断医疗器械的文件在制造商规定的医疗器械或体外诊断医疗器械的有效期届满后至少保留一年。
4.2.5 记录控制 | 4.2.5 Control of records | 4.2.4 Control of records.- | 4.2.4记录控制
应保持记录以提供符合要求和质量管理体系有效运行的证据。 | Records shall be maintained to provide evidence of conformity to requirements and of the effective operation of the quality management system. | Records shall be established and maintained to provide evidence of conformity to the requirements and of the effective operation of the quality management system. | 应建立并保持记录，以提供符合要求和质量管理体系有效运行的证据。
组织应建立程序并形成文件，以规定记录的标识、存储、安全和完整性、检索、保留时间和处置所需的控制。 | The organization shall document procedures to define the controls needed for the identification, storage, security and integrity, retrieval, retention time and disposition of records. | Records shall remain legible, readily identifiable and retrievable. A documented procedure shall be established to define the controls needed for the identification, storage, protection, retrieval, retention time and disposition of records. The manufacturer shall retain the records for a period of time at least one year after the date of expiry of the medical device or in vitro diagnostic medical devices as defined by the manufacturer, but not less than two years from the date of product release by the manufacturer. | 记录应保持清晰，易于识别和检索。应建立形成文件的程序，规定记录的识别、储存、保护、检索、保留时间和处置所需的控制。 生产企业应当自医疗器械或者生产企业规定的体外诊断医疗器械使用期限届满之日起至少1年，但自生产企业放行产品之日起不少于2年。
按照适用的法规要求，组织应对记录中包含的保密健康信息规定并实施保护方法。 | The organization shall define and implement methods for protecting confidential health information contained in records in accordance with the applicable regulatory requirements.
记录应保持清晰、易于识别和检索。记录的更改应保持可识别。 | Records shall remain legible, readily identifiable and retrievable. Changes to a record shall remain identifiable.
组织应保持记录的期限至少为组织所规定的或适用的法规要求所规定的医疗器械的寿命期，而且还应从组织放行医疗器械起不少于两年。 | The organization shall retain the records for at least the lifetime of the medical device as defined by the organization, or as specified by applicable regulatory requirements, but not less than two years from the medical device release by the organization.
5 管理职责 | 5 Management responsibility | 5 Management responsibility.- | 5 管理职责
5.1 管理承诺 | 5.1 Management commitment | 5.1 Management commitment: | 5.1管理承诺:
最高管理者应通过以下活动，对其建立、实施质量管理体系并保持其有效性的承诺提供证据； | Top management shall provide evidence of its commitment to the development and implementation of the quality management system and maintenance of its effectiveness by: | Top management of the manufacturer shall provide evidence of its commitment to the development and implementation of the quality management system and maintaining its effectiveness by:- | 制造商的最高管理者应提供证据，证明其致力于开发和实施质量管理体系，并通过以下方式保持其有效性
a） 向组织传达满足顾客要求以及适用的法规要求的重要性； | a) communicating to the organization the importance of meeting customer | (a) communicating to the employees the importance of meeting customer as well as statutory and regulatory requirements; | (a)向员工传达满足顾客以及法律法规要求的重要性;
b） 指定质量方针； | as well as applicable regulatory requirements; | (b) establishing the quality policy; | (b)制定质量方针;
c） 确保制定质量目标； | b) establishing the quality policy; | (c) ensuring that quality objectives are established; | (c)确保质量目标的确立;
d） 进行管理评审； | c) ensuring that quality objectives are established; | (d) conducting management reviews; and | (d)进行管理评审;和
e） 确保资源的可获得性。 | d) conducting management reviews; | (e) ensuring the availability of resources. | (e)确保资源的供应。
5.2 以顾客为关注焦点 | 5.2 Customer focus | 5.2 Customer focus: | 5.2 以顾客为关注焦点
最高管理者应确保顾客要求和适用的法规要求得到确定和满足。 | Top management shall ensure that customer requirements and applicable regulatory requirements are determined and met. | Top management of the manufacturer shall ensure that customer requirements are determined and are met. | 最高管理者应确保顾客要求和适用的法规要求得到确定和满足。
5.3 质量方针 | 5.3 Quality policy | 5.3 Quality policy: | 5.3质量方针:
最高管管理者应确保质量方针： | Top management shall ensure that the quality policy: | Top management of the manufacturer shall ensure that the quality policy:- | 制造商的最高管理者应确保质量方针
a) 适应组织的宗旨； | a) is applicable to the purpose of the organization; | (a) is appropriate to the purpose of the manufacturing facility; | (a)适合生产工厂的宗旨;
b) 包括对满足要求和保持质量管理体系有效性的承诺； | b) includes a commitment to comply with requirements and to maintainthe effectiveness of the quality management system; | (b) includes a commitment to comply with requirements and to maintain the effectiveness of the quality management system; | (b)包括遵守要求和保持质量管理体系有效性的承诺;
c) 为制定和评审质量目标提供框架； | c) provides a framework for establishing and reviewing quality objectives; | (c) provides a framework for establishing and reviewing quality objectives; | (c)提供建立和评审质量目标的框架;
d) 在组织内得到沟通和理解； | d) is communicated and understood within the organization; | (d) is communicated and understood within the manufacturer’s organisation; and | (d)在制造商组织内部进行沟通和理解;和
e) 在持续适宜性方面得到评审。 | e) is reviewed for continuing suitability. | (e) is reviewed for continuing suitability. | (e)是否持续适宜进行审查。
5.4 策划 | 5.4 Planning | 5.4 Planning.- | 5.4 策划
5.4.1 质量目标 | 5.4.1 Quality objectives | 5.4.1 Quality objectives: | 5.4.1 质量目标
最高管理者应确保在组织的相关职能和层次上建立质量目标，质量目标包括满足适用的法规要求和产品要求所需的内容。质量目标应是可测量的，并与质量方针保持一致。 | Top management shall ensure that quality objectives, including thoseneeded to meet applicable regulatory requirements and requirements forproduct, are established at relevant functions and levels within the organization.The quality objectives shall be measurable and consistent with thequality policy. | Top management of the manufacturer shall ensure that quality objectives, including those needed to meet requirements for product, are established at relevant functions and levels within the manufacturing organization. The quality objectives shall be measurable and consistent with the quality policy. | 制造商的最高管理者应确保在制造组织的相关职能和层次上建立质量目标，包括满足产品要求所需的质量目标。质量目标应是可测量的，并与质量方针相一致。
5.4.2 质量管理体系策划 | 5.4.2 Quality management system planning | 5.4.2 Quality management system planning: | 5.4.2质量管理体系策划:
最高管理者应确保： | Top management shall ensure that: | Top management of the manufacturer shall ensure that.- | 制造商的最高管理者应确保-
a) 对质量管理体系进行策划，以满足4.1的要求以及质量目标； | a) the planning of the quality management system is carried out in orderto meet the requirements given in 4.1, as well as the quality objectives; | (a) the planning of the quality management system is carried out in order to meet the specified requirements, as well as the quality objectives; and | (a)进行质量管理体系的策划，以满足规定的要求和质量目标;和
b) 在策划和实施质量管理体系变更时保持其完整性。 | b) the integrity of the quality management system is maintainedwhen changes to the quality management system are planned andimplemented. | (b) the integrity of the quality management system is maintained when changes to the quality management system are planned and implemented. | (b)在策划和实施质量管理体系变更时，保持质量管理体系的完整性。
5.5 职责、权限与沟通 | 5.5 Responsibility, authority and communication | 5.5 Responsibility, authority and communication.- | 5.5责任、权限和沟通
5.5.1 职责和权限 | 5.5.1 Responsibility and authority | 5.5.1 Responsibility and authority: | 5.5.1职责和权限:
最高管理者应确保职责和权限得到规定、形成文件并在组织内沟通 | Top management shall ensure that responsibilities and authorities are defined, documented and communicated within the organization. | Top management of the manufacturer shall ensure that responsibilities and authorities are defined, documented and communicated within the manufacturing organisation. | 制造商的最高管理者应确保责任和权限在生产组织内被定义、形成文件并进行沟通。
最高管理者应将所有从事对质量有影响的管理、执行和验证工作的人员的相互关系形成文件，并应确保其完成这些任务所必要的独立性和权限。 | Top management shall document the interrelation of all personnel who manage,perform and verify work affecting quality and shall ensure the independenceand authority necessary to perform these tasks. | Top management of the manufacturer shall establish the interrelation of all personnel who manage, perform and verify work affecting quality, and shall ensure the independence and authority necessary to perform these tasks. | 制造商的最高管理者应建立所有管理、执行和验证影响质量工作的人员之间的相互关系，并确保执行这些任务所需的独立性和权威性。
5.5.2 管理者代表 | 5.5.2 Management representative | 5.5.2 Management representative: | 5.5.2管理者代表:
最高管理者应在管理层中指定一名成员，无论该成员在其他方面的职责如何，应使其具有以下方面的职责和权限： | Top management shall appoint a member of management who, irrespectiveof other responsibilities, has responsibility and authority that includes: | Top management shall appoint a member of management who, irrespective of other responsibilities, shall have responsibility and authority that includes:- | 最高管理者应任命一名管理人员，无论其其他职责如何，其职责和权限应包括
a) 确保将质量管理体系所需的过程形成文件； | a) ensuring that processes needed for the quality management system are documented; | (a) ensuring that processes needed for the quality management system are established, implemented and maintained; | (a)确保质量管理体系所需过程的建立、实施和保持;
b) 向最高管理者报告质量管理体系的有效性和任何改进的需求； | b) reporting to top management on the effectiveness of the quality management system and any need for improvement; | (b) reporting to top management on the performance of the quality management system and any need for improvement; and | (b)向最高管理者报告质量管理体系的绩效和任何改进需求;和
c) 确保在整个组织内提高满足适用的法规要求和质量管理体系要求的意识。 | c) ensuring the promotion of awareness of applicable regulatory requirements and quality management system requirements throughout the organization. | (c) ensuring the promotion of awareness of regulatory and customer requirements throughout the manufacturing organization. | 确保在整个生产组织中提高对法规和顾客要求的认识。
5.5.3 内部沟通 | 5.5.3 Internal communication | 5.5.3 Internal communication: | 5.5.3 内部沟通
最高管理者应确保在组织内建立适当的沟通过程，并确保对质量管理体系的有效性进行沟通。 | Top management shall ensure that appropriate communication processesare established within the organization and that communication takes placeregarding the effectiveness of the quality management system. | Top management shall ensure that appropriate communication processes are established within the Manufacturing organisation and that communication takes place regarding the effectiveness of the quality management system. | 最高管理者应确保在生产组织内建立适当的沟通过程，并确保就质量管理体系的有效性进行沟通。
5.6 管理评审 | 5.6 Management review | 5.6 Management review.- | 5.6管理评审
5.6.1 总则 | 5.6.1 General | 5.6.1 General: | 5.6.1总则:
组织应将管理评审程序形成文件。最高管理者应按照形成文件的策划的时间间隔对组织的质量管理体系进行评审，以确保其持续的适宜性、充分性和有效性。评审应包括评价改进的机会和质量管理体系变更的需求，包括质量方针和质量目标变更的需求。 | The organization shall document procedures for management review. Topmanagement shall review the organization’s quality management systemat documented planned intervals to ensure its continuing suitability, adequacy,and effectiveness. The review shall include assessing opportunitiesfor improvement and the need for changes to the quality managementsystem, including the quality policy and quality objectives. | Top management shall review the organization’s quality management system, at planned intervals, to ensure its continuing suitability, adequacy and effectiveness. This review shall include assessing opportunities for improvement and the need for changes to the quality management system, including the quality policy and quality objectives. | 最高管理者应按策划的时间间隔对组织的质量管理体系进行评审，以确保其持续适宜性、充分性和有效性。评审应包括评价改进的机会和质量管理体系变更的需要，包括质量方针和质量目标。
应保留管理评审的记录（见4.2.5） | Records from management reviews shall be maintained (see 4.2.5). | Records from management reviews shall be maintained. | 管理评审的记录应予保持。
5.6.2 评审输入 | 5.6.2 Review input | 5.6.2 Review input: | 5.6.2评审输入:
管理评审的输入应包括但不限于由以下方面产生的信息： | The input to management review shall include, but is not limited to, information arising from: | The input to management review shall include information on:- | 管理评审的输入应包括以下信息
a) 反馈； | a) feedback; | (b) customer feedback, | (b)客户反馈;
b) 投诉处置； | b) complaint handling; | / | /
c) 给监管机构的报告； | c) reporting to regulatory authorities; | (c) process performance and product conformity, | (c)过程性能和产品符合性;
d) 审核； | d) audits; | (a) results of audits, | (a)审核结果;
e) 过程的监视和测量； | e) monitoring and measurement of processes; | / | /
f) 产品的监视和测量； | f) monitoring and measurement of product; | / | /
g) 纠正措施； | g) corrective action; | (d) status of preventive and corrective actions, | (d)预防和纠正措施的现况;
h) 预防措施 | h) preventive action;
i) 以往管理评审的跟踪措施； | i) follow-up actions from previous management reviews; | (e) follow-up actions from previous management reviews, | (e)以往管理审查的后续行动;
j) 可能影响质量管理体系的变更； | j) changes that could affect the quality management system; | (f) changes that could affect the quality management system, | (f)可能影响质量管理体系的变更;
k) 改进的建议； | k) recommendations for improvement; | (g) recommendations for improvement, and | (g)改进建议;以及
l) 适用的新的或修订的法规要求。 | l) applicable new or revised regulatory requirements. | (h) new or revised regulatory requirements as and when issued. | (h)发布时新的或修订的监管要求。
5.6.3 评审输出 | 5.6.3 Review output | 5.6.3 Review output: | 5.6.3评审输出:
管理评审的输出应予以记录（见4.2.5）并包括经评审的输入和以下方面有关的任何决定和措施： | The output from management review shall be recorded (see 4.2.5) and include the input reviewed and any decisions and actions related to: | The output from the management review shall include any decisions and actions related to:- | 管理评审的输出应包括与以下方面有关的任何决定和措施
a) 保持质量管理体系及其过程适宜性、充分性和有效性所需的改进； | a) improvement needed to maintain the suitability, adequacy, and effectiveness of the quality management system and its processes; | (a) improvements needed to maintain the effectiveness of the quality management system and its processes, | (a)保持质量管理体系及其过程有效性所需的改进;
b) 与顾客要求有关的产品的改进； | b) improvement of product related to customer requirements; | (b) improvement of product related to customer requirements, and | (b)与顾客要求有关的产品的改进
c) 响应适用的新的或修订的法规要求所需的变更； | c) changes needed to respond to applicable new or revised regulatory requirements; | / | /
d) 资源需求。 | d) resource needs. | (c) resource needs. | (c)资源需要。
6 资源管理 | 6 Resource management | 6 Resource management.- | 6资源管理
6.1 资源提供 | 6.1 Provision of resources | 6.1 Provision of resources: | 6.1提供资源:
组织应确定并提供所需的资源，以： | The organization shall determine and provide the resources needed to: | The manufacturing organization shall determine and provide the resources needed | (a)制造组织应确定并提供所需的资源
a) 实施质量管理体系并保持其有效性； | a) implement the quality management system and to maintain its effectiveness; | (a) to implement the quality management system and to maintain its effectiveness, and | (b)实施质量管理体系并保持其有效性
b) 满足适用的法规要求和顾客要求。 | b) meet applicable regulatory and customer requirements. | (b) to meet regulatory and customer requirements. | (c)满足法规和客户的要求。
6.2 人力资源 | 6.2 Human resources | 6.2 Human resources.- 6.2.1 General: | 6.2人力资源 6.2.1总则:
基于适当的教育、培训、技能和经验，从事影响产品质量工作的人员应是胜任的。 组织应将确立能力、提供所需的培训和确保人员的意识等一个或多个过程形成文件。 | Personnel performing work affecting product quality shall be competent on the basis of appropriate education, training, skills and experience. The organization shall document the process(es) for establishing competence, providing needed training, and ensuring awareness of personnel. | Personnel performing work affecting product quality shall be competent on the basis of appropriate education, training, skills and experience. Number of personnel employed shall be adequate and in direct proportion to the workload. Prior to employment, all personnel, shall undergo medical examination including eye examination, and shall be free from communicable or contagious diseases. Thereafter, they should be medically examined periodically, at least once a year. Records shall be maintained thereof. | 从事影响产品质量工作的人员应在接受过适当的教育、培训、技能和经验的基础上胜任工作。 雇用的人员数目应足够，并与工作量成正比。所有人员在入职前应进行体检，包括眼科检查，并无传染病。此后，他们应定期接受医疗检查，至少每年一次。 应保持记录。
|  | 6.2.2 Competence, awareness and training: | 6.2.2能力、意识和培训:
组织应： | The organization shall: | The manufacturer shall:- | 制造商应:-
a) 确定从事影响产品质量工作的人员所需具备的能力； | a) determine the necessary competence for personnel performing work affecting product quality; | (a) determine the necessary competence for personnel performing work affecting product quality, | (a)确定从事影响产品质量工作的人员的必要胜任能力;
b) 提供培训或采取其他措施以获得或保持所需的能力； | b) provide training or take other actions to achieve or maintain the necessary competence; | (b) provide training or take other actions to satisfy these needs, | (b)提供培训或采取其他行动以满足这些需求;
c) 评价所采取措施的有效性； | c) evaluate the effectiveness of the actions taken; | (c) evaluate the effectiveness of the actions taken, | (c)评价所采取行动的有效性;
d) 确保组织的人员知晓所从事活动的关联性和重要性，以及如何为实现质量目标做出贡献； | d) ensure that its personnel are aware of the relevance and importance of their activities and how they contribute to the achievement of the quality objectives; | (d) ensure that its personnel are aware of the relevance and importance of their activities and how they contribute to the achievement of the quality objectives, | (d)确保其人员意识到其活动的相关性和重要性，以及他们如何为实现质量目标作出贡献;
e) 保留教育、培训、技能和经验的适当记录（见4.2.5） | e) maintain appropriate records of education, training, skills and experience(see 4.2.5). | (e) maintain appropriate records of education, training, skills and experience, and | (e)保持教育、培训、技能和经验的适当记录
注：对于提供培训或采取其他措施的有效性的检查方法应与工作相关的风险相适应。 | NOTE The methodology used to check effectiveness is proportionate to the risk associated with the work for which the training or other action is being provided. | (f) establish documented procedures for identifying training needs and ensure that all personnel are trained to adequately perform their assigned responsibilities. | (f)建立文件化程序，以确定培训需求，并确保所有人员接受培训，以充分履行其分配的职责。
6.3 基础设施 | 6.3 Infrastructure | 6.3 Infrastructure: | 6.3基础设施:
为达到符合产品要求、防止产品混淆和确保产品有序处置，组织应将所需的基础设施的要求形成文件。适当时，基础设施包括： | The organization shall document the requirements for the infrastructure needed to achieve conformity to product requirements, prevent product mix-up and ensure orderly handling of product. Infrastructure includes, as appropriate: | The organisation shall determine, provide and maintain the infrastructure needed to achieve conformity to product requirements. Infrastructure includes, as applicable:- | 组织应确定、提供并保持满足产品要求所需的基础设施。基础设施包括(如适用)
a) 建筑物、工作场所和相关设施； | a) buildings, workspace and associated utilities; | (a) buildings, workspace and associated utilities, | (a)建筑物、工作空间及相关设施;
b) 过程设备（硬件和软件）； | b) process equipment (both hardware and software); | (b) process equipment (both hardware and software), and | (b)工艺设备(硬件和软件)
c) 支持性服务（如运输、通讯或信息系统）。 | c) supporting services (such as transport, communication, or information systems). | (c) supporting services (such as transport or communication). | (c)支援服务(例如运输或通讯)。
若维护活动或缺少维护活动可能影响产品质量，组织应将此类维护活动的要求 包括执行维护活动的时间间隔形成文件。适当时，要求应适用于生产设备、工作环境控制设备和监视测量设备。应保留此类维护的记录（见4.2.5）。 | The organization shall document requirements for the maintenance activities, including the interval of performing the maintenance activities, when such maintenance activities, or lack thereof, can affect product quality. As appropriate, the requirements shall apply to equipment used in production, the control of the work environment and monitoring and measurement. Records of such maintenance shall be maintained (see 4.2.5). | The manufacturer shall establish documented requirements for maintenance activities, including their frequency, when such activities or lack thereof can affect product quality. Records of such maintenance shall be maintained. | 当维护活动或缺乏此类活动会影响产品质量时，制造商应对维护活动建立文件化的要求，包括其频率。此种维护应保持记录。
6.4 工作环境和污染控制 | 6.4 Work environment and contamination control | 6.4 Work environment: | 6.4工作环境:
6.4.1 工作环境 | 6.4.1 Work environment
组织应将为达到符合产品要求所需工作环境的要求形成文件。 | The organization shall document the requirements for the work environment needed to achieve conformity to product requirements. | The organisation shall determine and manage the work environment needed to achieve conformity to product requirements. Following requirements shall apply, namely:- | 组织应确定和管理为满足产品要求所需的工作环境。应适用以下要求，即:-
如果工作环境条件可能对产品质量有不良的影响，组织应将工作环境要求以及监视和控制工作环境的程序形成文件。 | If the conditions for the work environment can have an adverse effect on product quality, the organization shall document the requirements for the work environment and the procedures to monitor and control the work environment. | (b) if work environment conditions can have an adverse effect on product quality, the manufacturer shall establish documented requirements as per Annexure- ‘A’ of this Schedule for the work environment conditions and documented procedures or work instructions to monitor and control these work environment condition; | (b)如果工作环境条件可能对产品质量产生不利影响，制造商应根据本附表附件A对工作环境条件和文件化的程序或工作指导书建立文件化的要求，以监视和控制这些工作环境条件;
组织应： | The organization shall:
a）将特定人员的健康、清洁和着装要求形成文件，此类人员与产品或工作环境的接触可能影响医疗器械的安全或性能； | a) document requirements for health, cleanliness and clothing of personnel if contact between such personnel and the product or work environment could affect medical device safety or performance; | (a) the manufacturer shall establish documented requirements for health, cleanliness and clothing of personnel if contact between such personnel and the product or work environment could adversely affect the quality of the product; | (a)如果这些人员与产品或工作环境的接触可能对产品质量产生不利影响，则制造商应对这些人员的健康、清洁和服装建立文件化要求;
b）确保需要在工作环境内的特殊环境条件下临时工作的所有人员是胜任的或在胜任人员监督下工作。 | b) ensure that all personnel who are required to work temporarily under special environmental conditions within the work environment are competent or supervised by a competent person. | (c) the manufacturer shall ensure that all personnel who are required to work temporarily under special environmental conditions within the work environment are appropriately trained and supervised by a trained person; | (c)制造商应确保所有被要求在工作环境内临时在特殊环境条件下工作的人员都经过适当的培训，并由经过培训的人员进行监督;
注：更多信息见ISO14644和ISO14698。 | NOTE Further information can be found in ISO 14644 and ISO 14698.
6.4.2 污染控制 | 6.4.2 Contamination control
适当时，为了防止对工作环境、人员或产品的污染，组织对受污染或易受污染产品的控制应进行策划并将安排形成文件。 | As appropriate, the organization shall plan and document arrangements for the control of contaminated or potentially contaminated product in order to prevent contamination of the work environment, personnel, or product. | (d) if appropriate, special arrangements shall be established and documented for the control of contaminated or potentially contaminated product in order to prevent contamination of other product, the work environment or personnel; | (d)为防止对其他产品、工作环境或人员的污染，在适当情况下，应对受污染或可能受污染的产品的控制作出特别安排并形成文件;
/ | / | (e) all personnel shall bear clean body covering appropriate to their duties. Smoking, eating, drinking, chewing or keeping food and drink shall not be permitted in production, laboratory and storage areas. | (e)所有人员应穿戴与其职务相称的干净衣服。禁止在生产、实验室和储存区域吸烟、吃、喝、嚼或者存放食品和饮料。
对于无菌医疗器械，组织应将控制微生物或微粒物污染的要求形成文件，在组装或包装过程中保持所要求的洁净度。 | For sterile medical devices, the organization shall document requirements for control of contamination with microorganisms or particulate matter and maintain the required cleanliness during assembly or packaging processes. | / | /
7 产品实现 | 7 Product realization | 7 Product realisation,- | 7 产品实现，-
7.1 产品实现的策划 | 7.1 Planning of product realization | 7.1 Planning of product realization: | 7.1产品实现的策划:
组织应策划和开发产品实现所需的过程。产品实现的策划应与质量管理体系其他过程的要求相一致。 | The organization shall plan and develop the processes needed for product realization. Planning of product realization shall be consistent with the requirements of the other processes of the quality management system. | The manufacturer shall plan and develop the processes needed for product realization. Planning of product realization shall be consistent with the requirements of the other processes of the quality management system. | 制造商应策划和开发产品实现所需的过程。产品实现的策划应与质量管理体系其他过程的要求相一致。
组织应在产品实现过程中，将风险管理的一个或多个过程形成文件。应保留风险管理活动的记录（见4.2.5）。 | The organization shall document one or more processes for risk management in product realization. Records of risk management activities shall be maintained (see 4.2.5). | / | /
在策划产品实现的过程中，适当时，组织应确定以下方面的内容： | In planning product realization, the organization shall determine the following, as appropriate: | In planning product realisation, the manufacturer shall determine the following, as appropriate:- | 在策划产品实现时，制造商应酌情确定以下内容
a） 产品的质量目标和要求； | a) quality objectives and requirements for the product; | (a) quality objectives and requirements for the product; | (a)产品的质量目标和要求;
b） 针对产品建立过程、文件（见4.2.4）和提供资源的需求，包括基础设施和工作环境； | b) the need to establish processes and documents (see 4.2.4) and to provide resources specific to the product, including infrastructure and work environment; | (b) the need to establish processes, documents, and provide resources specific to the product; | (b)针对产品建立流程、文件和提供资源的需求;
c） 针对产品所要求的验证、确认、监视、测量、检验和试验、处置、贮存、流通和可追溯性活动以及产品接收准则； | c) required verification, validation, monitoring, measurement, inspection and test, handling, storage, distribution and traceability activities specific to the product together with the criteria for product acceptance; | (c) required verification, validation, monitoring, inspection and test activities specific to the product and the criteria for product acceptance; | (c)产品所需的验证、确认、监视、检查和试验活动以及产品验收准则;
d） 为实现过程及其产品满足要求提供证据所需的记录（见4.2.5）。 | d) records needed to provide evidence that the realization processes and resulting product meet requirements (see 4.2.5). | (d) records needed to provide evidence that the realisation processes and resulting product meet requirements. | (d)为实现过程和最终产品满足要求提供证据所需的记录。
此策划的输出应以适合于组织运行方式的形式形成文件。 注：更多信息见ISO14971. | The output of this planning shall be documented in a form suitable for the organization’s method of operations. NOTE Further information can be found in ISO 14971. | The output of this planning shall be in a form suitable for the manufacturer’s method of operations. The manufacturer organisation shall establish documented requirements for risk management (as per the IS or ISO 14971) throughout product realisation. Records arising from risk management shall be maintained. | 该计划的输出应采用适合制造商操作方法的形式。 在整个产品实现过程中，制造商组织应建立文件化的风险管理要求(根据IS或ISO 14971)。风险管理产生的记录应予保存。
7.2 与顾客相关的过程 | 7.2 Customer-related processes | 7.2 Customer-related processes.- | 7.2 与顾客相关的过程
7.2.1 产品要求的确定 | 7.2.1 Determination of requirements related to product | 7.2.1 Determination of requirements related to the product: | 7.2.1 产品要求的确定
组织应确定： | The organization shall determine: | The manufacturer shall determine:- | 制造商应确定:-
a） 顾客规定的要求，包括对交付及交付后活动的要求； | a) requirements specified by the customer, including the requirements for delivery and post-delivery activities; | (a) requirements specified by the customer, including the requirements for delivery and post-delivery activities; | (a)顾客规定的要求，包括交付和交付后活动的要求;
b） 顾客虽然没有明示，但规定的用途或已知的预期用途所必需的要求； | b) requirements not stated by the customer but necessary for specified or intended use, as known; | (b) requirements not stated by the customer but necessary for specified or intended use, where known; | (b)顾客没有明示，但已知的规定用途或预期用途所必需的要求;
c） 与产品有关的适用的法规要求； | c) applicable regulatory requirements related to the product; | (c) statutory requirements related to the product, and | (c)与产品有关的法规要求
d） 确保医疗器械的特定性能和安全使用所需的任何用户培训； | d) any user training needed to ensure specified performance and safe use of the medical device; | / | /
e） 组织确定的任何附加要求。 | e) any additional requirements determined by the organization. | (d) any additional requirements determined by the manufacturer. | (d)制造商确定的任何附加要求。
7.2.2 产品要求的评审 | 7.2.2 Review of requirements related to product | 7.2.2 Review of requirements related to the product: | 7.2.2与产品有关的要求的评审:
组织应评审与产品有关的要求。评审应在组织向顾客做出提供产品的承诺（如：提交投标书、接受合同或订单以及接受合同或订单的更改）前进行并应确保； | The organization shall review the requirements related to product. This review shall be conducted prior to the organization’s commitment to supply product to the customer (e.g. submission of tenders, acceptance of contracts or orders, acceptance of changes to contracts or orders) and shall ensure that: | The manufacturer shall review the requirements related to the product. This review shall be conducted prior to the manufacturer's commitment to supply a product to the customer and shall ensure that:- | 制造商应评审与产品有关的要求。评审应在制造商承诺向顾客提供产品之前进行，并应确保
a) 产品要求已得到规定并形成文件； | a) product requirements are defined and documented; | (a) product requirements are defined and documented; | 产品要求已定义并形成文件;
b) 与以前表述不一致的合同或订单要求已得到解决； | b) contract or order requirements differing from those previously expressed are resolved; | (b) contract or order requirements differing from those previously expressed are resolved; and | (b)与先前表述不同的合同或订单要求得到解决;和
c) 满足适用的法规要求； | c) applicable regulatory requirements are met; | / | /
d) 依照7.2.1识别的任何用户培训是可获得的或按计划可获得的； | d) any user training identified in accordance with 7.2.1 is available or planned to be available; | / | /
e) 组织有能力满足规定的要求。 | e) the organization has the ability to meet the defined requirements. | (c) the manufacturer has the ability to meet the defined requirements. | (c)制造商有能力满足规定的要求。
应保留评审结果及评审所引起的措施的记录（见4.2.5）。 | Records of the results of the review and actions arising from the review shall be maintained (see 4.2.5). | Records of the results of the review and actions arising from the review shall be maintained. | 评审结果和评审产生的措施的记录应予保持。
若顾客没有提供形成文件的要求，组织在接受顾客要求前应对顾客要求进行确认。 | When the customer provides no documented statement of requirement, the customer requirements shall be confirmed by the organization before acceptance. | Where the customer provides no documented statement of requirement, the customer requirements shall be confirmed by the manufacturer before acceptance. | 当顾客没有提供文件化的要求时，顾客的要求应在接受前由制造商确认。
若产品要求发生变更，组织应确保相关文件得到修改，并确保相关人员知道已更改的要求。 | When product requirements are changed, the organization shall ensure that relevant documents are amended and that relevant personnel are made aware of the changed requirements. | Where product requirements are changed, the manufacturer shall ensure that relevant documents are amended and that relevant personnel are made aware of the changed requirements. | 当产品要求发生变更时，制造商应确保相关文件得到修改，并确保相关人员知道已变更的要求。
7.2.3 沟通 | 7.2.3 Communication | 7.2.3 Customer communication: | 7.2.3客户沟通:
组织应就以下方面与顾客的沟通进行策划并将安排形成文件 | The organization shall plan and document arrangements for communicating with customers in relation to: | The manufacturer shall determine and implement effective arrangements for communicating with customers in relation to:- | 制造商应确定并实施与顾客沟通的有效安排，包括
a） 产品信息； | a) product information; | (a) product information; | (a)产品信息;
b） 处理问询、合同或订单，包括更改； | b) enquiries, contracts or order handling, including amendments; | (b) enquiries, contracts or order handling, including amendments; | (b)查询、合同或订单处理，包括修改;
c） 顾客反馈，包括投诉； | c) customer feedback, including complaints; | (c) customer feedback, including customer complaints; and | (c)客户反馈，包括客户投诉;和
d） 忠告性通知。 | d) advisory notices. | (d) advisory notices. | d） 忠告性通知。
组织应按照适用的法规要求与监督机构沟通。 | The organization shall communicate with regulatory authorities in accordance with applicable regulatory requirements. | / | /
7.3 设计和开发 | 7.3 Design and development | 7.3 Design and development.- | 7.3设计和开发
7.3.1 总则 | 7.3.1 General | 7.3.1 Design and development planning: | 7.3.1设计开发规划:
组织应将设计和开发程序形成文件。 | The organization shall document procedures for design and development. | The manufacturer shall establish documented procedures for design and development. | 制造商应建立文件化的设计和开发程序。
7.3.2 设计和开发策划 | 7.3.2 Design and development planning
组织应对产品的设计和开发进行策划和控制。适当时，随着设计和开发的进展，应保持并更新设计和开发策划文件。 | The organization shall plan and control the design and development of product. As appropriate, design and development planning documents shall be maintained and updated as the design and development progresses. | The manufacturer shall plan and control the design and development of product. | 制造商应对产品的设计和开发进行策划和控制。
在设计和开发策划期间，组织应将以下方面形成文件： | During design and development planning, the organization shall document: | During the design and development planning, the manufacturer shall determine :- | 在设计和开发计划期间，制造商应确定:-
a） 设计可开发的各个阶段； | a) the design and development stages; | (a) the design and development stages; | (a)设计及发展阶段;
b） 每个设计和开发阶段所需的一个或多个评审； | b) the review(s) needed at each design and development stage; | (b) the review, verification, validation and design transfer activities that are appropriate at each design and development stage; and | (b)适合于每个设计和开发阶段的评审、验证、确认和设计转移活动;和
c） 适合于每个设计和开发阶段的验证、确认和设计转换活动； | c) the verification, validation, and design transfer activities that are appropriate at each design and development stage; | (c) the responsibilities and authorities for design and development. | (c)设计和开发的责任和权限。
d） 设计和开发的职责和权限； | d) the responsibilities and authorities for design and development; | The manufacturer shall manage the interfaces between different groups involved in design and development to ensure effective communication and clear assignment of responsibility. | 制造商应对参与设计和开发的不同小组之间的接口进行管理，以确保有效的沟通和明确的职责分配。
e） 确保设计和开发输出到设计和开发输入的可追溯的方法； | e) the methods to ensure traceability of design and development outputs to design and development inputs; | Planning output shall be documented, and updated as appropriate, as the design and development progresses. | 策划的输出应形成文件，并随着设计和开发的进展进行适当的更新。
f） 所需的资源，包括必要的人员能力。 | f) the resources needed including necessary competence of personnel. | NOTE: Design transfer activities during the design and development process ensure that design and development outputs are verified as suitable for manufacturing before becoming final production specifications. | 注:设计和开发过程中的设计转移活动确保设计和开发输出在成为最终生产规范之前被验证为适合生产。
7.3.3 设计和开发输入 | 7.3.3 Design and development inputs | 7.3.2 Design and development inputs: | 7.3.2设计和开发输入:
应确定与产品要求有关的输入，并保留记录（见4.2.5），这些输入应包括： | Inputs relating to product requirements shall be determined and records maintained (see 4.2.5). These inputs shall include: | Inputs relating to product requirements shall be determined and records maintained. The design requirements relating to a device are appropriate and address the intended use of the device, including the needs of the user and patients. These inputs shall include:- | 应确定与产品要求有关的输入，并保持记录。与器械相关的设计要求是适当的，并满足器械的预期用途，包括用户和患者的需求。 这些输入应包括
a） 根据预期用途所确定的功能、性能、可用性和安全要求； | a) functional, performance, usability and safety requirements, according to the intended use; | (a) functional, performance and safety requirements, according to the intended use; | (a)根据预期用途的功能、性能和安全要求;
b） 适用的法规要求和标准； | b) applicable regulatory requirements and standards; | (b) applicable statutory and regulatory requirements; | (b)适用的法律法规要求;
c） 适用的风险管理的一个或多个输出； | c) applicable output(s) of risk management; | (e) output(s) of risk management. | (e)风险管理的输出。
d） 适当时，来源于以前类似设计的信息； | d) as appropriate, information derived from previous similar designs; | (c) where applicable, information derived from previous similar designs; | (c)在适用的情况下，来自以往类似设计的信息;
e） 产品和过程的设计和开发所必需的其他要求。 | e) other requirements essential for design and development of the product and processes. | (d) other requirements essential for design and development; and | (d)设计和开发所必需的其他要求;和
应对这些输入进行评审，以确保输入是充分和适宜的，并经批准。 | These inputs shall be reviewed for adequacy and approved. | These inputs shall be reviewed for adequacy and approved by designated individual. | 这些输入的充分性应由指定人员审核并批准。
这些要求应完整、清楚，能够被验证或确认，并且不能相互矛盾。 | Requirements shall be complete, unambiguous, able to be verified or validated, and not in conflict with each other. | Requirements shall be complete, unambiguous and not in conflict with each other. | 要求应完整、明确且不相互冲突。
注：更多信息见IEC 62366-1 | NOTE Further information can be found in IEC 62366–1.
7.3.4 设计和开发输出 | 7.3.4 Design and development outputs | 7.3.3 Design and development outputs: | 7.3.3设计和开发输出:
设计和开发输出应： | Design and development outputs shall: | The outputs of design and development shall be provided in a form that enables verification against the design and development input and shall be documented, reviewed, and approved prior to release. Design and development outputs shall:- | 设计和开发的输出应以能够对照设计和开发输入进行验证的形式提供，并应在发布前形成文件、进行评审和批准。设计和开发输出应
a) 满足设计和开发输入的要求； | a) meet the input requirements for design and development; | (a) meet the input requirements for design and development; | (a)满足设计和开发的输入要求;
b) 给出采购、生产和服务提供的适当信息； | b) provide appropriate information for purchasing, production and service provision; | (b) provide appropriate information for purchasing, production and for service provision; | (b)为采购、生产和服务提供适当的信息;
c) 包括或引用产品接收准则； | c) contain or reference product acceptance criteria; | (c) contain or reference product acceptance criteria; and | (c)包含或参考产品验收标准;和
d) 规定产品特性，该特性对于产品的安全和正确使用是必需的。 | d) specify the characteristics of the product that are essential for its safe and proper use. | (d) specify the characteristics of the product that are essential for its safe and proper use. | (d)规定产品的特性，这些特性对其安全和正确使用至关重要。
设计和开发输出的方式应适合于对照设计和开发输入进行验证，设计和开发输出应在发布前得到批准。 | The outputs of design and development shall be in a form suitable for verification against the design and development inputs and shall be approved prior to release. | Records of the design and development outputs shall be maintained. | 应保持设计和开发输出的记录。
应保留设计和开发输出的记录（见4.2.5） | Records of the design and development outputs shall be maintained (see 4.2.5). | Records of design and development outputs can include specifications, manufacturing procedures, engineering drawings, and engineering or research logbooks. | 设计和开发输出的记录可以包括规范、制造程序、工程图纸以及工程或研究日志。
7.3.5 设计和开发评审 | 7.3.5 Design and development review | 7.3.4 Design and development review: | 7.3.5 设计和开发评审
应依据所策划并形成文件的安排，在适宜的阶段对设计和开发进行系统评审，以： | At suitable stages, systematic reviews of design and development shall be performed in accordance with planned and documented arrangements to: | At suitable stages, systematic reviews of design and development shall be performed in accordance with planned arrangements:- | 在适当的阶段，应按照策划的安排对设计和开发进行系统评审
a) 评价设计和开发的结果满足要求的能力； | a) evaluate the ability of the results of design and development to meet requirements; | (a) to evaluate the ability of the results of design and development to meet requirements; and | (a)评价设计和开发结果满足要求的能力;和
b) 识别并提议必要的措施。 | b) identify and propose necessary actions. | (b) to identify any problems and propose necessary actions. | (b)查明任何问题并提出必要的行动建议。
评审的参加者包括与所评审的设计和开发阶段有关的职能的代表以及其他专业人员。应保留评审结果和任何必要措施的记录，包括所评审的设计、涉及的参加者和评审日期（见4.2.5）。 | Participants in such reviews shall include representatives of functions concerned with the design and development stage being reviewed, as well as other specialist personnel. Records of the results of the reviews and any necessary actions shall be maintained and include the identification of the design under review, the participants involved and the date of the review (see 4.2.5). | Participants in such reviews shall include representatives of functions concerned with the design and development stage being reviewed, as well as other specialist personnel. Records of the results of the reviews and any necessary actions shall be maintained. | 这种评审的参与者应包括与被评审的设计和开发阶段有关的职能代表，以及其他专业人员。评审结果和任何必要措施的记录应予保持。
7.3.6 设计和开发验证 | 7.3.6 Design and development verification | 7.3.5 Design and development verification: | 7.3.5设计和开发验证:
为确保设计和开发输出满足设计和开发输入的要求，应依据所策划并形成文件的安排对设计和开发进行验证。 | Design and development verification shall be performed in accordance with planned and documented arrangements to ensure that the design and development outputs have met the design and development input requirements. | Verification shall be performed in accordance with planned arrangements to ensure that the design and development outputs have met the design and development input requirements. Records of the results of the verification and any necessary actions shall be maintained. | 应按照策划的安排进行验证，以确保设计和开发输出满足设计和开发输入的要求。验证结果和任何必要措施的记录应予保持。
组织应将验证计划形成文件，验证计划包括方法、接收准则，适当时包括包含样本量原理的统计技术。 | The organization shall document verification plans that include methods, acceptance criteria and, as appropriate, statistical techniques with rationale for sample size. | / | /
如果预期用途要求医疗器械连接至或通过接口连接至其他一个或多个医疗器械，验证应包括证实当这样连接或通过接口连接时设计输出满足设计输入的要求。应保留验证结果和结论及必要措施的记录（见4.2.4和 4.2.5）。 | If the intended use requires that the medical device be connected to, or have an interface with, other medical device(s), verification shall include confirmation that the design outputs meet design inputs when so connected or interfaced.Records of the results and conclusions of the verification and necessary actions shall be maintained (see 4.2.4 and 4.2.5). | / | /
7.3.7 设计和开发确认 | 7.3.7 Design and development validation | 7.3.6 Design and development validation: | 7.3.6 设计和开发确认
为确保产品能够满足规定的应用要求或预期用途要求，应依据策划并形成文件的安排对设计和开发进行确认。 | Design and development validation shall be performed in accordance with planned and documented arrangements to ensure that the resulting product is capable of meeting the requirements for the specified application or intended use. | Design and development validation shall be performed in accordance with planned arrangements to ensure that the resulting product is capable of meeting the requirements for the specified application or intended use. | 为确保产品能够满足规定的应用要求或预期用途要求，应依据策划并形成文件的安排对设计和开发进行确认。
组织应将确认计划形成文件，确认计划包括方法、接收准则，适当时包括包含样本量原理的统计技术。 | The organization shall document validation plans that include methods, acceptance criteria, and, as appropriate, statistical techniques with rationale for sample size. | / | /
设计确认应选择有代表性产品进行。有代表性产品包括最初的生产单元、批次或其等同品。应记录用于确认的产品选择是理由说明（见4.2.5）。 | Design validation shall be conducted on representative product. Representative product includes initial production units, batches or their equivalents. The rationale for the choice of product used for validation shall be recorded (see 4.2.5). | Design validation shall be performed under defined operating conditions on initial production units, lots, or batches or their equivalence. Design validation shall include software validation and risk analysis, where appropriate validation shall be completed prior to the delivery or implementation of the product. Records of the results of validation and any necessary actions shall be maintained. | 设计验证应在规定的操作条件下对初始生产单元、批次或批次或其等效性进行。设计确认应包括软件确认和风险分析，适当的确认应在产品交付或实施之前完成。 确认结果和任何必要措施的记录应予保持。
作为设计和开发确认的一部分，组织应按照适用的法规要求进行医疗器械临床评价或性能评价。用于临床评价或性能评价的医疗器械不视为放行给顾客使用。 | As part of design and development validation, the organization shall perform clinical evaluations or performance evaluations of the medical device in accordance with applicable regulatory requirements. A medical device used for clinical evaluation or performance evaluation is not considered to be released for use to the customer. | As part of design and development validation, the manufacturer shall perform clinical evaluations and/or evaluation of performance of the medical device or in vitro diagnostic medical devices. | 作为设计和开发验证的一部分，制造商应对医疗器械或体外诊断医疗器械的性能进行临床评估和/或评估。
如果预期用途要求医疗器械连接至或通过接口连接至其他一个或多个医疗器械，确认应包括证实当这样连接或通过接口连接时已满足规定的应用要求或预期用途。 | If the intended use requires that the medical device be connected to, or have an interface with, other medical device(s), validation shall include confirmation that the requirements for the specified application or intended use have been met when so connected or interfaced. | NOTE 1.-If a medical device or in vitro diagnostic medical devices can only be validated following assembly and installation at point of use, delivery is not considered to be complete until the product has been formally transferred to the customer. | 注1。-如果医疗器械或体外诊断医疗器械只能在装配和安装后才能在使用点进行验证，则在产品正式移交给客户之前，不认为交付完成。
确认应在向顾客放行产品使用前完成。 | Validation shall be completed prior to release for use of the product to the customer. | NOTE 2.-Provision of the medical device for purposes of clinical evaluations and/or evaluation of performance is not considered to be delivery. | 注2。-为临床评价和/或性能评价目的而提供医疗器械不被视为交付。
应保留确认应结果和结论及必要措施的记录（见4.2.4和4.2.5）。 | Records of the results and conclusion of validation and necessary actions shall be maintained (see 4.2.4 and 4.2.5). | / | /
7.3.8 设计和开发转换 | 7.3.8 Design and development transfer | / | /
组织应将设计和开发输出向制造转换的程序形成文件。这些程序应确保设计和开发输出在成为最终生产规范前经验证适合于制造并确保生产能力能满足产品要求。 | The organization shall document procedures for transfer of design and development outputs to manufacturing. These procedures shall ensure that design and development outputs are verified as suitable for manufacturing before becoming final production specifications and that production capability can meet product requirements. | / | /
应记录转换的结果和结论（见4.2.5）。 | Results and conclusions of the transfer shall be recorded (see 4.2.5). | / | /
7.3.9 设计和开发更改的控制 | 7.3.9 Control of design and development changes | 7.3.7 Control of design and development changes: | 7.3.7设计和开发更改的控制:
组织应将控制设计和开发更改的程序形成文件。组织应确定更改对于医疗器械功能、性能、可用性、安全、适用的法规要求及其预期用途等的重要程度。 | The organization shall document procedures to control design and development changes. The organization shall determine the significance of the change to function, performance, usability, safety and applicable regulatory requirements for the medical device and its intended use. | Design and development changes shall be identified and records maintained. The changes shall be reviewed, verified and validated, as appropriate, and approved before implementation. The review of design and development changes shall include evaluation of the effect of the changes on constituent parts and product already delivered. Records of the results of the review of changes and any necessary actions shall be maintained. | 设计和开发变更应加以识别并保持记录。变更应在适当时进行评审、核实和确认，并在实施前予以批准。设计和开发变更的评审应包括评价变更对组成部分和已交付产品的影响。变更评审的结果和任何必要措施的记录应予保持。
应识别设计和开发的更改。更改在实施前应经： | Design and development changes shall be identified. Before implementation, the changes shall be:
a) 评审； | a) reviewed;
b) 验证； | b) verified;
c) 适当时，确认； | c) validated, as appropriate;
d) 批准。 | d) approved.
设计和开发更改的评审应包括评价更改对在制的或已交付的组成部件和产品的影响，以及对风险管理的输入或输出和产品实现过程的影响。应保留更改及其评审和任何必要的措施的记录（见4.2.5）。 | The review of design and development changes shall include evaluation of the effect of the changes on constituent parts and product in process or already delivered, inputs or outputs of risk management and product realization processes.Records of changes, their review and any necessary actions shall be maintained (see 4.2.5). | Note.-Each manufacturer shall establish and maintain a Design History File for each type of device. The Design History File shall contain or reference the records necessary to demonstrate that the design was developed in accordance with the approved design plan and the requirements of design and development. | 请注意。-每个制造商应为每种类型的器械建立并保持设计历史文件。设计历史文件应包含或引用必要的记录，以证明设计是按照批准的设计方案和设计开发的要求进行的。
7.3.10 设计和开发文档 | 7.3.10 Design and development files | / | /
组织应为每个医疗器械类型或医疗器械族保留设计和开发文档。该文档应包含或引用形成的记录以证明符合设计和开发要求，该文档还应包含设计和开发更改的记录。 | The organization shall maintain a design and development file for each medical device type or medical device family. This file shall include or reference records generated to demonstrate conformity to the requirements for design and development and records for design and development changes. | / | /
7.4 采购 | 7.4 Purchasing | 7.4 Purchasing,- | 7.4采购,
7.4.1 采购过程 | 7.4.1 Purchasing process | 7.4.1 Purchasing process: | 7.4.1采购流程:
组织应将确保采购的产品符合规定的采购信息的程序形成文件（见4.2.4）。 | The organization shall document procedures (see 4.2.4) to ensure that purchased product conforms to specified purchasing information. | The manufacturer organisation shall establish documented procedures to ensure that purchased product conforms to specified purchase requirements. The type and extent of control applied to the supplier and the purchased product shall be dependent upon the effect of the purchased product on subsequent product realisation or the final product. | 制造商组织应建立形成文件的程序，以确保采购的产品符合规定的采购要求。对供方和采购产品控制的类型和程度应取决于采购产品对随后的产品实现或最终产品的影响。
组织应建立评价和选择供方的准则。准则应： | The organization shall establish criteria for the evaluation and selection of suppliers. The criteria shall be: | / | /
a) 基于供方提供满足组织要求的产品的能力； | a) based on the supplier’s ability to provide product that meets the organizations’ requirements; | / | /
b) 基于供方的绩效； | b) based on the performance of the supplier; | / | /
c) 基于采购产品对医疗器械质量的影响； | c) based on the effect of the purchased product on the quality of the medical device; | / | /
d) 与医疗器械相关风险相适应。 | d) proportionate to the risk associated with the medical device. | / | /
组织应对供方的监视和再评价进行策划。应监视供方满足采购产品的要求的绩效。监视结果应为供方再评价过程提供输入。 | The organization shall plan the monitoring and re-evaluation of suppliers. Supplier performance in meeting requirements for the purchased product shall be monitored. The results of the monitoring shall provide an input into the supplier re-evaluation process. | The manufacturer shall evaluate and select suppliers based on their ability to supply product in accordance with the manufacturer’s requirements. Criteria for selection, evaluation and re-evaluation shall be established. | 制造商应根据供应商按照制造商的要求提供产品的能力来评估和选择供应商。应建立选择、评价和再评价的标准。
对未实现采购要求的供方的处置应与所采购产品有关的风险相适应，并符合使用的法规要求。 | Non-fulfilment of purchasing requirements shall be addressed with the supplier proportionate to the risk associated with the purchased product and compliance with applicable regulatory requirements. | / | /
应保留供方能力或绩效的评价、选择、监视和再评价的结果及由这些活动所引起的任何必要措施的记录（见4.2.5）。 | Records of the results of evaluation, selection, monitoring and re-evaluation of supplier capability or performance and any necessary actions arising from these activities shall be maintained (see 4.2.5). | Records of the results of evaluations and any necessary actions arising from the evaluation shall be maintained. | 评价结果和评价引起的任何必要措施的记录应予保持。
7.4.2 采购信息 | 7.4.2 Purchasing information | 7.4.2 Purchasing information: | 7.4.2采购信息:
拟采购产品的采购信息应表述或引用，适当时包括： | Purchasing information shall describe or reference the product to be purchased, including as appropriate: | Purchasing information shall describe the product to be purchased, including where appropriate:- | 采购信息应描述拟采购的产品，适当时包括
a) 产品规范； | a) product specifications;
b) 产品接收、程序、过程和设备的要求； | b) requirements for product acceptance, procedures, processes and equipment; | (a) requirements for approval of product, procedures, processes and equipment; | (a)产品、程序、工艺和设备的批准要求;
c) 供方人员资格要求； | c) requirements for qualification of supplier personnel; | (b) requirements for qualification of personnel; and | (b)对人员资格的要求;和
d) 质量管理体系要求。 | d) quality management system requirements. | (c) quality management system requirements. | (c)质量管理体系要求。
组织应确保在与供方沟通前所规定的采购要求是充分和适宜的。 | The organization shall ensure the adequacy of specified purchasing requirements prior to their communication to the supplier. | The manufacturer shall ensure the adequacy of specified purchase requirements prior to their communication to the supplier. | 在与供方沟通之前，制造商应确保规定的采购要求的充分性。
适用时，采购信息应包括书面协议，该协议明确了在影响采购产品满足规定的采购要求的能力的任何更改实施前，供方应将采购产品方面的更改通知组织。 | Purchasing information shall include, as applicable, a written agreement that the supplier notify the organization of changes in the purchased product prior to implementation of any changes that affect the ability of the purchased product to meet specified purchase requirements. | / | /
按照7.5.9规定的可追溯性要求的范围和程度，组织应以文件（见4.2.4）和记录（见4.2.5）的形式保持相关采购信息。 | To the extent required for traceability given in 7.5.9, the organization shall maintain relevant purchasing information in the form of documents (see 4.2.4) and records (see 4.2.5). | To the extent required for traceability, the manufacturer shall maintain documents and records of relevant purchasing information. | 在可追溯性要求的范围内，企业应当保持有关采购信息的文件和记录。
7.4.3 采购产品的验证 | 7.4.3 Verification of purchased product | 7.4.3 Verification of purchased product: | 7.4.3采购产品的验证:
组织应建立并实施检验或其他必要的活动，以确保采购的产品满足规定的采购要求。验证活动的范围应基于供方评价结果，并与采购产品有关的风险相适应。 | The organization shall establish and implement the inspection or other activities necessary for ensuring that purchased product meets specified purchasing requirements. The extent of verification activities shall be based on the supplier evaluation results and proportionate to the risks associated with the purchased product. | The manufacturer shall establish and implement the inspection or other activities necessary for ensuring that purchased product meets specified purchase requirements. | 制造商应建立并实施检验或其他必要的活动，以确保采购的产品满足规定的采购要求。
若组织觉察到采购产品的任何更改，组织应确定这些更改是否影响产品实现过程或医疗器械最终产品。 | When the organization becomes aware of any changes to the purchased product, the organization shall determine whether these changes affect the product realization process or the medical device. | / | /
若组织或其顾客拟在供方的现场实施验证，组织应在采购信息中对拟验证的活动和产品放行方法做出规定。 | When the organization or its customer intends to perform verification at the supplier’s premises, the organization shall state the intended verification activities and method of product release in the purchasing information. | Where the manufacturer intends to perform verification at the supplier’s premises, the manufacturer shall state the intended verification arrangements and method of product release in the purchasing information. | 在供方现场进行验证的，应当在采购信息中说明拟进行的验证安排和产品放行方法。
应保留验证记录（见4.2.5）。 | Records of the verification shall be maintained (see 4.2.5). | Records of the verification shall be maintained. | 应保持验证的记录。
7.5 生产和服务提供 | 7.5 Production and service provision | 7.5 Production and service provision.- | 7.5生产和服务提供
7.5.1 生产和服务提供的控制 | 7.5.1 Control of production and service provision | 7.5.1 Control of production and service provision: | 7.5.1生产和服务提供的控制:
|  | 7.5.1.1 General requirements: | 7.5.1.1一般要求:
生产和服务提供应予策划、实施、监视和控制以确保产品符合规范。适当时，生产控制应包括但不限于： | Production and service provision shall be planned, carried out, monitored and controlled to ensure that product conforms to specification. As appropriate, production controls shall include but are not limited to: | The manufacturer shall plan and carry out production and service provision under controlled conditions. Controlled conditions shall include, as applicable:- | 制造商应计划并在受控条件下进行生产和服务提供。如适用，受控条件应包括
|  | (a) the availability of information that describes the characteristics of the product, | (a)描述产品特性的信息的可获得性;
a） 编制生产控制程序和控制方法的文件（见4.2.4）； | a) documentation of procedures and methods for the control of production (see 4.2.4); | (b) the availability of documented procedures, documented requirements, work instructions; and reference materials and reference measurement procedures as necessary; | (b)文件化程序、文件化要求、作业指导书的可得性;必要时提供参考物质和参考测量程序;
b） 基础设施鉴定； | b) qualification of infrastructure; | (c) the use of suitable equipment; | (c)使用合适的设备;
c） 实施过程参数和产品特性的监视和测量； | c) implementation of monitoring and measurement of process parameters and product characteristics; | (d) the availability and use of monitoring and measuring devices; | (d)监测和测量装置的可用性和使用情况;
d） 获得和使用监视和测量设备； | d) availability and use of monitoring and measuring equipment; | (e) the implementation of monitoring and measurement; | (e)监测和测量的实施;
e） 对标记和包装实施规定的操作； | e) implementation of defined operations for labelling and packaging; | (f) the implementation of release, delivery and post-delivery activities; and | (g)标签和包装规定操作的实施。
f） 实施产品放行、交付和交付后活动。 | f) implementation of product release, delivery and post-delivery activities. | (g) the implementation of defined operations for labeling and packaging. | (f)放行、交付和交付后活动的实施;和
组织应为每一台或每一批医疗器械建立并保留记录（见4.2.5），该记录提供了7.5.9规定的可追溯性范围和程度的信息并标明制造数量和批准销售数量。应核验和批准该记录。 | The organization shall establish and maintain a record (see 4.2.5) for each medical device or batch of medical devices that provides traceability to the extent specified in 7.5.9 and identifies the amount manufactured and amount approved for distribution. The record shall be verified and approved. | The manufacturer shall establish and maintain a record for each batch of medical device or in vitro diagnostic medical devices that provides traceability and identifies the amount manufactured and amount approved for distribution. The batch record shall be verified and approved. | 生产企业应当对每批医疗器械或者体外诊断医疗器械建立并保持记录，具有可追溯性，标识生产数量和批准销售数量。批记录应当进行审核和批准。
|  | 7.5.1.2 Control of production and service provision — Specific requirements | 7.5.1.2生产和服务提供的控制——具体要求
7.5.2 产品的清洁 | 7.5.2 Cleanliness of product | 7.5.1.2.1 Cleanliness of product and contamination control: | 7.5.1.2.1产品清洁度及污染控制:
在下列情况下，组织应将产品的清洁或污染控制要求形成文件； | The organization shall document requirements for cleanliness of product or contamination control of product if: | The manufacturer shall establish documented requirements for cleanliness of product if:- | 在以下情况下，制造商应建立产品清洁度的文件化要求
a） 产品在灭菌或使用前有组织进行清洁； | a) product is cleaned by the organization prior to sterilization or its use; | (a) product is cleaned by the manufacturer prior to sterilisation or its use; or | (a)产品在灭菌或使用前由制造商清洗;或
b） 产品是以非无菌形式提供且需在灭菌或使用前进行清洁处理； | b) product is supplied non-sterile and is to be subjected to a cleaning process prior to sterilization or its use; | (b) product is supplied non-sterile to be subjected to a cleaning process prior to sterilisation or its use; or | b)产品是以非无菌形式提供且需在灭菌或使用前进行清洁处理；
c） 产品在灭菌或使用前不能进行清洁处理，使用时其清洁是至关重要的； | c) product cannot be cleaned prior to sterilization or its use, and its cleanliness is of significance in use; | / | /
d） 提供产品为非无菌使用，使用时其清洁时至关重要的； | d) produ ct is supplied to be used non-sterile, and its cleanliness is of significance in use; | (c) product is supplied to be used non-sterile and its cleanliness is of significance in use; or | c) 提供产品为非无菌使用，使用时其清洁时至关重要的；
e） 在制造过程中从产品中除去过程添加物。 | e) process agents are to be removed from product during manufacture. | (d) process agents are to be removed from product during manufacture. | d)在制造过程中从产品中除去过程添加物。
如果产品是按照上述a)或b）的要求进行清洁，则6.4.1中包含的要求不适用于清洁处理前的过程。 | If product is cleaned in accordance with a) or b) above, the requirements contained in 6.4.1 do not apply prior to the cleaning process. | If the product is cleaned in accordance with clause (a) or clause (b) above, the requirements content in clause (a) and (b) of sub-paragraph 6.4 do not apply prior to the cleaning process. | 如果产品按照上述(a)或(b)条款进行清洗，则在清洗过程之前不适用6.4款(a)和(b)条款中的要求内容。
7.5.3 安装活动 | 7.5.3 Installation activities | 7.5.1.2.2 Installation activities: | 7.5.1.2.2安装活动:
适当时，组织应将医疗器械安装要求和安装验证接受准则形成文件。 | The organization shall document requirements for medical device installation and acceptance criteria for verification of installation, as appropriate. | If appropriate, the manufacturer shall establish documented requirements which contain acceptance criteria for installing and verifying the installation of the medical device or in vitro diagnostic medical devices. | 如果合适，制造商应建立文件化的要求，其中包含安装和验证医疗器械或体外诊断医疗器械安装的验收标准。
如果经同意的顾客要求允许除组织或其供方以外的外部方安装医疗器械，则组织应提供医疗器械安装和安装验证的形成文件的要求。 | If the agreed customer requirements allow installation of the medical device to be performed by an external party other than the organization or its supplier, the organization shall provide documented requirements for medical device installation and verification of installation. | If the agreed customer requirements allow installation to be performed other than by manufacturer or its authorised agent, the manufacturer shall provide documented requirements for installation and verification. | 如果商定的客户要求允许非由制造商或其授权的代理进行安装，制造商应提供文件化的安装和验证要求。
应保留由组织或其供方完成的医疗器械安装和安装验证的记录（见4.2.5）。 | Records of medical device installation and verification of installation performed by the organization or its suppliers shall be maintained (see 4.2.5). | Records of installation and verification performed by the manufacturer or its authorized agent shall be maintained. | 应保存由制造商或其授权代理人进行的安装和验证的记录。
7.5.4 服务活动 | 7.5.4 Servicing activities | / | /
如果对医疗器械服务有规定的要求，必要时，组织应将服务程序、所涉及的材料和所涉及的测量形成文件，用于实施服 | If servicing of the medical device is a specified requirement, the organization shall document servicing procedures, reference materials, and reference measurements, as necessary, for performing servicing activities and verifying that product requirements are met. | / | /
务活动并验证产品要求得到满足。 |  | / | /
组织应分析有组织或其供方实施的服务活动记录： | The organization shall analyse records of servicing activities carried out by the organization or its supplier: | / | /
a） 以确定该信息是否作为投诉进行处置； | a) to determine if the information is to be handled as a complaint; | / | /
b） 适当时，为改进过程形成输入。 | b) as appropriate, for input to the improvement process. | / | /
应保留由组织或其供方实施的服务活动的记录（见4.2.5）。 | Records of servicing activities carried out by the organization or its supplier shall be maintained (see 4.2.5). | / | /
7.5.5 无菌医疗器械的专用要求 | 7.5.5 Particular requirements for sterile medical devices | 7.5.1.3 Particular requirements for sterile medical devices: | 7.5.1.3无菌医疗器械的专用要求:
组织应保留每一灭菌批的灭菌过程参数的记录（见4.2.5）。灭菌记录应可追溯到医疗器械的每一生产批。 | The organization shall maintain records of the sterilization process parameters used for each sterilization batch (see 4.2.5). Sterilization records shall be traceable to each production batch of medical devices. | The manufacturer shall maintain records of the process parameters for the sterilization process which was used for each sterilization batch. Sterilization records shall be traceable to each production batch of medical device. | 生产企业应当保存每批灭菌工艺的工艺参数记录。灭菌记录应当溯源到医疗器械的每一个生产批次。
7.5.6 生产和服务提供过程的确认 | 7.5.6 Validation of processes for production and service provision | 7.5.2 Validation of processes for production and service provision.- 7.5.2.1 General: | 7.5.2生产和服务提供过程的确认 7.5.2.1总则:
当生产和服务提供过程的输出不能或不是由后续的监视或测量加以验证，并因此使问题仅在产品使用后或服务交付后才显现时，组织应对任何这样的过程进行确认。确认应证实这些过程具有稳定地实现所策划的结果的能力。 | The organization shall validate any processes for production and service provision where the resulting output cannot be or is not verified by subsequent monitoring or measurement and, as a consequence, deficiencies become apparent only after the product is in use or the service has been delivered.Validation shall demonstrate the ability of these processes to achieve planned results consistently. | The manufacturer shall validate any processes for production and service provision where the resulting output cannot be verified by subsequent monitoring or measurement. This includes any processes where deficiencies become apparent only after the product is in use. Validation shall demonstrate the ability of these processes to achieve planned results. | 当产生的输出不能通过后续的监视或测量来验证时，制造商应对生产和服务提供的任何过程进行验证。这包括只有在产品使用后缺陷才变得明显的任何过程。验证应证明这些过程实现所策划的结果的能力。
组织应将过程确认程序形成文件，过程确认包括： | The organization shall document procedures for validation of processes including: | The manufacturer shall establish arrangements for these processes including, as applicable:- | 制造商应建立这些过程的安排，如适用，包括
a) 为过程的评审和批准所规定的准则； | a) defined criteria for review and approval of the processes; | (a) defined criteria for review and approval of the processes; | (a)确定的审核和批准流程的标准;
b) 设备鉴定和人员资格鉴定； | b) equipment qualification and qualification of personnel; | (b) approval of equipment and qualification of personnel | (b)设备和人员资格的批准
c) 使用特定的方法、程序和接收准则； | c) use of specific methods, procedures and acceptance criteria; | (c) use of specific methods and procedures,; | (c)具体方法和程序的使用;
d) 适当时包括包含样本量原理的统计技术； | d) as appropriate, statistical techniques with rationale for sample sizes; | / | /
e) 记录的要求（见4.2.5）； | e) requirements for records (see 4.2.5); | (d) requirements for records; and | (d)记录要求;和
f) 再确认，包括再确认的准则； | f) revalidation, including criteria for revalidation; | (e) revalidation. | (e)重新生效。
g) 对过程更改的批准。 | g) approval of changes to the processes. | / | /
组织应将用于生产和服务提供的计算机软件应用的确认程序形成文件。此类软件的应用在首次使用前应予确认，适当时，此类软件或其应用更改后也应予确认。有关软件确认和再确认的特定方法和活动应与软件使用有关的风险（包括对产品符合规范的能力的影响）相适应。 | The organization shall document procedures for the validation of the application of computer software used in production and service provision. Such software applications shall be validated prior to initial use and, as appropriate, after changes to such software or its application. The specific approach and activities associated with software validation and revalidation shall be proportionate to the risk associated with the use of the software including the effect on the ability of the product to conform to specifications. | The manufacturer shall establish documented procedures for the validation of the application of computer software (and its changes to such software or its application) for production and service provision that affect the ability of the product conform to specified requirements. Such software applications shall be validated prior to initial use. | 对于影响产品符合规定要求的能力的生产和服务提供，制造商应建立文件化的程序，以验证计算机软件的应用(以及对该软件或其应用的变更)。此类软件应用程序应在首次使用前进行验证。
应保留确认的结果和结论以及确认所采取的必要措施的记录（见4.2.4和4.2.5）。 | Records of the results and conclusion of validation and necessary actions from the validation shall be maintained (see 4.2.4 and 4.2.5). | Records of validation shall be maintained. | 应保持验证记录。
7.5.7 灭菌过程和无菌屏障系统确认的专用要求 | 7.5.7 Particular requirements for validation of processes for sterilization and sterile barrier systems | 7.5.2.2 Particular requirements for sterile medical devices: | 7.5.2.2无菌医疗器械的特殊要求:
组织应将灭菌过程和无菌屏障系统的确认程序形成文件（见4.2.4）。 | The organization shall document procedures (see 4.2.4) for the validation of processes for sterilization and sterile barrier systems. | The manufacturer shall establish documented procedures for the validation of sterilization processes. | 制造商应建立文件化的程序来验证灭菌过程。
灭菌过程和无菌屏障系统应在实施前得到确认，适当时，还应在后续的产品或过程更改实施前得到确认。 应保留确认的结果和结论以及确认所采取的必要措施的记录（见4.2.4和4.2.5）。注：更多信息见ISO11607-1和ISO11607-2。 | Processes for sterilization and sterile barrier systems shall be validated prior to implementation and following product or process changes, as appropriate. Records of the results and, conclusion of validation and necessary actions from the validation shall be maintained (see 4.2.4 and 4.2.5).NOTE Further information can be found in ISO 11607-1 and ISO 11607-2. | Sterilization processes shall be validated prior to initial use. The records of validation of each sterilization process shall be maintained. | 灭菌过程应在初次使用前进行验证。每个灭菌过程的验证记录应予保存。
|  | 7.5.3 Identification and traceability.- | 7.5.3识别和可追溯性
7.5.8 标识 | 7.5.8 Identification | 7.5.3.1 Identification: | 7.5.3.1标识:
组织应将产品标识程序形成文件，并在产品实现的整个过程中使用适当的方法识别产品。 | The organization shall document procedures for product identification and identify product by suitable means throughout product realization. | The manufacturer shall identify the product by suitable means throughout product realization, and shall establish documented procedures for such product identification. The manufacturer shall establish documented procedures to ensure that medical devices and in vitro diagnostic medical devices returned to the manufacturer are identified and distinguished from conforming product. | 在产品实现的整个过程中，制造商应通过适当的方法识别产品，并应建立产品识别的文件化程序。制造商应建立文件化的程序，以确保医疗器械和体外诊断医疗器械从合格产品中被识别和区分。
组织应在产品实现的整个过程中按照监视和测量要求识别产品的状态。在产品的生产、贮存、安装和服务的全过程中应保持产品状态的标识，以确保只有通过所要求的检验和试验或经授权让步放行的产品才能被发送、使用或安装。 | The organization shall identify product status with respect to monitoring and measurement requirements throughout product realization. Identification of product status shall be maintained throughout production, storage, installation and servicing of product to ensure that only product that has passed the required inspections and tests or released under an authorized concession is dispatched, used or installed.
如果有适用的法规要求，组织应将为医疗器械指定唯一器械标识的系统形成文件。 | If required by applicable regulatory requirements, the organization shall document a system to assign unique device identification to the medical device.
组织应建立程序并形成文件以确保返回组织的医疗器械能被识别且能与合格的产品区分开。 | The organization shall document procedures to ensure that medical devices returned to the organization are identified and distinguished from conforming product.
7.5.9 可追溯性 | 7.5.9 Traceability | 7.5.3.2 Traceability.- | 7.5.3.2 可追溯性
7.5.9.1 总则 | 7.5.9.1 General | 7.5.3.2.1 General: | 7.5.3.2.1总则:
组织应将可追溯性程序形成文件。这些程序应依据适用的法规要求规定可追溯性的范围和程序以及拟保留的记录（见4.2.5）。 | The organization shall document procedures for traceability. These procedures shall define the extent of traceability in accordance with applicable regulatory requirements and the records to be maintained (see 4.2.5). | The manufacturer shall establish documented procedures for traceability. Such procedures shall define the extent of product traceability and the records required. Where traceability is a requirement, the manufacturer shall control and record the unique identification of the product. NOTE.-Configuration management is a means by which identification and traceability can be maintained. | 制造商应建立文件化的可追溯性程序。这些程序应规定产品可追溯性的程度和所需的记录。 有可追溯性要求的，生产企业应当控制并记录产品的唯一标识。 请注意。配置管理是一种可以保持识别和可追溯性的手段。
7.5.9.2 植入性医疗器械的专用要求 | 7.5.9.2 Particular requirements for implantable medical devices | 7.5.3.2.2 Particular requirements for active implantable medical devices and implantable medical devices: | 7.5.3.2.2有源植入式医疗器械和植入式医疗器械的特殊要求:
如果所使用的组件、材料和工作环境条件等因素可能导致医疗器械不满足其规定的安全和性能要求，可追溯性所要求的记录应包括这些相关因素的记录。 | The records required for traceability shall include records of components, materials, and conditions for the work environment used, if these could cause the medical device not to satisfy its specified safety and performance requirements. | In defining the records required for traceability, the manufacturer shall include records of all components, materials and work environment conditions, if these could cause the medical device not to satisfy its specified requirements. | 在确定可追溯性所需的记录时，如果可能导致医疗器械不满足其规定要求，制造商应包括所有组件、材料和工作环境条件的记录。
组织应要求提供流通服务的供方或经销商保留医疗器械的流通记录以便追溯，若检查需要，可获得这些记录。 | The organization shall require that suppliers of distribution services or distributors maintain records of the distribution of medical devices to allow traceability and that these records are available for inspection. | The manufacturer shall require that its agents or distributors maintain records of the distribution of active implantable medical devices and implantable medical devices to allow traceability and that such records are available for inspection. | 制造商应要求其代理商或经销商保持有源植入式医疗器械和植入式医疗器械的销售记录，以允许可追溯性，并可提供此类记录以供检查。
应保留货运包装收件人的名字和地址的记录（见 4.2.5）。 | Records of the name and address of the shipping package consignee shall be maintained (see 4.2.5). | Records of the name and address of the shipping package consignee shall be maintained. | 应当保存装运包裹收货人的姓名和地址记录。
/ | / | 7.5.3.3 Status identification: The manufacturer shall identify the product status with respect to monitoring and measurement requirements. The identification of product status shall be maintained throughout production, storage, implant, usage and installation of the product to ensure that only product that has passed the required inspections and tests (or released under an authorized concession) is despatched, used or installed. | 7.5.3.3状态识别: 制造商应根据监视和测量要求识别产品状态。在产品的生产、储存、植入、使用和安装过程中，应保持产品状态标识，以确保只有通过要求的检验和测试(或在授权特许下放行)的产品才被发运、使用或安装。
7.5.10 顾客财产 | 7.5.10 Customer property | 7.5.4 Customer property: | 7.5.4顾客财产:
若顾客财产在组织控制下或由组织使用，对组织使用的或构成产品一部分的顾客财产，组织应予以识别、验证、保护和防护。若任何顾客财产发生丢失、损坏或发现不适用情况，组织应向顾客报告并保留记录（见4.2.5）。 | The organization shall identify, verify, protect, and safeguard customer property provided for use or incorporation into the product while it is under the organization’s control or being used by the organization. If any customer property is lost, damaged or otherwise found to be unsuitable for use, the organization shall report this to the customer and maintain records(see 4.2.5). | The manufacturer shall exercise care with customer property while it is under the manufacturer’s control or being used by the manufacturer. The manufacturer shall identify, verify, protect and safeguard customer property provided for use or incorporation into the product. If any customer property is lost, damaged or otherwise found to be unsuitable for use, this shall be reported to the customer and records maintained. NOTE.-Customer property can include intellectual property or confidential health information. | 当顾客财产在制造商的控制下或由制造商使用时，制造商应谨慎对待。制造商应识别、验证、保护和维护供其使用或纳入产品的顾客财产。如发现顾客财产丢失、损坏或不适合使用，应向顾客报告并保持记录。 请注意。-客户财产可包括知识产权或机密健康信息。
7.5.11 产品保护 | 7.5.11 Preservation of product | 7.5.5 Preservation of product: | 7.5.5产品保护:
在产品的加工、贮存、处置和流通期间，组织应将为产品要求提供防护的程序形成文件。防护应适用于医疗器械的组成部分。 | The organization shall document procedures for preserving the conformity of product to requirements during processing, storage, handling, and distribution. Preservation shall apply to the constituent parts of a medical device. | The manufacturer shall establish documented procedures or documented work instructions for preserving the conformity of product during internal processing and delivery to the intended destination. This preservation shall include identification, handling, packaging, storage and protection. Preservation shall also apply to the constituent parts of a product. | 制造商应建立文件化的程序或文件化的工作指导书，以在内部加工和交付到预定目的地期间保持产品的符合性。这种保存应包括标识、处理、包装、储存和保护。保存也应适用于产品的组成部分。
在产品的加工、贮存、处置和流通期间，当其暴露于预期条件和危险（源）时，组织应通过以下法师放置产品发生变化、污染或损坏： | The organization shall protect product from alteration, contamination or damage when exposed to expected conditions and hazards during processing, storage, handling, and distribution by: | / | /
a）设计和制作适当的包装和货运包装箱； | a) designing and constructing suitable packaging and shipping containers; | / | /
b）如果包装本身不能提供防护，将所需的特殊条件要求形成文件。 | b) documenting requirements for special conditions needed if packaging alone cannot provide preservation. | / | /
如要求特殊条件，其应受控并予以记录（见4.2.5）。 | If special conditions are required, they shall be controlled and recorded (see 4.2.5). | The manufacturer shall establish documented procedures or documented work instructions for the control of product with a limited shelf-life or requiring special storage conditions. Such special storage conditions shall be controlled and recorded. | 制造商应建立文件化的程序或文件化的工作指导书，以控制有限保质期或需要特殊储存条件的产品。这种特殊的储存条件应当加以控制和记录。
7.6 监视和测量设备的控制 | 7.6 Control of monitoring and measuring equipment | 7.6 Control of monitoring and measuring devices: | 7.6监视和测量装置的控制:
组织应确定需实施的监视和测量以及所需的监视和测量设备，为产品符合确定的要求提供证据。 | The organization shall determine the monitoring and measurement to be undertaken and the monitoring and measuring equipment needed to provide evidence of conformity of product to determined requirements. | The manufacturer shall determine the monitoring and measurement to be undertaken and the monitoring and measuring devices needed to provide evidence of conformity of product to determined requirements. | 制造商应确定所要进行的监视和测量，以及为产品符合所确定的要求提供证据所需的监视和测量装置。
组织应建立程序并形成文件，以确保监视和测量活动可行并以与监视和测量要求相一致的方式实施。 | The organization shall document procedures to ensure that monitoring and measurement can be carried out and are carried out in a manner that is consistent with the monitoring and measurement requirements. | The manufacturer shall establish documented procedures to ensure that monitoring and measurement can be carried out and are carried out in a manner that is consistent with the monitoring and measurement requirements. | 制造商应建立形成文件的程序，以确保监视和测量能够进行，并以符合监视和测量要求的方式进行。
为确保结果有效，必要时，测量设备应： | As necessary to ensure valid results, measuring equipment shall: | Where necessary to ensure valid results, measuring equipment shall be:- | 必要时，测量设备应:-
a ) 对照能溯源到国际或国家标准的测量标准，按照规定的时间间隔或在使用前进行校准和（或）检定，当不存在上述标准时，应记录校准或验证依据（见4.2.5）： | a) be calibrated or verified, or both, at specified intervals, or prior to use, against measurement standards traceable to international or national measurement standards: when no such standards exist, the basis used for calibration or verification shall be recorded (see 4.2.5); | (a) calibrated or verified at specified intervals, or prior to use, against measurement standards traceable to Bureau of Indian Standards wherever available ; where no such standards exist, the basis used for calibration or verification shall be recorded; | (a)按照可追溯到印度标准局(Bureau of Indian standards)的测量标准，按规定的时间间隔或在使用前进行校准或验证;没有标准的，应当记录校准、验证的依据;
b ) 必要时得到调整或再调整；应记录这种调整或再调整（见4.2.5）； | b) be adjusted or re-adjusted as necessary: such adjustments or readjustments shall be recorded (see 4.2.5); | (b) adjusted or re-adjusted as necessary; | (b)根据需要进行调整或重新调整;
c ) 具有标识，以确定其校准状态； | c) have identification in order to determine its calibration status; | (c) identified to enable the calibration status to be determined; | (c)加以标识，以便确定校准状态;
d ) 予以防护，防止由于调整使测量结果失效； | d) be safeguarded from adjustments that would invalidate the measurement result; | (d) safeguarded from adjustments that would invalidate the measurement result; | (d)防止使测量结果无效的调整;
e ) 予以保护，防止处置、维护和贮存期间的损坏和衰减。 | e) be protected from damage and deterioration during handling, maintenance and storage. | (e) protected from damage and deterioration during handling, maintenance and storage. | (e)在搬运、维护和储存过程中防止损坏和变质。
组织应按照形成文件的程序执行校准或检定。 | The organization shall perform calibration or verification in accordance with documented procedures. | In addition, the manufacturer shall assess and record the validity of the previous measuring results when the equipment is found not to conform to requirements. The manufacturer shall take appropriate action on the equipment and any product affected. Records of the results of calibration and verification shall be maintained. | 此外，当发现设备不符合要求时，制造商应评估和记录先前测量结果的有效性。制造商应对受影响的设备和任何产品采取适当的措施。校准和验证结果的记录应予保持。
此外，当发现设备不符合要求时，组织应对以往测量结果的有效性进行评定和记录。组织应对该设备和任何受影响的产品采取适当的措施。 | In addition, the organization shall assess and record the validity of the previous measuring results when the equipment is found not to conform to requirements. The organization shall take appropriate action in regard to the equipment and any product affected. | When used in the monitoring and measurement of specified requirements, the ability of computer software to satisfy the intended application shall be confirmed. This shall be undertaken prior to initial use and reconfirmed as necessary. | 当计算机软件用于规定要求的监视和测量时，应确认其满足预期用途的能力。这应在初次使用前进行，并在必要时再次确认。
应保留校准和检定（验证）结果的记录（见4.2.5）。 | Records of the results of calibration and verification shall be maintained (see 4.2.5). | / | /
组织应将用于监视和测量要求的计算机软件应用的确认程序形成文件。此类软件的应用在首次使用前应予确认，适当时，此类软件或其应用更改后也应予确认。有关软件确认和再确认的特定方法和活动应与软件使用有关的风险（包括对产品符合规范的能力的影响）相适应。 | The organization shall document procedures for the validation of the application of computer software used for the monitoring and measurement of requirements. Such software applications shall be validated prior to initial use and, as appropriate, after changes to such software or its application.The specific approach and activities associated with software validation and revalidation shall be proportionate to the risk associated with the use of the software including the effect on the ability of the product to conform to specifications. | / | /
应保留确认的结果和结论以及确认所采取的必要措施的记录（见4.2.4和4.2.5）。注：更多信息见ISO 10012 | Records of the results and conclusion of validation and necessary actions from the validation shall be maintained (see 4.2.4 and 4.2.5). NOTE Further information can be found in ISO 10012. | / | /
8 测量、分析和改进 | 8 Measurement, analysis and improvement | 8 Measurement, analysis and improvement.- | 8 测量、分析和改进
8.1 总则 | 8.1 General | 8.1 General: | 8.1总则:
组织应策划并实施所需的监视、测量、分析和改进过程以： | The organization shall plan and implement the monitoring, measurement, analysis and improvement processes needed to: | The manufacturer shall plan and implement the monitoring, measurement, analysis and improvement processes needed:- | 制造商应策划和实施所需的监视、测量、分析和改进过程
a ) 证实产品的符合性； | a) demonstrate conformity of product; | (a) to demonstrate conformity of the product; | (a)证明产品的合格性;
b ) 确保质量管理体系的符合性； | b) ensure conformity of the quality management system; | (b) to ensure conformity of the quality management system; and | (b)确保质量管理体系的符合性;和
C ) 保持质量管理体系的有效性。 | c) maintain the effectiveness of the quality management system. | (c) to maintain the effectiveness of the quality management system. | (c)保持质量管理体系的有效性。
这应包括对统计技术在内的适当方法及其使用程度的确定。 | This shall include determination of appropriate methods, including statistical techniques, and the extent of their use. | This shall include determination of applicable methods, including statistical techniques, and the extent of their use. Note.-If relevant Indian standards are not available, International standards are applicable. In case no Indian or International standards are available, validated testing process of the manufacturer is applicable. | 这应包括确定包括统计技术在内的适用方法及其使用程度。 请注意。-如果没有相关的印度标准，可以使用国际标准。如果没有印度或国际标准，则适用制造商的验证测试过程。
8.2 监视和测量 | 8.2 Monitoring and measurement | 8.2 Monitoring and measurement.- | 8.2监视和测量
8.2.1 反馈 | 8.2.1 Feedback | 8.2.1 Feedback: | 8.2.1反馈:
作为对质量管理体系有效性的一种测量，组织应收集和监视组织是否满足顾客要求的相关信息，并应将获取和利用这种信息的方法形成文件。 | As one of the measurements of the effectiveness of the quality management system, the organization shall gather and monitor information relating to whether the organization has met customer requirements. The methods for obtaining and using this information shall be documented. | As one of the measurements of the performance of the quality management system, the manufacturer shall monitor information relating to whether the manufacturer has met customer or regulatory requirements. The methods for obtaining and using this information shall be determined. | 作为对质量管理体系绩效的测量之一，制造商应监视与制造商是否满足顾客或法规要求有关的信息。应确定获取和使用这些信息的方法。
组织应将反馈过程程序形成文件。该反馈过程应包括从生产和生产后活动中收集数据的规定。 | The organization shall document procedures for the feedback process. This feedback process shall include provisions to gather data from production as well as post-production activities. | / | /
从反馈过程中收集的信息应用作监视和保护产品要求的风险管理的潜在输入以及产品实现或改进过程的潜在输入。 | The information gathered in the feedback process shall serve as potential input into risk management for monitoring and maintaining the product requirements as well as the product realization or improvement processes. | The manufacturer shall establish a documented procedure for a feedback system to provide early warning of quality problems and for input into the corrective and preventive action processes. | 生产企业应建立文件化的反馈系统程序，提供质量问题的早期预警，并输入到纠正和预防措施过程中。
如果适用的法规要求要求组织从生产后活动获取特定经验，则对该经验的评审应构成反馈过程的一部分。 | If applicable regulatory requirements require the organization to gain specific experience from post-production activities, the review of this experience shall form part of the feedback process. | / | /
8.2.2 投诉处置 | 8.2.2 Complaint handling | / | /
组织应按照适用的法规要求将及时处置投诉的程序形成文件。 | The organization shall document procedures for timely complaint handling in accordance with applicable regulatory requirements. | / | /
这些程序应包括对以下方面的最低要求和职责： | These procedures shall include at a minimum requirements and responsibilities for: | / | /
a)接收和记录信息； | a) receiving and recording information; | / | /
b)评价信息以确定反馈是否构成投诉； | b) evaluating information to determine if the feedback constitutes a complaint; | / | /
c)调查投诉； | c) investigating complaints; | / | /
d)确定是否需要向适当的监管机构报告信息； | d) determining the need to report the information to the appropriate regulatory authorities; | / | /
e)处置与投诉有关的产品； | e) handling of complaint-related product; | / | /
f)确定是否需要启动纠正或纠正措施。 | f) determining the need to initiate corrections or corrective actions. | / | /
如果有任何投诉未经调查，应记录理由。应记录由投诉处置过程形成的任何纠正或纠正措施。 | If any complaint is not investigated, justification shall be documented. Any correction or corrective action resulting from the complaint handlingprocess shall be documented. | / | /
如果一项调查确定是组织外的活动导致了投诉，则应在组织和所涉及的外部方之间交换相关信息。应保留投诉处置记录（见4.2.5）。 | If an investigation determines activities outside the organization contributed to the complaint, relevant information shall be exchanged between the organization and the external party involved. | / | /
8.2.3 向监管机构报告 | 8.2.3 Reporting to regulatory authorities | / | /
如果适用的法规有要求，将符合不良事件报告准则的投诉或发布的忠告性通知报告，则组织应将向有关的监管机构报告的程序形成文件。 | If applicable regulatory requirements require notification of complaints that meet specified reporting criteria of adverse events or issuance of advisory notices, the organization shall document procedures for providing notification to the appropriate regulatory authorities. | / | /
应保留向监管机构报告的记录（见4.2.5）。 | Records of reporting to regulatory authorities shall be maintained (see 4.2.5). | / | /
8.2.4 内部审核 | 8.2.4 Internal audit | 8.2.2 Internal audit: | 8.2.2内部审核:
组织应按策划的时间间隔进行内部审核以确定质量管理体系是否： | The organization shall conduct internal audits at planned intervals to determine whether the quality management system: | The manufacturer shall conduct internal audits at planned intervals to determine whether the quality management system:- | 制造商应按策划的时间间隔进行内部审核，以确定质量管理体系是否
a）符合策划并形成文件的安排、本标准的要求以及组织所确定的质量管理体系要求和适用的法规要求； | a) conforms to planned and documented arrangements, requirements of this International Standard, quality management system requirements established by the organization, and applicable regulatory requirements; | (a) conforms to the planned arrangements, to the requirements of this Schedule and to the quality management system requirements established by the manufacturer; and | (a)符合策划的安排、本附表的要求和制造商制定的质量管理体系要求;和
b）得到有效实施与保持。 | b) is effectively implemented and maintained. | (b) is effectively implemented and maintained. | (b）得到有效实施与保持。
组织应建立程序并形成文件以说明策划和实施审核以及记录和报告审核结果的职责和要求。 | The organization shall document a procedure to describe the responsibilities and requirements for planning and conducting audits and recording and reporting audit results. | / | /
组织应策划审核方案，策划时应考虑拟审核的过程和区域的状况和重要性以及以往的审核结果。应规定并记录审核的准则、范围、时间间隔和方法（见 4.2.5）。审核员的选择和审核的实施应确保审核过程客观公正。审核员不应审核自己的工作。 | An audit program shall be planned, taking into consideration the status and importance of the processes and area to be audited, as well as the results of previous audits. The audit criteria, scope, interval and methods shall be defined and recorded (see 4.2.5). The selection of auditors and conduct of audits shall ensure objectivity and impartiality of the audit process. Auditors shall not audit their own work. | An audit programme shall be planned, taking into consideration the status and importance of the processes and areas to be audited, as well as the results of previous audits. The audit criteria, scope, frequency and methods shall be defined.Selection of auditors and conduct of audits shall ensure objectivity and impartiality of the audit process. Auditors shall not audit their own work. | 应考虑待审核过程和领域的现状和重要性，以及以往审核的结果，策划审核方案。应确定审核标准、范围、频率和方法。审核员的选择和审核的实施应确保审核过程的客观性和公正性。审核员不得审核自己的工作。
应保留审核和审核结果的记录（见 4.2.5），包括过程、受审核区域和结论。 | Records of the audits and their results, including identification of the processes and areas audited and the conclusions, shall be maintained (see 4.2.5). | / | /
负责受审核区域的管理者应确保采取任何必要的纠正和纠正措施，应无不当拖延，以消除所发现的不合格及其原因。后续活动应包括验证所采取的措施并报告验证结果。 | The management responsible for the area being audited shall ensure that any necessary corrections and corrective actions are taken without undue delay to eliminate detected nonconformities and their causes. Follow-up activities shall include the verification of the actions taken and the reporting of verification results. | The responsibilities and requirements for planning and conducting audits, and for reporting results and maintaining records shall be defined in a documented procedure. The management responsible for the area being audited shall ensure that actions are taken without undue delay to eliminate detected nonconformities and their causes. Follow-up activities shall include the verification of the actions taken and the reporting of verification results. | 策划和实施审核以及报告结果和保持记录的职责和要求应在形成文件的程序中加以规定。负责被审核区域的管理者应确保及时采取措施，消除发现的不符合及其原因。跟踪活动应包括对所采取措施的验证和验证结果的报告。
注：更多信息见 ISO 19011。 | NOTE Further information can be found in ISO 19011.
8.2.5 过程的监视和测量 | 8.2.5 Monitoring and measurement of processes | 8.2.3 Monitoring and measurement of processes: | 8.2.3过程的监视和测量:
组织应采用适宜的方法对质量管理体系过程进行监视，并在适当时进行测量。这些方法应证实过程实施策划结果的能力。若未能实施策划结果，适当时，应采取纠正和纠正措施。 | The organization shall apply suitable methods for monitoring and, as appropriate, measurement of the quality management system processes. These methods shall demonstrate the ability of the processes to achieve planed results. When planned results are not achieved, correction and corrective action shall be taken, as appropriate. | The manufacturer shall apply suitable methods for monitoring and, where applicable, measurement of the quality management system processes. These methods shall demonstrate the ability of the processes to achieve planned results.When planned results are not achieved, correction and corrective action shall be taken, as appropriate, to ensure conformity of the product. | 制造商应采用适当的方法对质量管理体系过程进行监视，并在适用时进行测量。这些方法应证明过程实现所策划的结果的能力。当未能达到所策划的结果时，应采取适当的纠正和纠正措施，以确保产品的符合性。
8.2.6 产品的监视和测量 | 8.2.6 Monitoring and measurement of product | 8.2.4 Monitoring and measurement of product.- 8.2.4.1 General requirements: | 8.2.4产品的监视和测量 8.2.4.1一般要求:
组织应对产品的特性进行监视和测量，以验证产品要求已得到满足。这种件事和测量应依据策划并形成文件的安排和形成文件的程序，在产品实现过程的适当阶段进行。 | The organization shall monitor and measure the characteristics of the product to verify that product requirements have been met. This shall be carried out at applicable stages of the product realization process in accordance with the planned and documented arrangements and documented procedures. | The manufacturer shall monitor and measure the characteristics of the product to verify that product requirements have been met. This shall be carried out at appropriate stages of the product realization process in accordance with the planned arrangements and documented procedures. | 制造商应监视和测量产品的特性，以验证产品要求已得到满足。这应根据策划的安排和形成文件的程序，在产品实现过程的适当阶段进行。
应保持符合接收准则的证据。应记录有权放行产品的人员的身份（见4.2.5）。适当时，记录应识别用于执行测量活动的检测设备。 | Evidence of conformity with the acceptance criteria shall be maintained. The identity of the person authorizing release of product shall be recorded(see 4.2.5). As appropriate, records shall identify the test equipment used to perform measurement activities. | Evidence of conformity with the acceptance criteria shall be maintained. Records shall indicate the person(s) authorizing release of product. Product release shall not proceed until the planned arrangements have been satisfactorily completed. | 应保持符合验收准则的证据。记录应注明授权放行产品的人员。在计划的安排已令人满意地完成之前，产品放行不得进行。
在策划并形成文件的安排已圆满完成前不应放行产品和交付服务。 | Product release and service delivery shall not proceed until the planned and documented arrangements have been satisfactorily completed.
对于植入性医疗器械，组织应记录进行任何检验或试验的人员的身份。 | For implantable medical devices, the organization shall record the identity of personnel performing any inspection or testing. | 8.2.4.2 Particular requirement for active implantable medical devices and implantable medical Devices wherever applicable: The manufacturer shall record the identity of personnel performing any inspection or testing. | 8.2.4.2有源植入医疗器械和适用的植入医疗器械的特殊要求: 制造商应记录执行任何检验或测试的人员的身份。
8.3 不合格品控制 | 8.3 Control of nonconforming product | 8.3 Control of nonconforming product | 8.3不合格品控制
8.3.1 总则 | 8.3.1 General
组织应确保对不符合产品要求的产品进行识别和控制，以防止非预期的使用或交付。组织应建立程序并形成文件以规定不合格品控制以及与不合格品识别、记录、隔离、评价和处置有关的职责和权限。不合格的评价因包括确定是否需要调查和通知对不合格负责的所有外部方。 | The organization shall ensure that product which does not conform to product requirements is identified and controlled to prevent its unintended use or delivery. The organization shall document a procedure to define the controls and related responsibilities and authorities for the identification, documentation, segregation, evaluation, and disposition of nonconforming product. The evaluation of nonconformity shall include a determination of the need for an investigation and notification of any external party responsible for the nonconformity. | The manufacturer shall ensure that product which does not conform to product requirements is identified and controlled to prevent its unintended use or delivery. The controls and related responsibilities and authorities for dealing with nonconforming product shall be defined in a documented procedure. | 制造商应确保不符合产品要求的产品得到识别和控制，以防止其非预期的使用或交付。处理不合格品的控制措施及相关职责和权限应在形成文件的程序中加以规定。
应保留不合格的性质以及随后所采取的任何措施的记录，包括评价、任何调查和决策的理由说明（4.2.5）。 | Records of the nature of the nonconformities and any subsequent action taken, including the evaluation, any investigation and the rationale for decisions shall be maintained (see 4.2.5). | / | /
8.3.2 交付前发现不合格品的响应措施 | 8.3.2 Actions in response to nonconforming product detected before delivery | / | /
组织应通过下列一种或几种途径处置不合格品： | The organization shall deal with nonconforming product by one or more of the following ways: | The manufacturer shall deal with nonconforming product by one or more of the following ways: | 制造商应对不合格品采取下列一种或几种处理方法:
a）采取措施以消除已发现的不合格； | a) taking action to eliminate the detected nonconformity; | (a) by taking action to eliminate the detected nonconformity; | (a)采取措施消除发现的不符合;
b）采取措施以防止其原预期的使用或应用； | b) taking action to preclude its original intended use or application; | (b) by authorizing its use, release or acceptance under concession; | (c)采取行动阻止其最初的预期用途或应用。
c）授权让步使用、放行或接收。 | c) authorizing its use, release or acceptance under concession. | (c) by taking action to preclude its original intended use or application. | (b)授权其在特许权下使用、释放或接受;
组织应确保不合格品仅在提供理由、或得到批准和满足适用的法规要求的情况下才能让步接收。应保留让步接收和授权让步人员身份的记录（4.2.5.） | The organization shall ensure that nonconforming product is accepted by concession only if the justification is provided, approval is obtained, and applicable regulatory requirements are met. Records of the acceptance by concession and the identity of the person authorizing the concession shall be maintained (see 4.2.5). | The manufacturer shall ensure that nonconforming product is accepted by concession only if regulatory requirements are met. Records of the identity of the person authorising the concession shall be maintained. | 制造商应确保只有在符合法规要求的情况下，才以让步的方式接受不合格品。应保留授权特许权人的身份记录。
8.3.3 交付后发现不合格品的响应措施 | 8.3.3 Actions in response to nonconforming product detected after delivery | Records of the nature of nonconformities and any subsequent actions taken, including concessions obtained, shall be maintained. | 不符合的性质和随后采取的任何措施的记录，包括所获得的让步，应予以保持。
当交付后或开始使用后发现不合格品时，组织应采取与不合格的影响或潜在影响的程度相适应的措施。应保留所采取措施的记录（4.2.5）。 | When nonconforming product is detected after delivery or use has started, the organization shall take action appropriate to the effects, or potential effects, of the nonconformity. Records of actions taken shall be maintained (see 4.2.5). | When nonconforming product is corrected it shall be subject to re-verification to demonstrate conformity to the requirements. When nonconforming product is detected after delivery or use has started, the manufacturer shall take action appropriate to the effects, or potential effects, of the non-conformity. | 不合格品纠正后，应重新验证，以证明符合要求。当不合格品在交付或开始使用后被发现时，制造商应采取与不合格品的影响或潜在影响相适应的措施。
组织应按照适用的法规要求将忠告性通知的发布程序形成文件。这些程序应能随时付诸实施。应保留与发布忠告性通知相关的措施的记录（见4.2.5）。 | The organization shall document procedures for issuing advisory notices in accordance with applicable regulatory requirements. These procedures shall be capable of being put into effect at any time. Records of actions relating to the issuance of advisory notices shall be maintained (see 4.2.5). | If product needs to be reworked (one or more times), the manufacturer shall document the rework process in a work instruction that has undergone the same authorisation and approval procedure as the original work instruction. Prior to authorisation and approval of the work instruction, a determination of any adverse effect of the rework upon product shall be made and documented. | 如果产品需要返工(一次或多次)，生产企业应将返工过程记录在与原作业指导书经过相同授权和批准程序的作业指导书中。在作业指导书授权和批准之前，应确定返工对产品的任何不利影响，并形成文件。
8.4 数据分析 | 8.4 Analysis of data | 8.4 Analysis of data: | 8.4数据分析:
组织应将确定、收集和分析适当数据的程序形成文件以证实质量管理体系的适宜性、充分性和有效性。这些程序应包括对统计技术在内的适当方法及其使用程度的确定。 | The organization shall document procedures to determine, collect and analyse appropriate data to demonstrate the suitability, adequacy and effectiveness of the quality management system. The procedures shall include determination of appropriate methods, including statistical technique sand the extent of their use. | The manufacturer shall establish documented procedures to determine, collect and analyze appropriate data to demonstrate the suitability and effectiveness of the quality management system and to evaluate whether improvement of the effectiveness of the quality management system can be made. | 制造商应建立形成文件的程序，以确定、收集和分析适当的数据，以证明质量管理体系的适宜性和有效性，并评价是否可以改进质量管理体系的有效性。
数据分析应包括来自监视和测量的结果以及其他有关来源的数据，并至少包括以下方面的输入： | The analysis of data shall include data generated as a result of monitoring and measurement and from other relevant sources and include, at a minimum, input from: | This shall include data generated as a result of monitoring and measurement and from other relevant sources.The analysis of data shall provide information relating to:- | 这应包括由于监视和测量以及从其他有关来源产生的数据。 数据分析应提供以下方面的信息
a)反馈； | a) feedback; | (a) feedback; | (a)反馈;
b)产品要求的符合性； | b) conformity to product requirements; | (b) conformity to product requirements; | (b)符合产品要求;
c)过程和产品的特性及趋势，包括改进的机会； | c) characteristics and trends of processes and product including opportunities for improvement; | (c) characteristics and trends of processes and products including opportunities for preventive action; and | (c)过程和产品的特性和趋势，包括采取预防措施的机会;和
d)供方； | d) suppliers; | (d) suppliers. | (d)供应商。
e)审核； | e) audits; | / | /
f)适当时，服务报告。 | f) service reports, as appropriate. | / | /
如果数据分析表明质量管理体系不是适宜的、充分的或有效的，组织应按照8.5的要求将此分析结果用作改进的输入。 | If the analysis of data shows that the quality management system is not suitable, adequate or effective, the organization shall use this analysis as input for improvement as required in 8.5. | / | /
应保留分析结果的记录（见4.2.5）。 | Records of the results of analyses shall be maintained (see 4.2.5). | Records of the results of the analysis of data shall be maintained. | 数据分析结果的记录应予保持。
8.5 改进 | 8.5 Improvement | 8.5 Improvement.- | 8.5 改进
8.5.1 总则 | 8.5.1 General | 8.5.1 General: | 8.5.1 总则
组织应利用质量方针、质量目标、审核结果、上市后监督、数据分析、纠正措施、预防措施和管理评审来识别和实验任何必要的更改，以确保和保持质量管理体系的持续适宜性、充分性和有效性以及医疗器械的安全和性能。 | The organization shall identify and implement any changes necessary to ensure and maintain the continued suitability, adequacy and effectiveness of the quality management system as well as medical device safety and performance through the use of the quality policy, quality objectives, audit results, post-market surveillance, analysis of data, corrective actions, preventive actions and management review. | The manufacturer shall identify and implement any changes necessary to ensure and maintain the continued suitability and effectiveness of the quality management system through the use of the quality policy, quality objectives, audit results, analysis of data, corrective and preventive actions and management review. | 制造商应识别并实施任何必要的变更，以通过使用质量方针、质量目标、审核结果、数据分析、纠正和预防措施以及管理评审，确保和保持质量管理体系的持续适宜性和有效性。
/ | / | The manufacturer shall establish documented procedures for the issue and implementation of advisory notices. These procedures shall be capable of being implemented at any time. Records of all customer complaint investigations shall be maintained. If investigation determines that the activities outside the manufacturer’s organisation contributed to the customer complaint, relevant information shall be exchanged between the organisations involved. | 制造商应为咨询通知的发布和实施建立文件化的程序。这些程序应能够在任何时候实施。应保留所有客户投诉调查的记录。如果调查确定制造商组织之外的活动导致了客户投诉，则应在涉及的组织之间交换相关信息。
/ | / | If any complaint is not investigated, justification shall be documented. Any correction or corrective action resulting from the compliant handling process shall be documented. Manufacturer shall notify the adverse event to the regulatory authority and establish documented procedures for the same. | 如果投诉未被调查，应将理由记录在案。由合规处理过程产生的任何纠正或纠正措施均应形成文件。制造商应将不良事件通知监管机构，并建立相应的文件化程序。
8.5.2 纠正措施 | 8.5.2 Corrective action | 8.5.2 Corrective action: | 8.5.2纠正措施:
组织应采取措施消除不合格的原因以防止不合格的再发生。组织应采取任何必要的纠正措施，应无不当拖延。纠正措施应与不合格的影响程度相适应。 | The organization shall take action to eliminate the cause of nonconformities in order to prevent recurrence. Any necessary corrective actions shall betaken without undue delay. Corrective actions shall be proportionate to the effects of the nonconformities encountered. | The manufacturer shall take action to eliminate the cause of nonconformities in order to prevent recurrence. Corrective actions shall be appropriate to the effects of the nonconformities encountered. | 制造商应采取措施消除不合格的原因，以防止不合格的再次发生。纠正措施应与所遇到的不符合的影响相适应。
组织应将规定以下方面要求的程序形成文件： | The organization shall document a procedure to define requirements for: | A documented procedure shall be established to define requirements for:- | 应建立形成文件的程序，以确定以下要求
a)评审不合格（包括投诉）; | a) reviewing nonconformities (including complaints); | (a) reviewing nonconformities (including customer complaints); | (a)评审不符合项(包括顾客投诉);
b)确定不合格的原因； | b) determining the causes of nonconformities; | (b) determining the causes of nonconformities; | (b)确定不合格的原因;
c)评价确保不合格不再发生的措施的需求； | c) evaluating the need for action to ensure that nonconformities do not recur; | (c) evaluating the need for action to ensure that nonconformities do not recur; | (c)评价采取措施确保不符合不再发生的必要性;
d)对所需的措施进行策划、形成文件并实施，适当时，包括更新文件； | d) planning and documenting action needed and implementing such action, including, as appropriate, updating documentation; | (d) determining and implementing action needed, including, if appropriate, updating documentation; | (d)确定和实施所需的行动，包括酌情更新文件;
e)验证纠正措施对满足适用的法规要求的能力和对医疗器械的安全和性能无不良影响； | e) verifying that the corrective action does not adversely affect the ability to meet applicable regulatory requirements or the safety and performance of the medical device; | (e) recording of the results of any investigation and of action taken; and | (e)记录任何调查的结果及所采取的行动;和
f)评审所采取的纠正措施的有效性。 | f) reviewing the effectiveness of corrective action taken. | (f) reviewing the corrective action taken and its effectiveness. | (f)审查所采取的纠正措施及其有效性。
应保留任何调查的结果和所采取措施的记录（见4.2.5）。 | Records of the results of any investigation and action taken shall be maintained (see 4.2.5). | / | /
8.5.3 预防措施 | 8.5.3 Preventive action | 8.5.3 Preventive action: | 8.5.3预防措施:
组织应确定措施消除潜在不合格的原因以防止不合格的发生。预防措施应与潜在问题的影响程度相适应。 | The organization shall determine action to eliminate the causes of potential nonconformities in order to prevent their occurrence. Preventive actions shall be proportionate to the effects of the potential problems. | The manufacturer shall determine action to eliminate the causes of potential nonconformities in order to prevent their occurrence. Preventive actions shall be appropriate to the effects of the potential problems. | 制造商应确定消除潜在不合格原因的措施，以防止不合格的发生。预防措施应与潜在问题的影响相适应。
组织应将说明以下发面要求的程序形成文件： | The organization shall document a procedure to describe requirements for: | A documented procedure shall be established to define requirements for | 应建立形成文件的程序，以确定以下方面的要求
a)确定潜在不合格及其原因； | a) determining potential nonconformities and their causes; | (a) determining potential nonconformities and their causes, | (a)确定潜在不合格及其原因;
b)评价防止不合格发生的措施的需求； | b) evaluating the need for action to prevent occurrence of nonconformities; | (b) evaluating the need for action to prevent occurrence of nonconformities, | (b)评价为防止不符合发生而采取措施的必要性;
c)对所需的措施进行策划、形成文件并实施，适当时，包括更新文件； | c) planning and documenting action needed and implementing such action, including, as appropriate, updating documentation; | (c) determining and implementing action needed, | (c)确定和执行所需的行动;
d)验证预防措施对满足适用的法规要求的能力和对医疗器械的安全和性能无不良影响； | d) verifying that the action does not adversely affect the ability to meet applicable regulatory requirements or the safety and performance of the medical device; | (d) recording of the results of any investigations and of action taken, and | (d)记录任何调查和所采取行动的结果，以及
e)适当时，评审所采取的预防措施的有效性。 | e) reviewing the effectiveness of the preventive action taken, as appropriate. | (e) reviewing preventive action taken and its effectiveness. | (e)审查所采取的预防行动及其效力。
应保留任何调查的结果和所采取措施的记录（见4.2.5)。 | Records of the results of any investigations and of action taken shall be maintained (see 4.2.5). | / | /

## ISO13485 & Singapore-general
Health Sciences Authority requires a company or organization to establish and maintain the Quality Management System (QMS) throughout the supply chain to ensure that the quality of medical devices is not adversely affected during their distribution and to ensure that appropriate records are kept. For manufacturers, this is achieved through establishing and maintaining QMS such as those conforming to standards like ISO 13485. This ensures that medical devices manufactured or released for distribution are of appropriate quality. Importers and wholesalers of medical devices of risk categories other than Class A medical devices are required to be certified to Good Distribution Practice for Medical Devices (GDPMDS). Certification to GDPMDS is performed by third party certification bodies that are accredited by the Singapore Accreditation Council (SAC).
Requirements for GDPMDS A Quality Management System shall include documentary evidence of the following: Site master file Control of documents, where there should be suitable control to ensure accuracy, availability, legibility and traceability of documents. Control of records, where the organisation should maintain all records required to illustrate their compliance to GDPMDS All documented requirements; procedures and activities, once established by the organization must be implemented and maintained. The scopes included in the QMS are as following: -Management Responsibility -Resource Management -Premises and Facilities -Secondary Assembly -Traceability -Counterfeit, Adulterated, Unwholesome or Tampered Medical Devices -Complaint Handling -Field Safety Corrective Action -Internal Audit -Outsourced Activities For more details refer to the Medical Device Guidance. GN-33: Guidance on the Application of Singapore Standard Good Distribution Practice for Medical Devices
Distribution records: The person who registers, manufacture, import and supply by wholesale medical devices in Singapore should keep proper and appropriate distribution records. Keeping and maintaining distribution records is an important component of QMS to ensure traceability of products in the market. Distribution records should be maintained for each batch of medical device in order to facilitate the recall of the batch if necessary. For more information on distribution records, please refer to GN-06: Guidance on Distribution Records for Medical Devices

## ISO13485 & Singapore-detail
ISO13485: 2016 | GN-33:Guidance on the Application of Singapore Standard Good Distribution Practice for Medical Devices
1 Scope | 1. SCOPE
This International Standard specifies requirements for a quality management system where an | 1.1 General
organization needs to demonstrate its ability to provide medical devices and related services that | The company should develop a quality management system that conforms to the requirements of SS GDPMDS and demonstrate its ability to maintain the quality and integrity of medical devices throughout the supply chain.
consistently meet customer and applicable regulatory requirements. Such organizations can be involved | 1.2 Application
in one or more stages of the life-cycle, including design and development, production, storage and | SS GDPMDS is applicable to importers, wholesalers and outsourced service providers dealing with medical devices in Singapore. For more information on the scope of SS GDPMDS certification, please refer to Annex A and Annex B of this document.
distribution, installation, or servicing of a medical device and design and development or provision of
associated activities (e.g. technical support). This International Standard can also be used by suppliers
or external parties that provide product, including quality management system-related services to such
organizations.
Requirements of this International Standard are applicable to organizations regardless of their size
and regardless of their type except where explicitly stated. Wherever requirements are specified as
applying to medical devices, the requirements apply equally to associated services as supplied by the
organization.
The processes required by this International Standard that are applicable to the organization, but
are not performed by the organization, are the responsibility of the organization and are accounted
for in the organization’s quality management system by monitoring, maintaining, and controlling the
processes.
If applicable regulatory requirements permit exclusions of design and development controls, this can
be used as a justification for their exclusion from the quality management system. These regulatory
requirements can provide alternative approaches that are to be addressed in the quality management
system. It is the responsibility of the organization to ensure that claims of conformity to this
International Standard reflect any exclusion of design and development controls.
If any requirement in Clauses 6, 7 or 8 of this International Standard is not applicable due to the
activities undertaken by the organization or the nature of the medical device for which the quality
management system is applied, the organization does not need to include such a requirement in its
quality management system. For any clause that is determined to be not applicable, the organization
records the justification as described in 4.2.2.
2 Normative references | 2. NORMATIVE REFERENCES
The following documents, in whole or in part, are normatively referenced in this document and are | Refer to SS GDPMDS
indispensable for its application. For dated references, only the edition cited applies. For undated
references, the latest edition of the referenced document (including any amendments) applies.
ISO 9000:20151), Quality management systems — Fundamentals and vocabulary
3 Terms and definitions | 3. TERMS AND DEFINITIONS
For the purposes of this document, the terms and definitions given in ISO 9000:2015 and the | Refer to SS GDPMDS.
following apply.
3.1 advisory notice
notice issued by the organization, subsequent to delivery of the medical device, to provide supplementary information or to advise on action to be taken in the: — use of a medical device, — modification of a medical device, — return of the medical device to the organization that supplied it, or — destruction of a medical device Note 1 to entry: Issuance of an advisory notice can be required to comply with applicable regulatory requirements.
3.2 authorized representative
natural or legal person established within a country or jurisdiction who has received a written mandate
from the manufacturer to act on his behalf for specified tasks with regard to the latter’s obligations
under that country or jurisdiction’s legislation
[SOURCE: GHTF/SG1/N055:2009, 5.2]
3.3 clinical evaluation
assessment and analysis of clinical data pertaining to a medical device to verify the clinical safety and
performance of the device when used as intended by the manufacturer
[SOURCE: GHTF/SG5/N4:2010, Clause 4]
3.4 complaint
written, electronic or oral communication that alleges deficiencies related to the identity, quality,
durability, reliability, usability, safety or performance of a medical device that has been released from
the organization’s control or related to a service that affects the performance of such medical devices
Note 1 to entry: This definition of “complaint” differs from the definition given in ISO 9000:2015.
3.5 distributor
natural or legal person in the supply chain who, on his own behalf, furthers the availability of a medical
device to the end user
Note 1 to entry: More than one distributor may be involved in the supply chain.
Note 2 to entry: Persons in the supply chain involved in activities such as storage and transport on behalf of the
manufacturer, importer or distributor, are not distributors under this definition.
[SOURCE: GHTF/SG1/N055:2009, 5.3]
3.6 implantable medical device
medical device which can only be removed by medical or surgical intervention and which is intended to:
— be totally or partially introduced into the human body or a natural orifice, or
— replace an epithelial surface or the surface of the eye, and
— remain after the procedure for at least 30 days
Note 1 to entry: This definition of implantable medical device includes active implantable medical device
3.7 importer
natural or legal person in the supply chain who is the first in a supply chain to make a medical device,
manufactured in another country or jurisdiction, available in the country or jurisdiction where it is to
be marketed
[SOURCE: GHTF/SG1/N055:2009, 5.4]
3.8 labelling
label, instructions for use, and any other information that is related to identification, technical
description, intended purpose and proper use of the medical device, but excluding shipping documents
[SOURCE: GHTF/SG1/N70:2011, Clause 4]
3.9 life-cycle
all phases in the life of a medical device, from the initial conception to final decommissioning and
disposal
[SOURCE: ISO 14971:2007, 2.7]
3.1
manufacturer
natural or legal person with responsibility for design and/or manufacture of a medical device with the
intention of making the medical device available for use, under his name; whether or not such a medical
device is designed and/or manufactured by that person himself or on his behalf by another person(s)
Note 1 to entry: This “natural or legal person” has ultimate legal responsibility for ensuring compliance with
all applicable regulatory requirements for the medical devices in the countries or jurisdictions where it is
intended to be made available or sold, unless this responsibility is specifically imposed on another person by the
Regulatory Authority (RA) within that jurisdiction.
Note 2 to entry: The manufacturer’s responsibilities are described in other GHTF guidance documents. These
responsibilities include meeting both pre-market requirements and post-market requirements, such as adverse
event reporting and notification of corrective actions.
Note 3 to entry: “Design and/or manufacture”, as referred to in the above definition, may include specification
development, production, fabrication, assembly, processing, packaging, repackaging, labelling, relabelling,
sterilization, installation, or remanufacturing of a medical device; or putting a collection of devices, and possibly
other products, together for a medical purpose.
Note 4 to entry: Any person who assembles or adapts a medical device that has already been supplied by another
person for an individual patient, in accordance with the instructions for use, is not the manufacturer, provided
the assembly or adaptation does not change the intended use of the medical device.
Note 5 to entry: Any person who changes the intended use of, or modifies, a medical device without acting on
behalf of the original manufacturer and who makes it available for use under his own name, should be considered
the manufacturer of the modified medical device.
Note 6 to entry: An authorized representative, distributor or importer who only adds its own address and
contact details to the medical device or the packaging, without covering or changing the existing labelling, is not
considered a manufacturer.
Note 7 to entry: To the extent that an accessory is subject to the regulatory requirements of a medical device, the
person responsible for the design and/or manufacture of that accessory is considered to be a manufacturer.
[SOURCE: GHTF/SG1/N055:2009, 5.1]
3.11
medical device
instrument, apparatus, implement, machine, appliance, implant, reagent for in vitro use, software,
material or other similar or related article, intended by the manufacturer to be used, alone or in
combination, for human beings, for one or more of the specific medical purpose(s) of:
— diagnosis, prevention, monitoring, treatment or alleviation of disease;
— diagnosis, monitoring, treatment, alleviation of or compensation for an injury;
— investigation, replacement, modification, or support of the anatomy or of a physiological process;
— supporting or sustaining life;
— control of conception;
— disinfection of medical devices;
— providing information by means of in vitro examination of specimens derived from the human body;
and does not achieve its primary intended action by pharmacological, immunological or metabolic
means, in or on the human body, but which may be assisted in its intended function by such means
Note 1 to entry: Products which may be considered to be medical devices in some jurisdictions but not in others
include:
— disinfection substances;
— aids for persons with disabilities;
— devices incorporating animal and/or human tissues;
— devices for in vitro fertilization or assisted reproduction technologies.
[SOURCE: GHTF/SG1/N071:2012, 5.1]
3.12
medical device family
group of medical devices manufactured by or for the same organization and having the same basic
design and performance characteristics related to safety, intended use and function
3.13
performance evaluation
assessment and analysis of data to establish or verify the ability of an in vitro diagnostic medical device
to achieve its intended use
3.14
post-market surveillance
systematic process to collect and analyse experience gained from medical devices that have been placed
on the market
3.15
product
result of a process
Note 1 to entry: There are four generic product categories, as follows:
— services (e.g. transport);
— software (e.g. computer program, dictionary);
— hardware (e.g. engine mechanical part);
— processed materials (e.g. lubricant).
Many products comprise elements belonging to different generic product categories. Whether the product is
then called service, software, hardware or processed material depends on the dominant element. For example,
the offered product “automobile” consists of hardware (e.g. tyres), processed materials (e.g. fuel, cooling liquid),
software (e.g. engine control software, driver’s manual), and service (e.g. operating explanations given by the
salesman).
Note 2 to entry: Service is the result of at least one activity necessarily performed at the interface between the
supplier and customer and is generally intangible. Provision of a service can involve, for example, the following:
— an activity performed on a customer-supplied tangible product (e.g. automobile to be repaired);
— an activity performed on a customer-supplied intangible product (e.g. the income statement needed to
prepare a tax return);
— the delivery of an intangible product (e.g. the delivery of information in the context of knowledge
transmission);
— the creation of ambience for the customer (e.g. in hotels and restaurants).
Software consists of information and is generally intangible and can be in the form of approaches, transactions
or procedures.
Hardware is generally tangible and its amount is a countable characteristic. Processed materials are generally
tangible and their amount is a continuous characteristic. Hardware and processed materials often are referred
to as goods.
Note 3 to entry: This definition of “product” differs from the definition given in ISO 9000:2015.
[SOURCE: ISO 9000:20052), 3.4.2, modified]
3.16
purchased product
product provided by a party outside the organization’s quality management system
Note 1 to entry: The provision of product does not necessarily infer a commercial or financial arrangement.
3.17
risk
combination of the probability of occurrence of harm and the severity of that harm
Note 1 to entry: This definition of “risk” differs from the definition given in ISO 9000:2015.
[SOURCE: ISO 14971:2007, 2.16]
3.18
risk management
systematic application of management policies, procedures and practices to the tasks of analysing,
evaluating, controlling and monitoring risk
[SOURCE: ISO 14971:2007, 2.22]
3.19
sterile barrier system
minimum package that prevents ingress of microorganisms and allows aseptic presentation of the
product at the point of use
[SOURCE: ISO 11607-1:2006, 3.22]
2) Superseded by ISO 9000:2015.
3.2
sterile medical device
medical device intended to meet the requirements for sterility
Note 1 to entry: The requirements for sterility of a medical device can be subject to applicable regulatory
requirements or standards.
4 Quality management system | 4. QUALITY MANAGEMENT SYSTEM
4.1 General requirements | 4.1 General requirements
4.1.1 The organization shall document a quality management system and maintain its effectiveness in accordance with the requirements of this International Standard and applicable regulatory requirements. | The quality management system established should be sufficiently robust to meet external and internal considerations such as changes in regulatory requirements, customer feedback, changes to key personnel, facilities, etc.
The organization shall establish, implement and maintain any requirement, procedure, activity or arrangement required to be documented by this International Standard or applicable regulatory requirements.
The organization shall document the role(s) undertaken by the organization under the applicable regulatory requirements.
NOTE Roles undertaken by the organization can include manufacturer, authorized representative, importer or distributor.
4.1.2 The organization shall:
a) determine the processes needed for the quality management system and the application of these processes throughout the organization taking into account the roles undertaken by the organization;
b) apply a risk based approach to the control of the appropriate processes needed for the quality management system;
c) determine the sequence and interaction of these processes.
4.1.3 For each quality management system process, the organization shall:
a) determine criteria and methods needed to ensure that both the operation and control of these processes are effective;
b) ensure the availability of resources and information necessary to support the operation and monitoring of these processes;
c) implement actions necessary to achieve planned results and maintain the effectiveness of these processes;
d) monitor, measure as appropriate, and analyse these processes;
e) establish and maintain records needed to demonstrate conformance to this International Standard and compliance with applicable regulatory requirements (see 4.2.5).
4.1.4 The organization shall manage these quality management system processes in accordance with the requirements of this International Standard and applicable regulatory requirements. Changes to be made to these processes shall be:
a) evaluated for their impact on the quality management system;
b) evaluated for their impact on the medical devices produced under this quality management system;
c) controlled in accordance with the requirements of this International Standard and applicable regulatory requirements.
4.1.5 When the organization chooses to outsource any process that affects product conformity to requirements, it shall monitor and ensure control over such processes. The organization shall retain responsibility of conformity to this International Standard and to customer and applicable regulatory requirements for outsourced processes. The controls shall be proportionate to the risk involved and the ability of the external party to meet the requirements in accordance with 7.4. The controls shall include written quality agreements. | 14. OUTSOURCED ACTIVITIES
All service provider of outsourced activities should be audited by the certification body unless they are already certified to this standard. Certification bodies will normally conduct an on-site audit for service providers of storage and secondary assembly.
The organisation should control and review the outsourced activities as part of its internal audit. The organisation is ultimately responsible to put in place processes to assure the control of outsourced activities.
Prior to outsourcing activities, the organisation is responsible for assessing the suitability and the competence of the service provider to carry out the outsourced activities. When monitoring the performance of the service provider, the organisation should consider compliance trends and conformance history. The performance of the service provider should be monitored on a regular basis.
R1.1 ► Certain activities relating to cleaning, pest control, installation and servicing, calibration and transportation may be outsourced to third party service providers. The service provider shall be audited as part of the internal audit of the organisation if they have not been certified to this standard. For outsourced activities relating to storage, warehousing, stock handling and secondary assembly, the service provider of outsourced activities shall be audited by the certification body as part of organisation’s system unless they are already certified to this standard. ◄
For outsourced activities covered under Clause 7 (Premises and facilities) and Clause 8 (Secondary assembly) of the SS GDPMDS, there should be a written contract to define the roles and responsibilities of each party, technical specification and contractual agreements to avoid any misunderstandings which could result in a product or work of unsatisfactory quality. Technical specification may include any special storage and handling conditions for storage or transportation, or packing configuration of medical devices to be assembled.
4.1.6 The organization shall document procedures for the validation of the application of computer software used in the quality management system. Such software applications shall be validated prior to initial use and, as appropriate, after changes to such software or its application.
The specific approach and activities associated with software validation and revalidation shall be proportionate to the risk associated with the use of the software.
Records of such activities shall be maintained (see 4.2.5).
4.2 Documentation requirements | 4.2 Documentation requirements
4.2.1 General | 4.2.1 General
The quality management system documentation (see 4.2.4) shall include: a) documented statements of a quality policy and quality objectives; b) a quality manual; c) documented procedures and records required by this International Standard; d) documents, including records, determined by the organization to be necessary to ensure the effective planning, operation, and control of its processes; e) other documentation specified by applicable regulatory requirements. | It is important to recognise that the structure and level of detail required in the procedures should be tailored to the needs of the organisation. It is also dependent on the methodology or approach taken to implement the SS GDPMDS requirements, and the skills and qualifications of the personnel involved in the activities. Procedures or instructions may be presented in any form (text, graphic or audio-visual) or medium.
4.2.2 Quality manual
The organization shall document a quality manual that includes:
a) the scope of the quality management system, including details of and justification for any exclusion or non-application;
b) the documented procedures for the quality management system, or reference to them;
c) a description of the interaction between the processes of the quality management system.
The quality manual shall outline the structure of the documentation used in the quality management system.
/ | 4.2.2 Site master file
Refer to the HSA medical devices regulatory guidance GN-03: Guidance on Preparation of a Site Master File for Licensing which is published on the HAS website.
4.2.4 Control of documents | 4.2.3 Control of documents
Documents required by the quality management system shall be controlled. Records are a special type of document and shall be controlled according to the requirements given in 4.2.5. A documented procedure shall define the controls needed to: a) review and approve documents for adequacy prior to issue; b) review, update as necessary and re-approve documents; c) ensure that the current revision status of and changes to documents are identified; d) ensure that relevant versions of applicable documents are available at points of use; e) ensure that documents remain legible and readily identifiable; f) ensure that documents of external origin, determined by the organization to be necessary for the planning and operation of the quality management system, are identified and their distribution controlled; g) prevent deterioration or loss of documents; h) prevent the unintended use of obsolete documents and apply suitable identification to them. The organization shall ensure that changes to documents are reviewed and approved either by the original approving function or another designated function that has access to pertinent background information upon which to base its decisions. The organization shall define the period for which at least one copy of obsolete documents shall be retained. This period shall ensure that documents to which medical devices have been manufactured and tested are available for at least the lifetime of the medical device as defined by the organization, but not less than the retention period of any resulting record (see 4.2.5), or as specified by applicable regulatory requirements. | There should be suitable control to ensure accuracy, availability, legibility and traceability of documents.
4.2.5 Control of records | 4.2.4 Control of records
Records shall be maintained to provide evidence of conformity to requirements and of the effective operation of the quality management system. The organization shall document procedures to define the controls needed for the identification, storage, security and integrity, retrieval, retention time and disposition of records. The organization shall define and implement methods for protecting confidential health information contained in records in accordance with the applicable regulatory requirements. Records shall remain legible, readily identifiable and retrievable. Changes to a record shall remain identifiable. The organization shall retain the records for at least the lifetime of the medical device as defined by the organization, or as specified by applicable regulatory requirements, but not less than two years from the medical device release by the organization.Records shall be maintained to provide evidence of conformity to requirements and of the effective | The organisation should maintain all records required to illustrate their compliance to the various relevant clauses in the SS GDPMDS.
Where an electronic records system is used in place of a paper-based system, the system utilised should have built-in checks and balances to ensure the integrity of the records and to protect against unauthorised entries.
Examples of records to be maintained and their expected retention periods:
| SS GDPMDS clauses | Examples of records | Expected retention periods
| 5.3 Management review | Attendance and minutes of management review | 3 years; in line with audit cycle and should be available at point of audit
| 6 Resource management | Training records; attendance and training effectiveness checks | At least for the period of employment and should be available at point of audit
| 7 Premises and facilities | Records of receipt / verification of stock | Projected useful life* of the medical device or 2 years from date of supply; whichever is longer
|  | Records of installation, commissioning and any test results to demonstrate proper installation | Projected useful life* of the medical device or 2 years from date of installation; whichever is longer
|  | Records of servicing activities | Projected useful life* of the medical device or 2 years from date of servicing; whichever is longer
| 8 Secondaryassembly | Records for each batch of medical devices that provide traceability | Projected useful life* of the medical device or 2 years from date of assembly; whichever is longer
|  | Batch assembly records should be kept for each batch or part of the batch assembled | Projected useful life* of the medical device or 3 years from date of assembly; whichever is longer
| 9 Traceability | Delivery orders, distribution or sales records, returns and disposal records, etc. | Projected useful life* of the medical device or 2 years from date of supply and/or disposal; whichever is longer
| 11 Complaint handling | Records of the complaint, investigation and any subsequent actions taken | At least 5 years after the expiry of the projected useful life* of the medical device
| 12 Field safety corrective action (FSCA) | Records of all actions taken in connection with the FSCA and their approval by the company and regulatory authority | Projected useful life* of the medical device or 2 years from date of completion of the FSCA; whichever is longer
| 13 Internal audit | Records of the internal audit and the verifications of actions taken | 3 years; in line with audit cycle and should be available at point of audit
* The projected useful life for medical device is determined by the product owner.
For all other types of records required by SS GDPMDS, the expected retention period should be defined by the organisation based on the type of medical device or nature of activities conducted.
5 Management responsibility | 5. MANAGEMENT RESPONSIBILITY
5.1 Management commitment
Top management shall provide evidence of its commitment to the development and implementation of the quality management system and maintenance of its effectiveness by:
a) communicating to the organization the importance of meeting customer as well as applicable regulatory requirements;
b) establishing the quality policy;
c) ensuring that quality objectives are established;
d) conducting management reviews;
e) ensuring the availability of resources.
5.2 Customer focus
Top management shall ensure that customer requirements and applicable regulatory requirements are determined and met.
5.3 Quality policy
Top management shall ensure that the quality policy:
a) is applicable to the purpose of the organization;
b) includes a commitment to comply with requirements and to maintain the effectiveness of the quality management system;
c) provides a framework for establishing and reviewing quality objectives;
d) is communicated and understood within the organization;
e) is reviewed for continuing suitability.
5.4 Planning
5.4.1 Quality objectives
Top management shall ensure that quality objectives, including those needed to meet applicable regulatory requirements and requirements for product, are established at relevant functions and levels within the organization. The quality objectives shall be measurable and consistent with the quality policy.
5.4.2 Quality management system planning
Top management shall ensure that:
a) the planning of the quality management system is carried out in order to meet the requirements given in 4.1, as well as the quality objectives;
b) the integrity of the quality management system is maintained when changes to the quality management system are planned and implemented.
5.5 Responsibility, authority and communication
5.5.1 Responsibility and authority | 5.1 Responsibility and authority
Top management shall ensure that responsibilities and authorities are defined, documented and communicated within the organization. | The organisation should document job descriptions, including key responsibilities and authorities, within the Quality Management System, e.g. in the Site Master File.
Top management shall document the interrelation of all personnel who manage, perform and verify work affecting quality and shall ensure the independence and authority necessary to perform these tasks.
5.5.2 Management representative | 5.2 Management representative
Top management shall appoint a member of management who, irrespective of other responsibilities,has responsibility and authority that includes: | The appointment of the management representative should be documented, e.g. in the Site Master File or on an official appointment letter. Only one member of management is to be authorised as the organisation’s management representative. A deputy may be appointed in the absence of the management representative.
a) ensuring that processes needed for the quality management system are documented; | There should be no conflict of interest between the other responsibilities and functions that the management representative performs and those relating to the quality management system.
b) reporting to top management on the effectiveness of the quality management system and any need for improvement;
c) ensuring the promotion of awareness of applicable regulatory requirements and quality management system requirements throughout the organization.
5.5.3 Internal communication
Top management shall ensure that appropriate communication processes are established within the organization and that communication takes place regarding the effectiveness of the quality management system.
5.6 Management review | 5.3 Management review
5.6.1 General | 5.3.1 General
The organization shall document procedures for management review. Top management shall review the organization’s quality management system at documented planned intervals to ensure its continuing suitability, adequacy and effectiveness. The review shall include assessing opportunities for improvement and the need for changes to the quality management system, including the quality policy and quality objectives. | The top management should review the quality management system on a regular basis, at least once a year to ensure that the quality management system remains effective.
Records from management reviews shall be maintained (see 4.2.5).
5.6.2 Review input | 5.3.2 Review input
The input to management review shall include, but is not limited to, information arising from: | Please refer to the sub-clause of SS GDPMDS for input to management review.
a) feedback;
b) complaint handling;
c) reporting to regulatory authorities;
d) audits;
e) monitoring and measurement of processes;
f) monitoring and measurement of product;
g) corrective action;
h) preventive action;
i) follow-up actions from previous management reviews;
j) changes that could affect the quality management system;
k) recommendations for improvement;
l) applicable new or revised regulatory requirements.
5.6.3 Review output | 5.3.3 Review output
The output from management review shall be recorded (see 4.2.5) and include the input reviewed and any decisions and actions related to: | Management records may be in any form that suits the organisation. These include formal meeting minutes or meeting notes, which should be produced, distributed and stored on paper or electronically.
a) improvement needed to maintain the suitability, adequacy, and effectiveness of the quality management system and its processes; b) improvement of product related to customer requirements; c) changes needed to respond to applicable new or revised regulatory requirements; d) resource needs. | Records of the management review should include the identity of those taking part in the review, date(s) of the review and all the review input and output aspects listed in the SS GDPMDS. This includes description of any corrective or preventive action to be taken. For any action determined, the responsibility for such actions, the resources, target dates for completion, etc. should be identified.
6 Resource management | 6. RESOURCE MANAGEMENT
6.1 Provision of resources
The organization shall determine and provide the resources needed to:
a) implement the quality management system and to maintain its effectiveness;
b) meet applicable regulatory and customer requirements.
6.2 Human resources | 6.1 Personnel
Personnel performing work affecting product quality shall be competent on the basis of appropriate education, training, skills and experience. | The organisation should consider the experience, qualifications, capabilities and abilities of personnel who are involved in the distribution of medical devices.
The organization shall document the process(es) for establishing competence, providing needed training, and ensuring awareness of personnel. | Personnel working within the quality system require a certain level of competence or training (internal or external) before they can perform the tasks properly. Training may take the form of on-the-job training/coaching by supervisors, instructional briefings, courses, workshops, etc.
The organization shall: a) determine the necessary competence for personnel performing work affecting product quality; b) provide training or take other actions to achieve or maintain the necessary competence; c) evaluate the effectiveness of the actions taken; d) ensure that its personnel are aware of the relevance and importance of their activities and how they contribute to the achievement of the quality objectives; e) maintain appropriate records of education, training, skills and experience (see 4.2.4). NOTE The methodology used to check effectiveness is proportionate to the risk associated with the work for which the training or other action is being provided. | Special training may be necessary for personnel dealing with certain categories of substances/materials such as chemicals, biological, radiation emitting or energy source components or products, and those requiring cold chain management.
6.2 Training Training for personnel should be tailored to the person’s roles and responsibilities. Typical training and education should cover: − nature of business; − functions of the personnel; − relevant policies, procedures and instructions; and − health, safety, environmental and any other applicable regulations.
Effectiveness of the training should be evaluated to ensure competency. Evaluation may consist of assessing the trainee’s feedback on the training, evaluating the work performed by the trainee after training, etc.
Records should be maintained to show the competencies that a personpossesses. Examples of such records are results of training assessments andcertifications.
6.3 Infrastructure | 7. PREMISES AND FACILITIES
The organization shall document the requirements for the infrastructure needed to achieve conformity to product requirements, prevent product mix-up and ensure orderly handling of product.Infrastructure includes, as appropriate:
a) buildings, workspace and associated utilities;
b) process equipment (both hardware and software);
c) supporting services (such as transport, communication, or information systems).
The organization shall document requirements for the maintenance activities, including the interval of performing the maintenance activities, when such maintenance activities, or lack thereof, can affect product quality. As appropriate, the requirements shall apply to equipment used in production, the control of the work environment and monitoring and measurement.
Records of such maintenance shall be maintained (see 4.2.5).
6.4 Work environment and contamination control
6.4.1 Work environment | 7.1.3 Cleanliness
The organization shall document the requirements for the work environment needed to achieve conformity to product requirements. | Storage areas should be kept clean and waste should be removed at regular intervals. Medical devices should be stored off the ground and suitably spaced to permit cleaning and inspection. Pallets should be well maintained and kept in a good state of cleanliness. Adequate precautions should be taken against spillage or breakage, contamination and cross-contamination.
If the conditions for the work environment can have an adverse effect on product quality, the organization shall document the requirements for the work environment and the procedures to monitor and control the work environment.
The organization shall: | 7.1.4 Pest control
a) document requirements for health, cleanliness and clothing of personnel if contact between such personnel and the product or work environment could affect medical device safety or performance; | Storage areas should be set up to prevent the entry of insects, rodents and other pests/animals. Examples of records to retain include service reports issued by third party pest control companies, records of in-house bait replacement, pest type and count.
b) ensure that all personnel who are required to work temporarily under special environmental conditions within the work environment are competent or supervised by a competent person.
NOTE Further information can be found in ISO 14644 and ISO 14698.
6.4.2 Contamination control
As appropriate, the organization shall plan and document arrangements for the control of contaminated or potentially contaminated product in order to prevent contamination of the work environment, personnel, or product.
For sterile medical devices, the organization shall document requirements for control of contamination with microorganisms or particulate matter and maintain the required cleanliness during assembly or packaging processes.
/ | 7.1 Storage, warehousing and stock handling
7.1.1 Receipt of stock
The receiving areas should protect deliveries from damage by the external environment (e.g. rain) during unloading. The receiving area should be separated from the storage area.
/ | 7.1.2 Storage conditions
Premises used for storage, warehousing and stock handling should have sufficient security to prevent unauthorised access and misappropriation of the medical devices. Premises should protect medical devices from contamination and deterioration, including protection from excessive heat or undue exposure to direct sunlight. The storage areas should have adequate lighting and ventilation. Smoking, eating and drinking should not be permitted.
There should be adequate storage areas, and where applicable, physically separated zones for the orderly segregation of medical devices. There should be designated areas for quarantined, saleable stock, expired, rejected/damaged, recalled and returned medical devices. Alternative means of segregation may be considered if proven, to prevent mix-up.
R1.2 ► The storage temperature as appropriate for the medical device should be based on the manufacturer’s recommendation. Where there are special storage requirements, these are typically indicated on the labels of the medical device. ◄
Temperature of the storage areas should be measured at suitable predetermined intervals to show the maximum and minimum temperatures for the day and recorded. Where applicable, humidity measurements should also be performed. Temperature monitors should be located in areas that are most likely to show fluctuations.
Medical devices requiring special storage conditions (i.e. storage conditions of 8°C and below) would require cold chain management. For more information on cold chain management of medical devices, please refer to Annex C.
Appropriate and suitable storage conditions should be provided for hazardous, sensitive and dangerous materials such as combustible liquids and solids, pressurised gases, highly toxic and radioactive substances. As there may be other applicable laws and their regulations, please check with the relevant controlling agencies to comply with the specific requirement.
/ | 7.1.5 Stock rotation
All records of incoming and outgoing medical devices according to batch/lot/serial number or equivalent should be maintained
Periodic stock reconciliation of medical devices should be performed by comparing the physical stock quantity against the recorded stock quantity. All significant stock discrepancies should be subjected to investigation to check for inadvertent mix-ups and supply errors.
The expiry of the projected useful life of medical devices should be monitored to ensure that they are only supplied within their projected useful lives. Medical devices should not be received or supplied near or after the expiry date.
A system should be in place to ensure medical devices due for expiry first are sold and/or distributed first (First-Expiry-First Out, FEFO). For cases where no expiry date exists for the medical devices, FIFO (First-In-First-Out) should be applied.
All labels and containers of medical devices should not be altered, tampered or changed. Medical devices with broken seals, damaged packaging or suspected tampering/contamination should not be sold or supplied.
To prevent any inadvertent mix-up with usable stock, damaged or expired medical devices should be clearly segregated, labelled as not for use, and disposed accordingly to documented procedure.
/ | 7.2 Delivery of medical devices
Procedures should be in place to achieve safe and secure delivery of all medical devices from their point of release to their point of delivery.
Due consideration should be given to ensure adequate security and proper storage conditions during transportation of the medical devices. The transport process should not affect the integrity and quality of the medical devices. Vehicles should not be used as a storage facility for medical devices.
Medical devices should be supplied to the designated party only. Deliveries should be made only to licensed wholesalers or distributors / persons authorised to supply medical devices.
For more information on delivery of medical devices requiring cold chain management, please refer to Annex C.
7.2 Customer-related processes
7.2.1 Determination of requirements related to product
The organization shall determine:
a) requirements specified by the customer, including the requirements for delivery and post-delivery activities;
b) requirements not stated by the customer but necessary for specified or intended use, as known;
c) applicable regulatory requirements related to the product;
d) any user training needed to ensure specified performance and safe use of the medical device;
e) any additional requirements determined by the organization.
7.2.2 Review of requirements related to product
The organization shall review the requirements related to product. This review shall be conducted prior to the organization’s commitment to supply product to the customer (e.g. submission of tenders, acceptance of contracts or orders, acceptance of changes to contracts or orders) and shall ensure that:
a) product requirements are defined and documented;
b) contract or order requirements differing from those previously expressed are resolved;
c) applicable regulatory requirements are met;
d) any user training identified in accordance with 7.2.1 is available or planned to be available;
e) the organization has the ability to meet the defined requirements.
Records of the results of the review and actions arising from the review shall be maintained (see 4.2.5).
When the customer provides no documented statement of requirement, the customer requirements shall be confirmed by the organization before acceptance.
When product requirements are changed, the organization shall ensure that relevant documents are amended and that relevant personnel are made aware of the changed requirements.
7.2.3 Communication
The organization shall plan and document arrangements for communicating with customers in relation to:
a) product information;
b) enquiries, contracts or order handling, including amendments;
c) customer feedback, including complaints;
d) advisory notices.
The organization shall communicate with regulatory authorities in accordance with applicable regulatory requirements.
7.4 Purchasing
7.4.1 Purchasing process
The organization shall document procedures (see 4.2.4) to ensure that purchased product conforms to specified purchasing information.
The organization shall establish criteria for the evaluation and selection of suppliers. The criteria shall be:
a) based on the supplier’s ability to provide product that meets the organization’s requirements;
b) based on the performance of the supplier;
c) based on the effect of the purchased product on the quality of the medical device;
d) proportionate to the risk associated with the medical device.
The organization shall plan the monitoring and re-evaluation of suppliers. Supplier performance in meeting requirements for the purchased product shall be monitored. The results of the monitoring shall provide an input into the supplier re-evaluation process.
Non-fulfilment of purchasing requirements shall be addressed with the supplier proportionate to the risk associated with the purchased product and compliance with applicable regulatory requirements.
Records of the results of evaluation, selection, monitoring and re-evaluation of supplier capability or performance and any necessary actions arising from these activities shall be maintained (see 4.2.5).
7.4.2 Purchasing information
Purchasing information shall describe or reference the product to be purchased, including as appropriate:
a) product specifications;
b) requirements for product acceptance, procedures, processes and equipment;
c) requirements for qualification of supplier personnel;
d) quality management system requirements.
The organization shall ensure the adequacy of specified purchasing requirements prior to their communication to the supplier.
Purchasing information shall include, as applicable, a written agreement that the supplier notify the organization of changes in the purchased product prior to implementation of any changes that affect the ability of the purchased product to meet specified purchase requirements.
To the extent required for traceability given in 7.5.9, the organization shall maintain relevant purchasing information in the form of documents (see 4.2.4) and records (see 4.2.5).
7.4.3 Verification of purchased product
The organization shall establish and implement the inspection or other activities necessary for ensuring that purchased product meets specified purchasing requirements. The extent of verification activities shall be based on the supplier evaluation results and proportionate to the risks associated with the purchased product.
When the organization becomes aware of any changes to the purchased product, the organization shall determine whether these changes affect the product realization process or the medical device.
When the organization or its customer intends to perform verification at the supplier’s premises, the organization shall state the intended verification activities and method of product release in the the organization shall state the intended verification activities and method of product release in the purchasing information.
Records of the verification shall be maintained (see 4.2.5).
7.5 Production and service provision | 7.3 Installation and servicing
7.5.3 Installation activities | 7.3.1 Installation activities
The organization shall document requirements for medical device installation and acceptance criteria for verification of installation, as appropriate. | Instructions and procedures for performing and verifying that the installation meets the specified requirements should be established. Only appropriately trained and qualified personnel should conduct installation of medical devices.
If the agreed customer requirements allow installation of the medical device to be performed by an external party other than the organization or its supplier, the organization shall provide documented requirements for medical device installation and verification of installation.
Records of medical device installation and verification of installation performed by the organization or its supplier shall be maintained (see 4.2.5).
7.5.4 Servicing activities | 7.3.2 Servicing activities
If servicing of the medical device is a specified requirement, the organization shall document servicing procedures, reference materials, and reference measurements, as necessary, for performing servicing activities and verifying that product requirements are met. | Instructions and procedures for performing and verifying that the servicing (including repair and maintenance) meets the specified requirements should be established. Only appropriately trained and qualified personnel should conduct servicing of medical devices.
The organization shall analyse records of servicing activities carried out by the organization or its supplier:
a) to determine if the information is to be handled as a complaint;
b) as appropriate, for input to the improvement process.
Records of servicing activities carried out by the organization or its supplier shall be maintained(see 4.2.5).
7.5.8 Identification
The organization shall document procedures for product identification and identify product by suitable means throughout product realization.
The organization shall identify product status with respect to monitoring and measurement requirements throughout product realization. Identification of product status shall be maintained throughout production, storage, installation and servicing of product to ensure that only product that has passed the required inspections and tests or released under an authorized concession is dispatched, used or installed.
If required by applicable regulatory requirements, the organization shall document a system to assign unique device identification to the medical device.
The organization shall document procedures to ensure that medical devices returned to the organization are identified and distinguished from conforming product.
/ | 8. SECONDARY ASSEMBLY | ANNEX D R1.1 ► Secondary assembly:
| Please refer to Annex D of this document for information on manufacturing activities that fall within the purview of secondary assembly. Manufacturing activities such as sterilisation do not fall within the definition of secondary assembly. | ‘‘secondary assembly’’ means enclosing the medical device(s) in its primary packaging, into another container or packaging material in which it is to be sold or supplied, without any breach of the primary packaging.
| The Quality Management System (QMS) Requirements for all other manufacturing activities other than secondary assembly is ISO 13485 Medical devices – Quality management systems – Requirements for regulatory purposes. | Any labelling information on the secondary packaging, where applicable, shall be aligned to the labelling on the primary packaging and as approved by the regulatory authority for the medical device.
| 8.1 General requirements | ‘‘primary packaging’’, in relation to a medical device, means packaging that maintains the sterility or integrity of the medical device. This maintains the wholesomeness of the medical device. Typically the appropriate packaging for a medical device is determined and validated by its manufacturer to ensure that it is effective in maintaining the sterility and wholesomeness of the medical device throughout its shelf life.
| The size of assembly area should reflect the volume of assembly. The adequacy of the working space should permit the orderly and logical positioning of equipment and materials so as to avoid confusion and to minimize the risk of mix-up between different medical devices or their components. Segregated areas should be provided for the storage of approved, quarantined, rejected, recalled and returned materials or products. | Examples of secondary assembly 1. Repackaging of single units of single type medical device into multiple unit configuration. 2. Repackaging of single units of different medical devices into multiple unit configuration. 3. Break-pack of multiple unit configuration of medical devices into smaller unit configurations.
| Assembly areas should be well-lit and effectively ventilated, with air handling system (including temperature, humidity and filtration controls) where appropriate. Smoking, eating and drinking should not be permitted in the assembly area. | Break-pack and repackaging as reduced number of medical device units or as single medical device units typically to meet local user demand for costeffectiveness or usability.
| In considering which controlled conditions are applicable for a given process, a due consideration an organisation should make is if in the absence of the control, whether there could be an adverse or potentially adverse effect on the quality of the medical device. If yes, then control is necessary. | Points to Note: • No breach of the original primary packaging of the individual medical devices and no assignment of new primary label (original primary label shall remain with individual medical devices). • All information on the label of the new secondary packaging, where applicable, shall be aligned to the labelling on the primary packaging and as approved by the regulatory authority for the medical device.
| The amount of control and level of detail should be commensurate with the degree of criticality (e.g. based on the output of risk management activities) of the process in achieving the requirements for quality and the degree of training of personnel involved in the secondary assembly line. | • No assignment of new expiry date, new product name, new product owner name or any other new information related to the identity or use of the medical device on the new repackaged format of the medical device.
| 8.2 Assembly documents | • Any expiry date indicated on the repackaged unit containing different medical devices shall reflect the shortest expiry date based on the expiry dates of the individual medical devices in it.
| Any secondary assembly work instructions should include: − name of medical device; − description of the applicable medical devices for assembly and pack size; − complete list of all the packaging materials required for a standard batch size, including quantities, sizes and types, with the code or reference number relating to the specifications of each packaging material; − where appropriate, an example or reproduction of the relevant printed packaging materials, and specimens indicating where to apply batch number references, and expiry date of the medical device; − special precautions to be observed, e.g. a careful examination of the area and equipment in order to ascertain the line clearance before operations begin; − description of the packaging operation, including any significant subsidiary operations, and equipment to be used; and − details of any in-process controls with instructions for sampling and acceptance limits, if applicable. | What is not secondary assembly:
| Records that facilitate traceability and review of the secondary assembly of a batch of medical device, derived during the secondary assembly of that batch, should be contained in a batch record, and frequently collated in a single file. Such files may be referred to as a “Device History Record”, “Batch Assembly Record”, “Lot History Record” or “Lot Record”. | 1. Repackaging of medical device units enclosed within original intact/sealed packaging into additional cartons intended for shipping and transportation purposes only.
If it is not practical to include all the relevant documents in the batch record, then the record should list the titles of those documents and their location(s).During secondary assembly, relevant information should be entered into the batch record. Such information should include: − quantity of raw materials, components and intermediate products, and their batch number, if appropriate, − date of start and completion of different stages of secondary assembly; − quantity of medical device assembled; − signed results of all inspections and tests; − designation of the product line used; and − deviation from the secondary assembly specifications, if applicable.
Data may be recorded by electronic data processing systems, photographic or other reliable means, but detailed procedures relating to the system in use should be available and the accuracy of the records should be checked.
| If documentation is handled by electronic data processing methods, only authorised personnel should be able to enter or modify data in the computer and there should be a record of changes and deletion; access should be restricted by passwords or other means and the result of entry of critical data should be independently checked. | 2. Manufacturing activities which are not deemed to be secondary assembly activities (non-exhaustive list):-
Batch records electronically stored should be protected by back-up transfer on magnetic tape, microfilm, paper or other means. It is particularly important that the records are readily available and to ensure their integrity throughout the period of retention.
| Records of batch/lot numbers and identities of all individual medical devices that are re-packaged or assembled during the secondary assembly process should be recorded and maintained to allow for traceability. Distribution records should be maintained for each batch of medical device in order to facilitate the recall of the batch if necessary. | • Assignment of new expiry date to medical device(s);
| 8.3 Materials control | • Assignment of new brand name and/or new product owner name to
| 8.3.1 Medical devices to be repacked | medical device(s);
| R1.1 ► Medical devices should only be purchased from approved suppliers and, where possible, directly from the manufacturer. A batch certificate of analysis or equivalent document with key information on the specifications of the medical device being re-packed should be obtained from the supplier for each batch of medical device to be re-packed. Inspection of breached primary packages should be performed. Any breached packages should be disposed of and should not be re-processed and returned to the secondary assemblyline. ◄ | • Assignment of new primary label to medical device(s); or
| Reference materials may be physical or visual, such as product examples indicating permissible colour variation, or visual such as photos of known nonconformities. Reference materials should be available at the point of use. An individual procedure may be in the form of a simple flowchart, or a processing sequence, combined with a checklist. | • Sterilisation of medical device(s)
8.3.2 Packaging materials
Particular attention should be paid to printed materials. They should be stored in adequately secure conditions so as to exclude unauthorised access. Cut labels and other loose printed materials should be stored and transported in separate closed containers so as to avoid mix-ups.
Packaging materials should be issued for used only by authorised personnel following an approved and documented procedure. Each delivery or batch of printed or packaging materials should be given a specific reference number or identification mark.
Outdated or obsolete packaging material or printed packaging material should be destroyed and the disposal should be recorded.
8.3.3 Medical device labelling
The appropriate marketing authorisation of the medical device to be assembled and the labelling and other requirements imposed by the regulatory authority on these products as part of the marketing authorisation, should be complied with, and where necessary, the approval of the marketing authorisation holder and/or regulatory authority should be sought.
The risk of labelling and packaging errors may be minimised by the introduction of appropriate controls such as: − segregation of secondary assembly (e.g. packaging and labelling) operations from other operations; − avoidance of packaging and labelling products of similar appearance in close proximity; − secondary assembly line identification; − application of line clearance procedures; − destruction of unused batch-coded materials on completion of packaging and labelling; − use of roll-feed labels; − use of a known number of labels and reconciliation of usage; − on-line printing, including batch coding; − use of electronic code encoders/readers and label counters; − use of labels designed to provide clear product differentiation; − inspection of label content before use; and − proper storage of labels in areas of restricted access.
8.4 Good assembly practices
8.4.1 Special considerations
On-line control of the product during packaging should include checking of the following: − general appearance of packages; − completeness of packages; − correctness of products and packaging materials used; and − correctness of over-printing, e.g. batch number, shelf life/expiry date.
Upon completion of a packaging operation, any unused batch-coded packaging materials should be destroyed and the destruction recorded. Any significant or unusual discrepancy observed during reconciliation of the amount of bulk product and printed packaging materials and the number of units produced should be investigated and satisfactorily accounted for before release. A documented procedure should be followed if uncoded printed materials are returned to stock.
Where any packaging process validation applies, the study should be carried out in accordance with pre-defined and authorised protocols.
8.4.2 Assembly equipment
Suitable equipment should be designed and selected so that process and product specifications are met. It should be verified that new and/or significantly modified equipment meets purchasing/design specifications and is capable of operating within its defined limits and process operating limits.
Before any assembly begins, line clearance should be carried out. There should be recorded checks that the equipment and work station are clear of previous products, documents or materials not required for the planned process, and that equipment is clean and suitable for use.
Each batch of medical device produced should be assigned a unique batch number. Normally, packing and sealing should be followed as quickly as possible by labelling. If it is not the case, appropriate procedures should be applied to ensure that no mix-ups or mislabelling can occur.
The performance of the printing operation should be checked and recorded. Attention should be paid to printing by hand which should be re-checked at regular intervals. Printed and embossed information on packaging materials should be distinct and resistant to fading or erasing.
8.5 Product release
There should be authorised and independent personnel in charge of product release. If sampling of the final finished pack is performed, the samples taken should be done in accordance with approved written procedures that describe: − method of sampling; − equipment to be used; − quantity of samples to be taken; − special storage condition for the samples taken, if applicable; and − recording of the results of in-process control.
7.5.9 Traceability | 9. TRACEABILITY
7.5.9.1 General | Traceability of medical devices in the distribution chain is important as it is necessary to track medical devices to the customer (e.g. clinics or hospitals) in the event of a field safety corrective action (FSCA). Traceability allows for investigation of quality-related problems and identification of other possible nonconforming medical devices. Identification of medical device by batch/lot/serial number permits end-to-end traceability, i.e. upstream to the manufacturer and downstream to the customers/disposal.
The organization shall document procedures for traceability. These procedures shall define the extent | Records providing traceability may include delivery orders from the supplier, distribution or sales records, returns and disposal records, etc. To facilitate traceability, the records should include the date, lot/batch/serial number, quantity, name of medical device, address of the supplier and customer. For more information on distribution records, please refer to GN-06: Guidance on Distribution Records for Medical Devices which is published on the HSA website.
of traceability in accordance with applicable regulatory requirements and the records to be maintained(see 4.2.5).
7.5.9.2 Particular requirements for implantable medical devices
The records required for traceability shall include records of components, materials, and conditions for the work environment used, if these could cause the medical device not to satisfy its specified safety and performance requirements.
The organization shall require that suppliers of distribution services or distributors maintain records of the distribution of medical devices to allow traceability and that these records are available for inspection.
Records of the name and address of the shipping package consignee shall be maintained (see 4.2.5).
7.5.10 Customer property
The organization shall identify, verify, protect, and safeguard customer property provided for use or incorporation into the product while it is under the organization’s control or being used by the organization. If any customer property is lost, damaged or otherwise found to be unsuitable for use, the organization shall report this to the customer and maintain records (see 4.2.5).
7.5.11 Preservation of product
The organization shall document procedures for preserving the conformity of product to requirements during processing, storage, handling, and distribution. Preservation shall apply to the constituent parts of a medical device.
The organization shall protect product from alteration, contamination or damage when exposed to expected conditions and hazards during processing, storage, handling, and distribution by:
a) designing and constructing suitable packaging and shipping containers;
b) documenting requirements for special conditions needed if packaging alone cannot provide preservation.
If special conditions are required, they shall be controlled and recorded (see 4.2.5).
7.6 Control of monitoring and measuring equipment | 7.4 Calibration
The organization shall determine the monitoring and measurement to be undertaken and the monitoring and measuring equipment needed to provide evidence of conformity of product to determined requirements. | The instruments used for measuring and monitoring (e.g. temperature, humidity, and electrical current) should be calibrated or verified for accuracy regularly and the results of such calibrations or verifications should be recorded and retained. Calibration should be performed by competent personnel or party.
The organization shall document procedures to ensure that monitoring and measurement can be carried out and are carried out in a manner that is consistent with the monitoring and measurement requirements.
As necessary to ensure valid results, measuring equipment shall:
a) be calibrated or verified, or both, at specified intervals, or prior to use, against measurement standards traceable to international or national measurement standards: when no such standards exist, the basis used for calibration or verification shall be recorded (see 4.2.5);
b) be adjusted or re-adjusted as necessary: such adjustments or re-adjustments shall be recorded(see 4.2.5);
c) have identification in order to determine its calibration status;
d) be safeguarded from adjustments that would invalidate the measurement result;
e) be protected from damage and deterioration during handling, maintenance and storage.
The organization shall perform calibration or verification in accordance with documented procedures.
In addition, the organization shall assess and record the validity of the previous measuring results when the equipment is found not to conform to requirements. The organization shall take appropriate action in regard to the equipment and any product affected.
Records of the results of calibration and verification shall be maintained (see 4.2.5).
The organization shall document procedures for the validation of the application of computer software used for the monitoring and measurement of requirements. Such software applications shall be validated prior to initial use and, as appropriate, after changes to such software or its application.
The specific approach and activities associated with software validation and revalidation shall be proportionate to the risk associated with the use of the software, including the effect on the ability of the product to conform to specifications.
Records of the results and conclusion of validation and necessary actions from the validation shall be maintained (see 4.2.4 and 4.2.5).
NOTE Further information can be found in ISO 10012.
/ | 10. COUNTERFEIT, ADULTERATED, UNWHOLESOME OR TAMPERED MEDICAL DEVICES
Reporting should be performed using the Medical Device Post-Market Information Report Form once the organisation becomes aware of possible counterfeit, adulterated, unwholesome or tampered medical devices. Key information (i.e. brand, place and period of purchase) should be provided in the form, to assist in the detection of counterfeit medical devices supplied in Singapore.
Appropriate measures should be undertaken for identified counterfeit, adulterated, unwholesome or tampered medical devices, which include, but are not limited to the following: − reporting to the regulatory authority; − segregation/quarantine of the affected medical devices, labelling them “Not for Use” or with other similar phrases/words; − investigation of supply chain breach; and − communication to all affected wholesalers / distributors / retailers / consignees possibly supplied with the affected medical devices (prior notification to the regulatory authority is required for such communication)
8 Measurement, analysis and improvement
8.1 General
The organization shall plan and implement the monitoring, measurement, analysis and improvement processes needed to:
a) demonstrate conformity of product;
b) ensure conformity of the quality management system;
c) maintain the effectiveness of the quality management system.
This shall include determination of appropriate methods, including statistical techniques, and the extent of their use.
8.2 Monitoring and measurement
8.2.1 Feedback
As one of the measurements of the effectiveness of the quality management system, the organization shall gather and monitor information relating to whether the organization has met customer requirements. The methods for obtaining and using this information shall be documented.
The organization shall document procedures for the feedback process. This feedback process shall include provisions to gather data from production as well as post-production activities.
The information gathered in the feedback process shall serve as potential input into risk management for monitoring and maintaining the product requirements as well as the product realization or improvement processes.
If applicable regulatory requirements require the organization to gain specific experience from postproduction activities, the review of this experience shall form part of the feedback process.
8.2.2 Complaint handling | 11. COMPLAINT HANDLING
The organization shall document procedures for timely complaint handling in accordance with applicable regulatory requirements. | A procedure describing the handling of all written and oral complaints regarding a medical device should be available. The procedure for handling complaints should ensure that the complaints received are investigated, followed through, and that corrective and preventive actions are taken. For more information on complaint handling, please refer to the GN-07: Guidance on Complaint Handling of Medical Devices which is published on the HSA website.
These procedures shall include at a minimum requirements and responsibilities for: a) receiving and recording information; b) evaluating information to determine if the feedback constitutes a complaint; c) investigating complaints; d) determining the need to report the information to the appropriate regulatory authorities; e) handling of complaint-related product; f) determining the need to initiate corrections or corrective actions. | A person within the company should be authorised to handle complaints and initiate investigations. All investigations should be documented in writing. Complaint investigation should take into consideration the condition and circumstances under which the medical device was distributed, stored and used. An investigation report should be prepared.
If any complaint is not investigated, justification shall be documented. Any correction or corrective action resulting from the complaint handling process shall be documented. | All applicable corrective and preventive actions should be clearly stated. If a medical device defect is detected or suspected, a review of all potentially affected batches should be considered. Where a decision is made to recall the medical device, the details of the recall should be documented. For more information on medical device recalls, please refer to the GN-04: Guidance on Medical Device Recall which is published on the HSA website. When the complaint is an adverse event related to a medical device, the regulatory authority must be notified promptly in accordance to GN-05: Guidance on Reporting of Adverse Events for Medical Devices which is published on the HSA website.
If an investigation determines activities outside the organization contributed to the complaint, relevant information shall be exchanged between the organization and the external party involved. | Individual complaint records should be maintained. Complaints records should be trended and reviewed regularly for any indication of specific or recurring problems requiring attention.
Complaint handling records shall be maintained (see 4.2.5).
8.2.3 Reporting to regulatory authorities
If applicable regulatory requirements require notification of complaints that meet specified reporting criteria of adverse events or issuance of advisory notices, the organization shall document procedures for providing notification to the appropriate regulatory authorities.
Records of reporting to regulatory authorities shall be maintained (see 4.2.5).
8.2.4 Internal audit | 13. INTERNAL AUDIT
The organization shall conduct internal audits at planned intervals to determine whether the quality management system: | Internal audits should be conducted on a regular basis, at least once a year, and recorded in order to monitor the implementation and conformance with the requirements of the SS GDPMDS.
a) conforms to planned and documented arrangements, requirements of this International Standard, quality management system requirements established by the organization, and applicable regulatory requirements; | The results of the audits are usually stated in a written report indicating the nonconformities found. Timely action should be taken to correct/eliminate the nonconformities and their root causes. The audit results and the status of the nonconformities should be communicated to management for review and be monitored till effective closure is achieved.
b) is effectively implemented and maintained. | The organisation should identify and implement any changes necessary to ensure and maintain the continued suitability and effectiveness of the quality management system
The organization shall document a procedure to describe the responsibilities and requirements for planning and conducting audits and recording and reporting audit results.
An audit program shall be planned, taking into consideration the status and importance of the processes and area to be audited, as well as the results of previous audits. The audit criteria, scope, interval and methods shall be defined and recorded (see 4.2.5). The selection of auditors and conduct of audits shall ensure objectivity and impartiality of the audit process. Auditors shall not audit their own work.
Records of the audits and their results, including identification of the processes and areas audited and the conclusions, shall be maintained (see 4.2.5).
The management responsible for the area being audited shall ensure that any necessary corrections and corrective actions are taken without undue delay to eliminate detected nonconformities and their causes. Follow-up activities shall include the verification of the actions taken and the reporting of verification results.
NOTE Further information can be found in ISO 19011.
8.2.5 Monitoring and measurement of processes
The organization shall apply suitable methods for monitoring and, as appropriate, measurement of the quality management system processes. These methods shall demonstrate the ability of the processes to achieve planned results. When planned results are not achieved, correction and corrective action shall be taken, as appropriate.
8.2.6 Monitoring and measurement of product
The organization shall monitor and measure the characteristics of the product to verify that product requirements have been met. This shall be carried out at applicable stages of the product realization process in accordance with the planned and documented arrangements and documented procedures.
Evidence of conformity to the acceptance criteria shall be maintained. The identity of the person authorizing release of product shall be recorded (see 4.2.5). As appropriate, records shall identify the test equipment used to perform measurement activities.
Product release and service delivery shall not proceed until the planned and documented arrangements have been satisfactorily completed.
For implantable medical devices, the organization shall record the identity of personnel performing any inspection or testing.
/ | 7.5 Return of medical devices
Any product returned to the organisation should be treated as a nonconforming product until the assessment of disposition has been determined.
The returned medical devices should only be formally released as saleable stock, following a satisfactory quality re-evaluation by authorised personnel.
This assessment should take into account the nature of the medical device, any special storage conditions required, and the time that has elapsed since it was distributed. Advice should be sought from the product owner as necessary. In addition, the following should be considered: − medical devices are in their original unopened containers and in good condition; − medical devices have been stored and handled under proper conditions; and − remaining projected useful life of the medical devices is acceptable to the customer.
Products that have been deemed suitable for release to saleable stock should be subject to the stock rotation system.
8.3 Control of nonconforming product
8.3.1 General | 7.6 Disposal of medical devices
The organization shall ensure that product which does not conform to product requirements is identified and controlled to prevent its unintended use or delivery. The organization shall document a procedure to define the controls and related responsibilities and authorities for the identification, documentation, segregation, evaluation and disposition of nonconforming product. | Medical devices that no longer meet specifications (for example, expired, contaminated, damaged, defective, etc.) should be disposed of properly.
The evaluation of nonconformity shall include a determination of the need for an investigation and notification of any external party responsible for the nonconformity. | Control should be established to ensure that the medical devices meant for disposal are not made available for re-supply and they are disposed off in accordance with the procedure and in line with the manufacturer’s instructions, if any.
Records of the nature of the nonconformities and any subsequent action taken, including the evaluation, any investigation and the rationale for decisions shall be maintained (see 4.2.5) | Methods of disposal may include but not limited to incineration, return to product owner, etc.
8.3.2 Actions in response to nonconforming product detected before delivery
The organization shall deal with nonconforming product by one or more of the following ways:
a) taking action to eliminate the detected nonconformity;
b) taking action to preclude its original intended use or application;
c) authorizing its use, release or acceptance under concession.
The organization shall ensure that nonconforming product is accepted by concession only if the justification is provided, approval is obtained and applicable regulatory requirements are met. Records of the acceptance by concession and the identity of the person authorizing the concession shall be maintained (see 4.2.5).
8.3.3 Actions in response to nonconforming product detected after delivery | 12. FIELD SAFETY CORRECTIVE ACTION (FSCA)
When nonconforming product is detected after delivery or use has started, the organization shall take action appropriate to the effects, or potential effects, of the nonconformity. Records of actions taken shall be maintained (see 4.2.5). | Once the product owner initiates the FSCA and affected medical devices have been supplied in Singapore, the implementation of the FSCA should not be delayed.
The organization shall document procedures for issuing advisory notices in accordance with applicable regulatory requirements. These procedures shall be capable of being put into effect at any time. Records of actions relating to the issuance of advisory notices shall be maintained (see 4.2.5). | Prior to the initiation of the FSCA in Singapore, it is a requirement that the FSCA is notified to the regulatory authority. After the notification, the dealer may proceed to initiate the FSCA or conduct stock recovery unless otherwise instructed by the regulatory authority. Dissemination of the product owner’s Field Safety Notice (FSN) and implementation of the corrective actions listed in the FSN are some measures (but not limited to) that constitute initiation of the FSCA.
Where the FSCA affects a particular batch, consideration should also be given to determine whether other batches/materials are also affected.
Upon completion of each FSCA, a final report should be provided to the regulatory authority. Reconciliation should be made between delivered and recovered quantities of medical devices. More information on FSCA is available in GN-10: Guidance on Medical Device Field Safety Corrective Action which is published on the HSA website.
8.4 Analysis of data
The organization shall document procedures to determine, collect and analyse appropriate data to demonstrate the suitability, adequacy and effectiveness of the quality management system. The procedures shall include determination of appropriate methods, including statistical techniques and the extent of their use.
The analysis of data shall include data generated as a result of monitoring and measurement and from other relevant sources and include, at a minimum, input from:
a) feedback;
b) conformity to product requirements;
c) characteristics and trends of processes and product, including opportunities for improvement;
d) suppliers;
e) audits;
f) service reports, as appropriate.
If the analysis of data shows that the quality management system is not suitable, adequate or effective, the organization shall use this analysis as input for improvement as required in 8.5.
Records of the results of analyses shall be maintained (see 4.2.5).
8.5 Improvement
8.5.1 General
The organization shall identify and implement any changes necessary to ensure and maintain the continued suitability, adequacy and effectiveness of the quality management system as well as medical device safety and performance through the use of the quality policy, quality objectives, audit results, postmarket surveillance, analysis of data, corrective actions, preventive actions and management review.
8.5.2 Corrective action
The organization shall take action to eliminate the cause of nonconformities in order to prevent recurrence. Any necessary corrective actions shall be taken without undue delay. Corrective actions shall be proportionate to the effects of the nonconformities encountered.
The organization shall document a procedure to define requirements for:
a) reviewing nonconformities (including complaints);
b) determining the causes of nonconformities;
c) evaluating the need for action to ensure that nonconformities do not recur;
d) planning and documenting action needed and implementing such action, including, as appropriate,
updating documentation;
e) verifying that the corrective action does not adversely affect the ability to meet applicable
regulatory requirements or the safety and performance of the medical device;
f) reviewing the effectiveness of corrective action taken.
Records of the results of any investigation and of action taken shall be maintained (see 4.2.5).
8.5.3 Preventive action
The organization shall determine action to eliminate the causes of potential nonconformities in order to prevent their occurrence. Preventive actions shall be proportionate to the effects of the potential problems.
The organization shall document a procedure to describe requirements for:
a) determining potential nonconformities and their causes;
b) evaluating the need for action to prevent occurrence of nonconformities;
c) planning and documenting action needed and implementing such action, including, as appropriate,
updating documentation;
d) verifying that the action does not adversely affect the ability to meet applicable regulatory
requirements or the safety and performance of the medical device;
e) reviewing the effectiveness of the preventive action taken, as appropriate.
Records of the results of any investigations and of action taken shall be maintained (see 4.2.5).
/ | ANNEX A |  | ANNEX C
| The scope of SS GDPMDS certification shall specify the following: |  | Cold Chain Management
| (A) Activities performed + categories of medical devices handled by the organisation |  | (1) Introduction
| • Applicable activities for organisation include: - Import - Storage - Distribution - Installation - Servicing - Secondary assembly • Refer to Annex B for categories of medical devices. |  | This annex is meant to provide guidance on the cold chain management of medical devices in relation to the SS GDPMDS.
| (B) Activities that are outsourced The name and address(es) of outsourced service providers for storage,distribution and secondary assembly. |  | Cold chain medical devices refer to those which require special storage and handling conditions of 8°C and below (e.g. refrigeration at 2°C to 8°C, deep freeze at below -15°C). The storage condition requirements are usually described on the product label.
| (C) Any special storage and handling conditions |  | (2) Documentation requirements
| For organisations with facilities to handle medical devices requiring cold chain management (storage conditions of 8°C and below), the following statement shall be provided: • There are special storage and handling conditions. |  | Procedures for receiving, storage and delivery should include instructions on the special storage and handling of medical devices requiring cold chain management.
| For organisations without facilities to handle medical devices requiring cold chain management, the following statement shall be provided: • There are no special storage and handling conditions. |  | There should be a written procedure for investigation and handling temperature excursions that may occur during receiving, storage and delivery. All records should be readily retrievable to enable early detection of temperature excursions. Temperature excursions or any possible product damage should be reported to the distributor and recipient of the affected medical devices.
(3) Personnel and Training
| ANNEX B |  | Specific training in cold chain management should be provided to all personnel, including delivery drivers, who are involved in operations that may affect the quality of cold chain medical devices. The training procedure should describe the training program required for personnel involved with cold chain management, and training records for relevant personnel should be retained.
| Medical device categories for inclusion in scope of certification |  | (4) Receipt of stock
| The certification is specific for categories of medical devices that are covered under the audit. Some replication may exist when it comes to classifying the medical device. For example, a medical device may be classified as dental device and single use device. Therefore, the category selected should be the one that is most appropriate for that medical device. |  | a) Immediate identification of cold chain medical devices: • A list of cold chain medical devices handled by the company should be available for reference by the personnel involved in receiving goods and other relevant store personnel. • Cold chain medical devices should be prioritised and transferred to the appropriate storage facilities immediately upon receipt.
| Name | Definition | b) Appropriate handling of cold chain medical devices: • Products transported in insulated boxes with cool-packs should be unpacked before storage in the cold room as internal temperature of such insulated boxes can drop below its minimum temperature allowable if kept in the cold room for extended periods. • Checks on incoming goods for signs of tampering, damage and non-compliance with cold chain storage condition, as well as other relevant information against the purchase order should be carried out under the storage conditions as recommended on the product label unless otherwise justified. • If temperature loggers are used to monitor the temperature of the medical devices during transportation, the temperature profile provided by the temperature loggers should be kept as records to show that the medical devices were transported under the required storage conditions. Any excursions should be investigated and the impact on the quality of the medical devices assessed. • Records of all checks performed should be kept.
| 01 Active implantable devices | Devices that operate with an integral power source (i.e., independent of energy from the human body or gravity), that are totally or partially introduced, surgically or medically, into the human body or body orifice, where they are intended to remain temporarily or permanently. Examples of devices in this category include cochlear implants, implantable defibrillators, implantable infusion pumps, implantable stimulators, pacemakers, and their accessories. | (5) Storage conditions
| 02 Anaesthetic and respiratory devices | Devices used to supply, condition, monitor, dispense, or deliver respiratory or anaesthetic gases, vapours or other substances to provide and/or control respiration and/or anaesthesia. Examples of devices in this category include airways, anaesthesia systems, breathing circuits, humidifiers, tracheal tubes, ventilators, and their accessories. | a) Storage facilities: Storage facilities for cold chain management include refrigerator, freezer and walk-in cold room. The facility should be able to provide adequate temperature controls for all areas used for storing products, as determined by a temperature mapping exercise. The storage conditions should be monitored and recorded continuously and installed with an alarm system to alert the relevant personnel to any temperature excursions. A maintenance program should be in place to ensure the functioning of the facility and in case of power failure, backup power supply or alternative storage solutions should provide means to maintain the storage conditions.
| 03 Dental devices | Devices used to diagnose, prevent, monitor, treat, or alleviate oral, maxillo-facial, and dental disease/disorders. Examples of devices in this category include dental amalgam, dental cements, dental hand instruments, dental implants, dental materials, dental tools/laboratory devices, and their accessories. | b) Temperature mapping: An initial temperature mapping exercise should be carried out on the facility under representative conditions. The mapping exercise should be repeated for significant changes according to the results of a risk assessment exercise. Temperature monitoring and recording instruments should be placed in various locations within the facility and monitored for a representative period of time, usually under the worstcase scenarios, e.g., areas that experience extreme fluctuations in temperature or the hottest time of the year. The results should be analysed to determine if all parts of the facility are able to meet the storage requirements and to identify the hottest and coldest spots, and other suitable locations for routine temperature monitoring.
| 04 Electro mechanical medical devices | Devices that operate on electrical energy (electromedical) and/or through some integrated physical mechanism or machinery (mechanical). Examples of devices in this category include specialized beds, defibrillators, dialysis systems, electrocardiographs (ECG), electroencephalographs (EEG), endoscopes, infusion pumps, lasers, operation/examination tables/lights, suction systems, and their accessories. | c) Temperature monitoring and recording: Acceptable temperature ranges for the storage conditions should be established based on the storage requirements of the medical devices. The routine monitoring should be taken from suitable locations identified from the temperature mapping exercise, unless otherwise justified. The temperature conditions of the storage facility should be monitored and recorded continuously using temperature monitoring and recording instruments. As far as possible, the actual storage temperature recorded should be expressed quantitatively. Temperature monitoring records should be checked and independently reviewed regularly.
| 05 Hospital hardware | Treatment-related devices that typically are not directly or actively involved in the diagnosis or treatment of patients, but that support or facilitate such activities. Examples of devices in this category include disinfectants for medical devices, patient beds, patient transfer equipment, sterilisers, and their accessories. | d) Alarm system: The storage facility should be installed with an alarm system to notify relevant personnel of any temperature excursions so that appropriate corrective or preventive actions can be initiated in a timely manner. Action and alert limits should be established based on the pre-defined storage conditions. The correct function of the alarm system should be subject to periodic testing.
| 06 In vitro diagnostic devices | Devices used to examine clinical samples taken from the human body to evaluate physiological or pathological conditions. Examples of devices in this category include analysers, blood glucose monitoring devices, in vitro diagnostic (IVD) test kits/calibrators/controls, dedicated laboratory equipment, microbial sensitivity systems, and their accessories. | e) Backup power: Backup power should be available for the storage facility to ensure that storage temperature conditions can be maintained and temperature monitoring and recording instruments will continue to function in the event of a power failure. Any backup generators used should be subject to periodic testing. Alternative backup plans that provide equivalent storage temperature conditions and monitoring system may be considered in the absence of backup generator.
| 07 Non-active implantable devices | Devices without an integral power source that are totally or partially introduced, surgically or medically, into the human body or body orifice, where they are intended to remain for longer than 30 days. Examples of devices in this category include cardiovascular clips, embolization implants, orthopaedic fixation systems, intrauterine devices, heart valves, bone prostheses, and their accessories. | f) Maintenance program: A regular maintenance programme should be established and carried out for the airconditioning system of the storage facility. For storage facilities with auto-defrost function, precautionary steps should be taken to ensure that the storage temperature is not affected during the defrost cycles.
| 08 Ophthalmic and optical devices | Devices used to diagnose, prevent, monitor, treat, correct, or alleviate diseases or disorders related to the eye. Examples of devices in this category include contact lenses, keratomes, intraocular lenses, slit lamps, ophthalmic test instruments, phacoemulsification systems, tonometers, and their accessories. | (6) Delivery of medical devices
| 09 Reusable devices | Devices that can be used for more than one application period, often involving cleaning and/or sterilisation between the periods (excluding capital equipment). Examples of devices in this category include drills, elastic bandages, haemostats, medicine administration kits, saws, scar management garments, reusable surgical instruments (chisels, scissors, retractors, scalpels), and their accessories. | Medical devices requiring special storage conditions should be transported by appropriate and suitable means to ensure that the required temperature conditions are maintained throughout delivery. Transportation providers should be informed of the special storage requirements for cold chain medical devices. The planning of the delivery route should be taken into consideration in the procedure for the delivery of cold chain medical devices. Where the transportation route includes unloading and reloading or transit storage at a transportation hub, particular attention should be paid to temperature monitoring, cleanliness and the security of any intermediate storage facilities.
| 10 Single-use devices # | Devices intended to be used only once, or for only one patient during one medical procedure or short term, and then discarded if not already rapidly absorbed. Examples of devices in this category include adhesive tapes, bandages, blood collection devices, catheters, condoms, dressings, electrodes, kits/sets (biopsy, intravenous infusion), needles, single use surgical instruments/products (cannulae, scalpels, absorbents), and disposable bedding. | a) Thermal packaging and insulated boxes: • Insulated boxes should be accompanied with temperature loggers which can monitor and record the temperature within each box throughout the delivery. Alternatively, if temperature loggers are not used, the delivery conditions using insulated boxes should be validated by simulating the possible worst-case scenarios during transportation to show that the desired temperature can be maintained under those conditions. • Based on the validation study, there should be written procedures to describe the packing materials, packing configuration of the insulated boxes, labelling on the product that clearly identifies the required special storage/transport conditions and acceptable delivery time required for the cold chain medical devices. • The packing operation for the cold chain medical devices should be recorded and there should be an independent check to ensure that the packing operation is carried out in accordance with the written procedure. The records for check should be kept. • If cool-packs are used in insulated boxes, they should be located such that the product does not come into direct contact with them. There should be a system in place to control the reuse of cool packs to ensure that incompletely cooled packs are not used in error. There should be adequate physical segregation between frozen and chilled cool-packs. • Special care should be exercised when using dry ice during transportation. Products should not come into direct contact with dry ice as this may cause freezing of the product.
| 11 Assistive products for persons with disability | Devices specially produced or adapted which compensate for, relieve, prevent, or neutralize an impairment, disability, or handicap. Examples of devices in this category include artificial limbs, audiometers, crutches, hearing aids, patient lifts, wheelchairs, and their accessories. | b) Temperature-controlled vehicles or transportation containers: • If refrigerated vehicles or transportation containers provide the primary means for environmental control, temperature mapping under representative conditions should be carried out. The storage conditions within the refrigerated vehicles or transportation containers used for cold chain medical devices should be monitored and recorded with calibrated temperature monitoring and recording instruments.
| 12 Diagnostic and therapeutic radiation devices | Devices that use radiation energy including in vivo isotopes, excited particle energy, magnetic resonance imaging, nuclear energy, ultrasound, and x-ray for the purpose of providing diagnostic imaging and/or therapeutic radiation treatment. Examples of devices in this category include accelerator systems, bone absorptiometric systems, computed tomography (CT) systems, magnetic resonance imaging (MRI) systems, positron emission tomography (PET) systems, X-ray systems, and their accessories. Radiant warming and phototherapy devices are excluded. | (7) Calibration
| 13 Medical software | Computer programs and related data designed for use in a parent medical device, or as a stand-alone medical device, intended to provide a variety of functions such as: to help drive and/or influence a process or procedure, manage [data manipulation (e.g., calculations)], assist in analysis of data. It is typically supplied pre-installed in a parent device, or as an update, or is installed on an electronic medium for installation into an existing computer(s)/network or Internet-based software. Examples of devices in this category include application program software, diagnostic digital imaging systems, IVD software, and web-based software. | Key equipment in cold chain management include cold rooms, air handling units, alarm system, refrigerators, temperature monitoring and recording instruments etc. Adequate records of repair, maintenance and calibration activities for key equipment should be maintained.
| # Category 10: Single-use devices is a generic category that shall be used only if all other categories do not apply to the medical device in question. |  | Temperature monitoring and recording instruments should be calibrated regularly to show that they are operating correctly. The frequency of calibration should be determined based on a risk assessment. A minimum of three-point calibration to cover the operating range and carried out annually is preferred. Certificates of calibration should be traceable to national or international standards and are reviewed and maintained.
(8) Return of medical devices
There should be procedures established to describe how return of cold chain medical devices should be handled. Return of cold chain medical devices to saleable stock is not advisable unless there is sufficient evidence to demonstrate that the cold chain and quality of the medical devices has not been compromised.
Records of evaluation should include documentation of the complete trail of the medical devices, including but not limited to the delivery to customer, opening of the package by customer, delivery back to the distributor and receipt of returned medical devices by the distributor.
(9) Secondary assembly of cold chain medical devices
The secondary assembly of cold chain medical devices should be carried out under the storage conditions as recommended on the product label unless otherwise justified by stability studies that such operations can be performed outside the recommended storage condition within an established timeframe do not compromise the quality of the medical devices.
(10) Outsourced activities
Written contract(s) should be established to describe arrangements and responsibilities between the company and the outsourced service provider with respect to contract warehousing, secondary assembly and transportation of cold chain medical devices. The contract should provide all the necessary information and define the conditions required for the contracted operations to be carried out.

## ISO 13485 & Pakistan GDPMD
ISO 13485: 2016 | Pakistan-GUIDELINES ON GOOD DISTRIBUTION PRACTICE FOR MEDICAL DEVICES (GDPMD) | Note
Introduction | 1. PRELIMINARY
0.1 General | 1.1 Objective
This International Standard specifies requirements for a quality management system that can be used by an organization involved in one or more stages of the life-cycle of a medical device, including design and development, production, storage and distribution, installation, servicing and final decommissioning and disposal of medical devices, and design and development, or provision of associated activities (e.g. technical support). The requirements in this International Standard can also be used by suppliers or other external parties providing product (e.g. raw materials, components, subassemblies, medical devices, sterilization services, calibration services, distribution services, maintenance services) to such organizations. The supplier or external party can voluntarily choose to conform to the requirements of this International Standard or can be required by contract to conform. | The objective of this document is to ensure the quality, safety and performance of medical device during all aspects of medical device supply-chain, which include, but not limited to procurement, storage, transportation, delivery, installation, commissioning, service and maintenance, calibration and after sale service, tracking, documentation and record- keeping practices.
1 SCOPE This International Standard specifies requirements for a quality management system where an organization needs to demonstrate its ability to provide medical devices and related services that consistently meet customer and applicable regulatory requirements. Such organizations can be involved in one or more stages of the life-cycle, including design and development, production, storage and distribution, installation, or servicing of a medical device and design and development or provision of associated activities (e.g. technical support). This International Standard can also be used by suppliers or external parties that provide product, including quality management system-related services to such organizations. Requirements of this International Standard are applicable to organizations regardless of their size and regardless of their type except where explicitly stated. Wherever requirements are specified as applying to medical devices, the requirements apply equally to associated services as supplied by the organization. The processes required by this International Standard that are applicable to the organization, but are not performed by the organization, are the responsibility of the organization and are accounted for in the organization’s quality management system by monitoring, maintaining, and controlling the processes. If applicable regulatory requirements permit exclusions of design and development controls, this can be used as a justification for their exclusion from the quality management system. These regulatory requirements can provide alternative approaches that are to be addressed in the quality management system. It is the responsibility of the organization to ensure that claims of conformity to this International Standard reflect any exclusion of design and development controls. If any requirement in Clauses 6, 7 or 8 of this International Standard is not applicable due to the activities undertaken by the organization or the nature of the medical device for which the quality management system is applied, the organization does not need to include such a requirement in its quality management system. For any clause that is determined to be not applicable, the organization records the justification as described in 4.2.2. | 1.2 Scope and Application This document is applicable to all parties involved in the supply-chain of medical device, covering authorized representatives of foreign manufacturers, importers or distributors of medical devices in Pakistan . The design and implementation of GDPMD by an establishment is dependent on the types, categories and classification of medical device, size and structure of the establishment and the processes employed. If any requirement in GDPMD is not applicable due to the type, category and classification of the medical device and supply-chain activities, a justification has to be provided for exclusion from fulfillment of that particular requirement.
4 Quality management system | 2. Quality management system
4.1 General requirements | 2.1 General Requirement
4.1.1 The organization shall document a quality management system and maintain its effectiveness in accordance with the requirements of this International Standard and applicable regulatory requirements. The organization shall establish, implement and maintain any requirement, procedure, activity or arrangement required to be documented by this International Standard or applicable regulatory requirements. The organization shall document the role(s) undertaken by the organization under the applicable regulatory requirements. NOTE Roles undertaken by the organization can include manufacturer, authorized representative, importer or distributor. | The establishment shall establish, document, implement and maintain a quality management system in accordance to the requirements of GDPMD.
4.1.2 The organization shall: a) determine the processes needed for the quality management system and the application of these processes throughout the organization taking into account the roles undertaken by the organization; b) apply a risk based approach to the control of the appropriate processes needed for the quality management system; c) determine the sequence and interaction of these processes. | /
4.1.3 For each quality management system process, the organization shall: a) determine criteria and methods needed to ensure that both the operation and control of these processes are effective; b) ensure the availability of resources and information necessary to support the operation and monitoring of these processes; c) implement actions necessary to achieve planned results and maintain the effectiveness of these processes; d) monitor, measure as appropriate, and analyse these processes; e) establish and maintain records needed to demonstrate conformance to this International Standard and compliance with applicable regulatory requirements (see 4.2.5). | /
4.1.4 The organization shall manage these quality management system processes in accordance with the requirements of this International Standard and applicable regulatory requirements. Changes to be made to these processes shall be: a) evaluated for their impact on the quality management system; b) evaluated for their impact on the medical devices produced under this quality management system; c) controlled in accordance with the requirements of this International Standard and applicable regulatory requirements. | /
4.1.5 When the organization chooses to outsource any process that affects product conformity to requirements, it shall monitor and ensure control over such processes. The organization shall retain responsibility of conformity to this International Standard and to customer and applicable regulatory requirements for outsourced processes. The controls shall be proportionate to the risk involved and the ability of the external party to meet the requirements in accordance with 7.4. The controls shall include written quality agreements. | Where an establishment chooses to outsource any activity that may affect the quality of medical devices, the establishment shall ensure control over such processes. 8. OUTSOURCE ACTIVITIES Where the establishment outsources any process within the scope of the GDPMD, the establishment shall ensure control over such processes. Certain processes such as cleaning, pest control, transportation etc may be outsourced to third party service providers. The establishment shall establish requirements to ensure that the outsourced activities conform to specified requirements.
4.1.6 The organization shall document procedures for the validation of the application of computer software used in the quality management system. Such software applications shall be validated prior to initial use and, as appropriate, after changes to such software or its application. The specific approach and activities associated with software validation and revalidation shall be proportionate to the risk associated with the use of the software. Records of such activities shall be maintained (see 4.2.5). | /
4.2 Documentation requirements | /
4.2.1 General | 2.1General Requirement
The quality management system documentation (see 4.2.4) shall include: a) documented statements of a quality policy and quality objectives; b) a quality manual; c) documented procedures and records required by this International Standard; d) documents, including records, determined by the organization to be necessary to ensure the effective planning, operation, and control of its processes; e) other documentation specified by applicable regulatory requirements. | The documentation required include: (i) a site master file, (ii) documented procedures required by the GDPMD, (iii) documents needed by the establishment to ensure the effective planning, operation and control of its processes, (iv) records required by the GDPMD, and (v) any other documentation specified by the regulatory authority.
4.2.2 Quality manual The organization shall document a quality manual that includes: a) the scope of the quality management system, including details of and justification for any exclusion or non-application; b) the documented procedures for the quality management system, or reference to them; c) a description of the interaction between the processes of the quality management system. The quality manual shall outline the structure of the documentation used in the quality management system. | 5 SITE MASTER FILE. The establishment shall establish and maintain a Site Master File as per format approved by the MDB.
4.2.3 Medical device file For each medical device type or medical device family, the organization shall establish and maintain one or more files either containing or referencing documents generated to demonstrate conformity to the requirement of this International Standard and compliance with applicable regulatory requirements. The content of the file(s) shall include, but is not limited to: a) general description of the medical device, intended use/purpose, and labelling, including any instructions for use; b) specifications for product; c) specifications or procedures for manufacturing, packaging, storage, handling and distribution; d) procedures for measuring and monitoring; e) as appropriate, requirements for installation; f) as appropriate, procedures for servicing. | /
4.2.4 Control of documents Documents required by the quality management system shall be controlled. Records are a special type of document and shall be controlled according to the requirements given in 4.2.5. A documented procedure shall define the controls needed to: a) review and approve documents for adequacy prior to issue; b) review, update as necessary and re-approve documents; c) ensure that the current revision status of and changes to documents are identified; d) ensure that relevant versions of applicable documents are available at points of use; e) ensure that documents remain legible and readily identifiable; f) ensure that documents of external origin, determined by the organization to be necessary for the planning and operation of the quality management system, are identified and their distribution controlled; g) prevent deterioration or loss of documents; h) prevent the unintended use of obsolete documents and apply suitable identification to them. The organization shall ensure that changes to documents are reviewed and approved either by the original approving function or another designated function that has access to pertinent background information upon which to base its decisions. The organization shall define the period for which at least one copy of obsolete documents shall be retained. This period shall ensure that documents to which medical devices have been manufactured and tested are available for at least the lifetime of the medical device as defined by the organization, but not less than the retention period of any resulting record (see 4.2.5), or as specified by applicable regulatory requirements. | 2.2 Control of Documents and Records Documents required for GDPMD shall be controlled and records for each purchase and sale, showing the date of purchase and supply, name of medical device, quantity received and supplied and name and address of supplier or consignee shall be established and maintained to provide evidence of conformity to requirements of GDPMD. Record shall be legible, readily identifiable and retrievable. A documented procedure shall be established for the identification, storage, protection, retrieval, retention time and disposal of record. The establishment shall retain the record for a period of time:- (i) specified by relevant regulatory requirements, or (ii) at least equivalent to the lifetime of the medical device as defined by the product owner of the medical devices, whichever is the longest. All documents shall be prepared, approved, signed and dated by an appropriate authorized person(s). Documents shall be reviewed regularly and kept up-to-date. If a document has been revised, a control system shall be established to prevent the unintended use of the superseded version. Where an electronic record system is used in place of a paper-based system, the system utilized should have built-in checks and balances to ensure the integrity of the record and to protect against unauthorized entries and the system should also incorporate audit trails for tracking changes.
4.2.5 Control of records Records shall be maintained to provide evidence of conformity to requirements and of the effective operation of the quality management system. The organization shall document procedures to define the controls needed for the identification, storage, security and integrity, retrieval, retention time and disposition of records. The organization shall define and implement methods for protecting confidential health information contained in records in accordance with the applicable regulatory requirements. Records shall remain legible, readily identifiable and retrievable. Changes to a record shall remain identifiable. The organization shall retain the records for at least the lifetime of the medical device as defined by the organization, or as specified by applicable regulatory requirements, but not less than two years from the medical device release by the organization. | Records providing traceability of medical devices from supplier and to the customers shall be maintained.
5 Management responsibility | /
5.1 Management commitment Top management shall provide evidence of its commitment to the development and implementation of the quality management system and maintenance of its effectiveness by: a) communicating to the organization the importance of meeting customer as well as applicable regulatory requirements; b) establishing the quality policy; c) ensuring that quality objectives are established; d) conducting management reviews; e) ensuring the availability of resources. | /
5.2 Customer focus Top management shall ensure that customer requirements and applicable regulatory requirements are determined and met. | /
5.3 Quality policy Top management shall ensure that the quality policy: a) is applicable to the purpose of the organization; b) includes a commitment to comply with requirements and to maintain the effectiveness of the quality management system; c) provides a framework for establishing and reviewing quality objectives; d) is communicated and understood within the organization; e) is reviewed for continuing suitability. | /
5.4 Planning
5.4.1 Quality objectives Top management shall ensure that quality objectives, including those needed to meet applicable regulatory requirements and requirements for product, are established at relevant functions and levels within the organization. The quality objectives shall be measurable and consistent with the quality policy. | /
5.4.2 Quality management system planning Top management shall ensure that: a) the planning of the quality management system is carried out in order to meet the requirements given in 4.1, as well as the quality objectives; b) the integrity of the quality management system is maintained when changes to the quality management system are planned and implemented. | /
5.5 Responsibility, authority and communication
5.5.1 Responsibility and authority Top management shall ensure that responsibilities and authorities are defined, documented and communicated within the organization. Top management shall document the interrelation of all personnel who manage, perform and verify work affecting quality and shall ensure the independence and authority necessary to perform these tasks. | 3.3 Responsibility and Authority The establishment shall ensure that responsibilities and authorities are defined, documented and communicated within the establishment. This should be documented in site master file as well.
5.5.2 Management representative Top management shall appoint a member of management who, irrespective of other responsibilities, has responsibility and authority that includes: a) ensuring that processes needed for the quality management system are documented; b) reporting to top management on the effectiveness of the quality management system and any need for improvement; c) ensuring the promotion of awareness of applicable regulatory requirements and quality management system requirements throughout the organization. | 3.4 Designated person The establishment shall appoint designated person who, irrespective of other responsibilities, shall have the ultimate responsibility of: (i) Ensuring that processes needed for the quality management system are established, implemented andmaintained and reporting to top management on the performance of the quality management system and any need for improvement, and (ii) Ensuring the promotion of awareness of regulatory and customer requirements throughout the establishment and liaising with external parties on matters relating to the quality managementsystem and regulatory requirements.
5.5.3 Internal communication Top management shall ensure that appropriate communication processes are established within the organization and that communication takes place regarding the effectiveness of the quality management system. | /
5.6 Management review | 2.7Internal Check System, Management Input and Output Review
5.6.1 General The organization shall document procedures for management review. Top management shall review the organization's quality management system at documented planned intervals to ensure its continuing suitability, adequacy and effectiveness. The review shall include assessing opportunities for improvement and the need for changes to the quality management system, including the quality policy and quality objectives. Records from management reviews shall be maintained (see 4.2.5). | The establishment top management should identify, review and implement any changes necessary to ensure and maintain the continued suitability and effectiveness of the quality system through the use of the quality policy, quality objectives, checking results, analysis of data, corrective and preventive actions and management review. This review shall include assessing opportunities for improvement and the need for changes to the quality management system. Records from management reviews shall be maintained. The input to management review may also include information on: (i) customer complaints/feedback; (ii) surveillance and vigilance activities including field corrective actions, advisory notes, recalls and adverse event /incident reporting; (iii) feedback from manufacturer; (iv) feedback and directives from the authority; (v) status of preventive and corrective actions; (vi) changes that could affect the GDPMD regulatory compliance system; and (vii) recommendations for compliance. The output from the management review shall include any decision and action related to:- (i) the corrective and preventive actions required; and (ii) the effectiveness of the GDPMD regulatory compliance system and its compliance with the regulatory requirements. Records of the management review should include the identity of those taking part in the review and all points of the review including description of any corrective or preventive action to be taken. For any action determined, the responsibility for such actions, the resources, target dates for completion, etc should be identified.
5.6.2 Review input The input to management review shall include, but is not limited to, information arising from: a) feedback; b) complaint handling; c) reporting to regulatory authorities; d) audits; e) monitoring and measurement of processes; f) monitoring and measurement of product; g) corrective action; h) preventive action; i) follow-up actions from previous management reviews; j) changes that could affect the quality management system; k) recommendations for improvement; l) applicable new or revised regulatory requirements.
5.6.3 Review output The output from management review shall be recorded (see 4.2.5) and include the input reviewed and any decisions and actions related to: a) improvement needed to maintain the suitability, adequacy, and effectiveness of the quality management system and its processes; b) improvement of product related to customer requirements; c) changes needed to respond to applicable new or revised regulatory requirements; d) resource needs.
6 Resource management | 3 RESOURCE MANAGEMENT
6.1 Provision of resources The organization shall determine and provide the resources needed to: a) implement the quality management system and to maintain its effectiveness; b) meet applicable regulatory and customer requirements. | /
6.2 Human resources Personnel performing work affecting product quality shall be competent on the basis of appropriate education, training, skills and experience. The organization shall document the process(es) for establishing competence, providing needed training, and ensuring awareness of personnel. The organization shall: a) determine the necessary competence for personnel performing work affecting product quality; b) provide training or take other actions to achieve or maintain the necessary competence; c) evaluate the effectiveness of the actions taken; d) ensure that its personnel are aware of the relevance and importance of their activities and how they contribute to the achievement of the quality objectives; e) maintain appropriate records of education, training, skills and experience (see 4.2.5). NOTE The methodology used to check effectiveness is proportionate to the risk associated with the work for which the training or other action is being provided. | 3.1 Personnel Personnel working at the premises shall be sufficient for complying GDPMD. Qualified person shall possess the prescribed qualification. 3.2 Training The establishment shall (i) provide training to key personnel to satisfyneeds, and (ii) maintain training record. Training for personnel will be tailored to the person’s assignment. Typical training and education should cover the: (i) nature of activities, (ii) health, safety and environmentalregulations and/or regulatory body related requirements, (iii) establishment’s policies, (iv) function of the personnel, and (v) procedures and instructions of relevance topersonnel. (vi) Special training may be necessary for personnel dealing with certain categories of substances/materials such as chemicals, biological, radiation emitting or energy source components and products.
6.3 Infrastructure The organization shall document the requirements for the infrastructure needed to achieve conformity to product requirements, prevent product mix-up and ensure orderly handling of product. Infrastructure includes, as appropriate: a) buildings, workspace and associated utilities; b) process equipment (both hardware and software); c) supporting services (such as transport, communication, or information systems). The organization shall document requirements for the maintenance activities, including the interval of performing the maintenance activities, when such maintenance activities, or lack thereof, can affect product quality. As appropriate, the requirements shall apply to equipment used in production, the control of the work environment and monitoring and measurement. Records of such maintenance shall be maintained (see 4.2.5). | 3.5 Premises and Facilities. The establishment shall ensure that the premises and equipment used are suitable and adequate to ensure proper conservation and distribution of medical devices. Fire extinguishers, smoke detector etc for control of fire shall be available. Storage areas should be cleaned and accumulated waste removed at regular intervals. The frequency and methods of cleaning the premises and areas should be recorded. No smoking, eating and drinking should be permitted in areas used for storage and handling of medical devices. Buildings should protect medical devices from contamination and deterioration, including protection from excessive heat or undue exposure to direct sunlight. Premises should be constructed, serviced and maintained regularly to protect stored medical devices from all potentially harmful influences such as undue variations of temperature and humidity.
6.4 Work environment and contamination control | /
6.4.1 Work environment The organization shall document the requirements for the work environment needed to achieve conformity to product requirements. If the conditions for the work environment can have an adverse effect on product quality, the organization shall document the requirements for the work environment and the procedures to monitor and control the work environment. The organization shall: a) document requirements for health, cleanliness and clothing of personnel if contact between such personnel and the product or work environment could affect medical device safety or performance; b) ensure that all personnel who are required to work temporarily under special environmental conditions within the work environment are competent or supervised by a competent person. NOTE Further information can be found in ISO 14644 and ISO 14698. | 3.6 Cleanliness and Pest Control The establishment shall establish documented procedure for cleaning of premises including frequency and methods. Records of cleaning shall be maintained. The establishment shall document and ensure a pest control mechanism to identify and prevent pest infestation. Record of pest control mechanism shall be maintained.
6.4.2 Contamination control As appropriate, the organization shall plan and document arrangements for the control of contaminated or potentially contaminated product in order to prevent contamination of the work environment, personnel, or product. For sterile medical devices, the organization shall document requirements for control of contamination with microorganisms or particulate matter and maintain the required cleanliness during assembly or packaging processes. | /
7 Product realization | /
7.1 Planning of product realization The organization shall plan and develop the processes needed for product realization. Planning of product realization shall be consistent with the requirements of the other processes of the quality management system. The organization shall document one or more processes for risk management in product realization. Records of risk management activities shall be maintained (see 4.2.5). | /
In planning product realization, the organization shall determine the following, as appropriate: a) quality objectives and requirements for the product; b) the need to establish processes and documents (see 4.2.4) and to provide resources specific to the product, including infrastructure and work environment; c) required verification, validation, monitoring, measurement, inspection and test, handling, storage, distribution and traceability activities specific to the product together with the criteria for product acceptance; d) records needed to provide evidence that the realization processes and resulting product meet requirements (see 4.2.5). The output of this planning shall be documented in a form suitable for the organization’s method of operations. NOTE Further information can be found in ISO 14971. | /
7.2 Customer-related processes | /
7.2.1 Determination of requirements related to product The organization shall determine: a) requirements specified by the customer, including the requirements for delivery and post-delivery activities; b) requirements not stated by the customer but necessary for specified or intended use, as known; c) applicable regulatory requirements related to the product; d) any user training needed to ensure specified performance and safe use of the medical device; e) any additional requirements determined by the organization. | /
7.2.2 Review of requirements related to product The organization shall review the requirements related to product. This review shall be conducted prior to the organization’s commitment to supply product to the customer (e.g. submission of tenders, acceptance of contracts or orders, acceptance of changes to contracts or orders) and shall ensure that: a) product requirements are defined and documented; b) contract or order requirements differing from those previously expressed are resolved; c) applicable regulatory requirements are met; d) any user training identified in accordance with 7.2.1 is available or planned to be available; e) the organization has the ability to meet the defined requirements. Records of the results of the review and actions arising from the review shall be maintained (see 4.2.5). When the customer provides no documented statement of requirement, the customer requirements shall be confirmed by the organization before acceptance. When product requirements are changed, the organization shall ensure that relevant documents are amended and that relevant personnel are made aware of the changed requirements. | /
7.2.3 Communication The organization shall plan and document arrangements for communicating with customers in relation to: a) product information; b) enquiries, contracts or order handling, including amendments; c) customer feedback, including complaints; d) advisory notices. The organization shall communicate with regulatory authorities in accordance with applicable regulatory requirements. | 4 AUTHORIZATION The establishment shall— (i) obtain appropriate authorization to become authorized representative, importer or distributor of medical devices; and (ii) establish and maintain written agreement with the relevant party regarding supply of information required for regulatory matters relating to medical devices it deals with. | 没有完全匹配的条款，与7.2.3 b) Contracts 最接近
7.3 Design and development
7.3.1 General The organization shall document procedures for design and development. | /
7.3.2 Design and development planning The organization shall plan and control the design and development of product. As appropriate, design and development planning documents shall be maintained and updated as the design and development progresses. During design and development planning, the organization shall document: a) the design and development stages; b) the review(s) needed at each design and development stage; c) the verification, validation, and design transfer activities that are appropriate at each design and development stage; d) the responsibilities and authorities for design and development; e) the methods to ensure traceability of design and development outputs to design and development inputs; f) the resources needed, including necessary competence of personnel. | /
7.3.3 Design and development inputs Inputs relating to product requirements shall be determined and records maintained (see 4.2.5). These inputs shall include: a) functional, performance, usability and safety requirements, according to the intended use; b) applicable regulatory requirements and standards; c) applicable output(s) of risk management; d) as appropriate, information derived from previous similar designs; e) other requirements essential for design and development of the product and processes. These inputs shall be reviewed for adequacy and approved. Requirements shall be complete, unambiguous, able to be verified or validated, and not in conflict with each other. NOTE Further information can be found in IEC 62366–1. | /
7.3.4 Design and development outputs Design and development outputs shall: a) meet the input requirements for design and development; b) provide appropriate information for purchasing, production and service provision; c) contain or reference product acceptance criteria; d) specify the characteristics of the product that are essential for its safe and proper use. The outputs of design and development shall be in a form suitable for verification against the design and development inputs and shall be approved prior to release. Records of the design and development outputs shall be maintained (see 4.2.5). | /
7.3.5 Design and development review At suitable stages, systematic reviews of design and development shall be performed in accordance with planned and documented arrangements to: a) evaluate the ability of the results of design and development to meet requirements; b) identify and propose necessary actions. Participants in such reviews shall include representatives of functions concerned with the design and development stage being reviewed, as well as other specialist personnel. Records of the results of the reviews and any necessary actions shall be maintained and include the identification of the design under review, the participants involved and the date of the review (see 4.2.5). | /
7.3.6 Design and development verification Design and development verification shall be performed in accordance with planned and documented arrangements to ensure that the design and development outputs have met the design and development input requirements. The organization shall document verification plans that include methods, acceptance criteria and, as appropriate, statistical techniques with rationale for sample size. If the intended use requires that the medical device be connected to, or have an interface with, other medical device(s), verification shall include confirmation that the design outputs meet design inputs when so connected or interfaced. Records of the results and conclusions of the verification and necessary actions shall be maintained (see 4.2.4 and 4.2.5). | /
7.3.7 Design and development validation Design and development validation shall be performed in accordance with planned and documented arrangements to ensure that the resulting product is capable of meeting the requirements for the specified application or intended use. The organization shall document validation plans that include methods, acceptance criteria and, as appropriate, statistical techniques with rationale for sample size. Design validation shall be conducted on representative product. Representative product includes initial production units, batches or their equivalents. The rationale for the choice of product used for validation shall be recorded (see 4.2.5). As part of design and development validation, the organization shall perform clinical evaluations or performance evaluations of the medical device in accordance with applicable regulatory requirements. A medical device used for clinical evaluation or performance evaluation is not considered to be released for use to the customer. If the intended use requires that the medical device be connected to, or have an interface with, other medical device(s), validation shall include confirmation that the requirements for the specified application or intended use have been met when so connected or interfaced. Validation shall be completed prior to release for use of the product to the customer. Records of the results and conclusion of validation and necessary actions shall be maintained (see 4.2.4 and 4.2.5). | /
7.3.8 Design and development transfer The organization shall document procedures for transfer of design and development outputs to manufacturing. These procedures shall ensure that design and development outputs are verified as suitable for manufacturing before becoming final production specifications and that production capability can meet product requirements. Results and conclusions of the transfer shall be recorded (see 4.2.5). | /
7.3.9 Control of design and development changes The organization shall document procedures to control design and development changes. The organization shall determine the significance of the change to function, performance, usability, safety and applicable regulatory requirements for the medical device and its intended use. Design and development changes shall be identified. Before implementation, the changes shall be: a) reviewed; b) verified; c) validated, as appropriate; d) approved. The review of design and development changes shall include evaluation of the effect of the changes on constituent parts and product in process or already delivered, inputs or outputs of risk management and product realization processes. Records of changes, their review and any necessary actions shall be maintained (see 4.2.5). | /
7.3.10 Design and development files The organization shall maintain a design and development file for each medical device type or medical device family. This file shall include or reference records generated to demonstrate conformity to the requirements for design and development and records for design and development changes. | /
7.4 Purchasing | /
7.4.1 Purchasing process The organization shall document procedures (see 4.2.4) to ensure that purchased product conforms to specified purchasing information. The organization shall establish criteria for the evaluation and selection of suppliers. The criteria shall be: a) based on the supplier’s ability to provide product that meets the organization’s requirements; b) based on the performance of the supplier; c) based on the effect of the purchased product on the quality of the medical device; d) proportionate to the risk associated with the medical device. The organization shall plan the monitoring and re-evaluation of suppliers. Supplier performance in meeting requirements for the purchased product shall be monitored. The results of the monitoring shall provide an input into the supplier re-evaluation process. Non-fulfilment of purchasing requirements shall be addressed with the supplier proportionate to the risk associated with the purchased product and compliance with applicable regulatory requirements. Records of the results of evaluation, selection, monitoring and re-evaluation of supplier capability or performance and any necessary actions arising from these activities shall be maintained (see 4.2.5). | /
7.4.2 Purchasing information Purchasing information shall describe or reference the product to be purchased, including as appropriate: a) product specifications; b) requirements for product acceptance, procedures, processes and equipment; c) requirements for qualification of supplier personnel; d) quality management system requirements. The organization shall ensure the adequacy of specified purchasing requirements prior to their communication to the supplier. Purchasing information shall include, as applicable, a written agreement that the supplier notify the organization of changes in the purchased product prior to implementation of any changes that affect the ability of the purchased product to meet specified purchase requirements. To the extent required for traceability given in 7.5.9, the organization shall maintain relevant purchasing information in the form of documents (see 4.2.4) and records (see 4.2.5). | /
7.4.3 Verification of purchased product The organization shall establish and implement the inspection or other activities necessary for ensuring that purchased product meets specified purchasing requirements. The extent of verification activities shall be based on the supplier evaluation results and proportionate to the risks associated with the purchased product. When the organization becomes aware of any changes to the purchased product, the organization shall determine whether these changes affect the product realization process or the medical device. When the organization or its customer intends to perform verification at the supplier’s premises, the organization shall state the intended verification activities and method of product release in the purchasing information. Records of the verification shall be maintained (see 4.2.5). | /
7.5 Production and service provision | /
7.5.1 Control of production and service provision Production and service provision shall be planned, carried out, monitored and controlled to ensure that product conforms to specification. As appropriate, production controls shall include but are not limited to: a) documentation of procedures and methods for the control of production (see 4.2.4); b) qualification of infrastructure; c) implementation of monitoring and measurement of process parameters and product characteristics; d) availability and use of monitoring and measuring equipment; e) implementation of defined operations for labelling and packaging; f) implementation of product release, delivery and post-delivery activities. The organization shall establish and maintain a record (see 4.2.5) for each medical device or batch of medical devices that provides traceability to the extent specified in 7.5.9 and identifies the amount manufactured and amount approved for distribution. The record shall be verified and approved. | /
7.5.2 Cleanliness of product The organization shall document requirements for cleanliness of product or contamination control of product if: a) product is cleaned by the organization prior to sterilization or its use; b) product is supplied non-sterile and is to be subjected to a cleaning process prior to sterilization or its use; c) product cannot be cleaned prior to sterilization or its use, and its cleanliness is of significance in use; d) product is supplied to be used non-sterile, and its cleanliness is of significance in use; e) process agents are to be removed from product during manufacture. If product is cleaned in accordance with a) or b) above, the requirements contained in 6.4.1 do not apply prior to the cleaning process. | /
7.5.3 Installation activities The organization shall document requirements for medical device installation and acceptance criteria for verification of installation, as appropriate. If the agreed customer requirements allow installation of the medical device to be performed by an external party other than the organization or its supplier, the organization shall provide documented requirements for medical device installation and verification of installation. Records of medical device installation and verification of installation performed by the organization or its supplier shall be maintained (see 4.2.5). | 6.1 Installation and Servicing The establishment shall where applicable— (i) establish and maintain documented procedures, work instructions and reference materials, tools and test equipment and reference measurement procedures, for performing servicing activities including calibration, repair, maintenance and verifying that they meet the regulatory requirements and applicable standards; (ii) establish documented requirements which contain acceptance criteria for installation, testing and commissioning of the medical device; (iii) establish installation qualification and maintain adequate installation and inspection instructions for medical devices requiring specified installation requirements, and where appropriate, test procedures; (iv) ensure proper installation, testing and commissioning; (v) ensure equipment used for testing, maintenance and conservation of medical devices are calibrated or verified at specific intervals; (vi) ensure the calibration and maintenance of test equipment conforms to the applicable standards; and (vii) maintain testing and commissioning, installation, calibration and maintenance service records. The establishment shall, as appropriate— (i) establish an appropriate technical support which include maintenance service, training, calibration, management of spare parts, workshop setup and management; (ii) establish maintenance management mechanism to support the customers; (iii) ensure the technical and maintenance support services for active medical devices conform to the applicable regulatory requirements.
7.5.4 Servicing activities If servicing of the medical device is a specified requirement, the organization shall document servicing procedures, reference materials, and reference measurements, as necessary, for performing servicing activities and verifying that product requirements are met. The organization shall analyse records of servicing activities carried out by the organization or its supplier: a) to determine if the information is to be handled as a complaint; b) as appropriate, for input to the improvement process. Records of servicing activities carried out by the organization or its supplier shall be maintained (see 4.2.5).
7.5.5 Particular requirements for sterile medical devices The organization shall maintain records of the sterilization process parameters used for each sterilization batch (see 4.2.5). Sterilization records shall be traceable to each production batch of medical devices. | /
7.5.6 Validation of processes for production and service provision The organization shall validate any processes for production and service provision where the resulting output cannot be or is not verified by subsequent monitoring or measurement and, as a consequence, deficiencies become apparent only after the product is in use or the service has been delivered. Validation shall demonstrate the ability of these processes to achieve planned results consistently. The organization shall document procedures for validation of processes, including: a) defined criteria for review and approval of the processes; b) equipment qualification and qualification of personnel; c) use of specific methods, procedures and acceptance criteria; d) as appropriate, statistical techniques with rationale for sample sizes; e) requirements for records (see 4.2.5); f) revalidation, including criteria for revalidation; g) approval of changes to the processes. The organization shall document procedures for the validation of the application of computer software used in production and service provision. Such software applications shall be validated prior to initial use and, as appropriate, after changes to such software or its application. The specific approach and activities associated with software validation and revalidation shall be proportionate to the risk associated with the use of the software, including the effect on the ability of the product to conform to specifications. Records of the results and conclusion of validation and necessary actions from the validation shall be maintained (see 4.2.4 and 4.2.5). | /
7.5.7 Particular requirements for validation of processes for sterilization and sterile barrier systems The organization shall document procedures (see 4.2.4) for the validation of processes for sterilization and sterile barrier systems. Processes for sterilization and sterile barrier systems shall be validated prior to implementation and following product or process changes, as appropriate. Records of the results and, conclusion of validation and necessary actions from the validation shall be maintained (see 4.2.4 and 4.2.5). NOTE Futher information can be found in ISO 11607-1 and ISO 11607-2. | /
7.5.8 Identification The organization shall document procedures for product identification and identify product by suitable means throughout product realization. The organization shall identify product status with respect to monitoring and measurement requirements throughout product realization. Identification of product status shall be maintained throughout production, storage, installation and servicing of product to ensure that only product that has passed the required inspections and tests or released under an authorized concession is dispatched, used or installed. If required by applicable regulatory requirements, the organization shall document a system to assign unique device identification to the medical device. The organization shall document procedures to ensure that medical devices returned to the organization are identified and distinguished from conforming product. | 7. SEGREGATION OF COUNTERFEIT, ADULTERATED, MISBRANDED, SPURIOUS AND SUBSTANDARD MEDICAL DEVICES Counterfeit, adulterated, misbranded, spurious and substandard medical devices, if found in the distribution network, shall be physically segregated from other medical devices to avoid any confusion. These shall be clearly labelled as “Not for Sale” or in other similar phrases/words. Appropriate corrective measures should be undertaken for identified counterfeit, adulterate, misbranded, spurious or substandard medical device, which includes but are not limited to the following:- (i) Segregation/quarantine of these medicaldevices, (ii) Investigation of supply chain breach. Communication is required to all affected wholesalers / distributors / retailers / consignees possibly supplied with these medical devices prior to notification to the regulatory authority for such communication
7.5.9 Traceability | /
7.5.9.1 General The organization shall document procedures for traceability. These procedures shall define the extent of traceability in accordance with applicable regulatory requirements and the records to be maintained (see 4.2.5). | /
7.5.9.2 Particular requirements for implantable medical devices The records required for traceability shall include records of components, materials, and conditions for the work environment used, if these could cause the medical device not to satisfy its specified safety and performance requirements. The organization shall require that suppliers of distribution services or distributors maintain records of the distribution of medical devices to allow traceability and that these records are available for inspection. Records of the name and address of the shipping package consignee shall be maintained (see 4.2.5). | /
7.5.10 Customer property The organization shall identify, verify, protect, and safeguard customer property provided for use or incorporation into the product while it is under the organization’s control or being used by the organization. If any customer property is lost, damaged or otherwise found to be unsuitable for use, the organization shall report this to the customer and maintain records (see 4.2.5). | /
7.5.11 Preservation of product The organization shall document procedures for preserving the conformity of product to requirements during processing, storage, handling, and distribution. Preservation shall apply to the constituent parts of a medical device. The organization shall protect product from alteration, contamination or damage when exposed to expected conditions and hazards during processing, storage, handling, and distribution by: a) designing and constructing suitable packaging and shipping containers; b) documenting requirements for special conditions needed if packaging alone cannot provide preservation. If special conditions are required, they shall be controlled and recorded (see 4.2.5). | 6 STORAGE AND STOCK HANDLING, TRACEABILITY AND SUPPLY CHAIN MANAGEMENT • The establishment shall provide suitable and adequate storage to ensure proper conservation of the medical devices and shall provide back up for electricity to ensure proper conservation of the medical devices during electric load shedding, wherever required. • The establishment shall ensure proper receipt of medical devices. • Medical devices requiring special storage conditions (e.g. temperature and/or humidity or narcotics requiring additional security measures) should be placed in separate areas equipped to provide the desired conditions. A list of such medical devices should be maintained and the medical devices properly identified. Storage conditions shall be monitored and recorded periodically and Records shall be maintained. During transportation special measure should be taken to maintain conditioned required to prevent effect on the integrity and quality of the medical devices. • There should be adequate storage areas, and where applicable, physically separated zones for the orderly segregation of saleable stock, quarantined, expired, rejected/damaged, recalled and returned medical devices ensured, stored off the ground and suitably spaced having adequate lighting and ventilation. • Appropriate and suitable storage conditions should be provided for hazardous, sensitive and dangerous materials such as combustible liquids and solids, pressurized gases, highly toxic and radioactive substances. • Adequate precautions should be taken against spillage or breakage, attack by micro- organisms, contamination and cross-contamination. • Where controlled environmental storage conditions are required, these conditions should be continuously monitored and documented. The actual storage temperature should be expressed quantitatively. Where the storage temperature is not expressed quantitatively or stated (in terms of a range) on the labels of the registered medical device. • The establishment shall establish a system to ensure stock rotation with proper label information. • The expired medical devices shall be disposed of as per procedure in this guidance document. • Medical devices bearing an expiry date must not be received or supplied close to or after the expiry date such that this date is likely to occur before the consumer uses the medical devices. • All labels and containers of medical devices should not be altered, tampered or changed. • The establishment shall establish adequate methods of transportation and deliveries should be made only to authorized wholesalers, distributors or person to achieve safe and secure delivery of all medical devices from their point of collection to their point of delivery. Medical devices shall be transported in such a way that:- (i) their identification is not lost; (ii) they do not contaminate, and are not contaminated by, other medical devices or materials/substances; (iii) adequate precautions are taken against spillage, breakage or theft; (iv) they are secure and not subjected to unacceptable degrees of heat, cold, light, moisture or other adverse influence, or to attack by microorganisms and pests. | The establishment shall ensure proper receipt of medical devices.对应7.5.9 The establishment shall establish a system to ensure stock rotation with proper label information.对应7.5.8 The expired medical devices shall be disposed of as per procedure in this guidance document.对应7.5.11 Medical devices bearing an expiry date must not be received or supplied close to or after the expiry date such that this date is likely to occur before the consumer uses the medical devices对应8.3.2 All labels and containers of medical devices should not be altered, tampered or changed.对应7.5.11和7.5.8 • The establishment shall establish adequate methods of transportation and deliveries should be made only to authorized wholesalers, distributors or person to achieve safe and secure delivery of all medical devices from their point of collection to their point of delivery. Medical devices shall be transported in such a way that:- (i) their identification is not lost; (ii) they do not contaminate, and are not contaminated by, other medical devices or materials/substances; (iii) adequate precautions are taken against spillage, breakage or theft; (iv) they are secure and not subjected to unacceptable degrees of heat, cold, light, moisture or other adverse influence, or to attack by microorganisms and pests.对应7.5.11
7.6 Control of monitoring and measuring equipment The organization shall determine the monitoring and measurement to be undertaken and the monitoring and measuring equipment needed to provide evidence of conformity of product to determined requirements. The organization shall document procedures to ensure that monitoring and measurement can be carried out and are carried out in a manner that is consistent with the monitoring and measurement requirements. As necessary to ensure valid results, measuring equipment shall: a) be calibrated or verified, or both, at specified intervals, or prior to use, against measurement standards traceable to international or national measurement standards: when no such standards exist, the basis used for calibration or verification shall be recorded (see 4.2.5); b) be adjusted or re-adjusted as necessary: such adjustments or re-adjustments shall be recorded (see 4.2.5); c) have identification in order to determine its calibration status; d) be safeguarded from adjustments that would invalidate the measurement result; e) be protected from damage and deterioration during handling, maintenance and storage. The organization shall perform calibration or verification in accordance with documented procedures. In addition, the organization shall assess and record the validity of the previous measuring results when the equipment is found not to conform to requirements. The organization shall take appropriate action in regard to the equipment and any product affected. Records of the results of calibration and verification shall be maintained (see 4.2.5). The organization shall document procedures for the validation of the application of computer software used for the monitoring and measurement of requirements. Such software applications shall be validated prior to initial use and, as appropriate, after changes to such software or its application. The specific approach and activities associated with software validation and revalidation shall be proportionate to the risk associated with the use of the software, including the effect on the ability of the product to conform to specifications. Records of the results and conclusion of validation and necessary actions from the validation shall be maintained (see 4.2.4 and 4.2.5). NOTE Futher information can be found in ISO 10012. | 3.7 Calibration The instruments used for measuring and monitoring temperature and humidity shall be calibrated or verified for accuracy from accredited laboratories at defined intervals, or prior to use and the results of such calibrations or verifications shall be recorded and retained.
8 Measurement, analysis and improvement | /
8.1 General The organization shall plan and implement the monitoring, measurement, analysis and improvement processes needed to: a) demonstrate conformity of product; b) ensure conformity of the quality management system; c) maintain the effectiveness of the quality management system. This shall include determination of appropriate methods, including statistical techniques, and the extent of their use. | /
8.2 Monitoring and measurement 8.2.1 Feedback As one of the measurements of the effectiveness of the quality management system, the organization shall gather and monitor information relating to whether the organization has met customer requirements. The methods for obtaining and using this information shall be documented. The organization shall document procedures for the feedback process. This feedback process shall include provisions to gather data from production as well as post-production activities. The information gathered in the feedback process shall serve as potential input into risk management for monitoring and maintaining the product requirements as well as the product realization or improvement processes. If applicable regulatory requirements require the organization to gain specific experience from post- production activities, the review of this experience shall form part of the feedback process. | /
8.2.2 Complaint handling The organization shall document procedures for timely complaint handling in accordance with applicable regulatory requirements. These procedures shall include at a minimum requirements and responsibilities for: a) receiving and recording information; b) evaluating information to determine if the feedback constitutes a complaint; c) investigating complaints; d) determining the need to report the information to the appropriate regulatory authorities; e) handling of complaint-related product; f) determining the need to initiate corrections or corrective actions. If any complaint is not investigated, justification shall be documented. Any correction or corrective action resulting from the complaint handling process shall be documented. If an investigation determines activities outside the organization contributed to the complaint, relevant information shall be exchanged between the organization and the external party involved. Complaint handling records shall be maintained (see 4.2.5). | 2.3 Complaints The establishment shall establish a documented procedure for handling of all written and oral complaints regarding medical devices and record should be established. The procedure for handling complaints shall ensure that the complaints received are investigated and followed through, and that corrective actions are taken to prevent repeated complaints, and, where a decision is made to recall the medical device, the details of the recall. Any report of adverse event which requires regulatory reporting shall be reported to the regulatory authority as per procedure and regulatory requirements. Records of the complaint, investigation and any subsequent actions taken shall be maintained. The investigation should take into consideration the condition and circumstances under which the medical device was distributed, stored and used.
8.2.3 Reporting to regulatory authorities If applicable regulatory requirements require notification of complaints that meet specified reporting criteria of adverse events or issuance of advisory notices, the organization shall document procedures for providing notification to the appropriate regulatory authorities. Records of reporting to regulatory authorities shall be maintained (see 4.2.5). | /
8.2.4 Internal audit The organization shall conduct internal audits at planned intervals to determine whether the quality management system: a) conforms to planned and documented arrangements, requirements of this International Standard, quality management system requirements established by the organization, and applicable regulatory requirements; b) is effectively implemented and maintained. The organization shall document a procedure to describe the responsibilities and requirements for planning and conducting audits and recording and reporting audit results. An audit program shall be planned, taking into consideration the status and importance of the processes and area to be audited, as well as the results of previous audits. The audit criteria, scope, interval and methods shall be defined and recorded (see 4.2.5). The selection of auditors and conduct of audits shall ensure objectivity and impartiality of the audit process. Auditors shall not audit their own work. Records of the audits and their results, including identification of the processes and areas audited and the conclusions, shall be maintained (see 4.2.5). The management responsible for the area being audited shall ensure that any necessary corrections and corrective actions are taken without undue delay to eliminate detected nonconformities and their causes. Follow-up activities shall include the verification of the actions taken and the reporting of verification results. NOTE Futher information can be found in ISO 19011 | 2.7 Internal Check System, Management Input and Output Review The establishment shall establish internal check system to monitor the implementation of and compliance with the requirements of GDPMD. Internal checking should normally be conducted once a year. The result of the checking is usually stated in a written report indicating the nonconformities found. Timely action should be taken to eliminate the nonconformities and their causes. The checking results should be communicated to management for review.
8.2.5 Monitoring and measurement of processes The organization shall apply suitable methods for monitoring and, as appropriate, measurement of the quality management system processes. These methods shall demonstrate the ability of the processes to achieve planned results. When planned results are not achieved, correction and corrective action shall be taken, as appropriate. | /
8.2.6 Monitoring and measurement of product The organization shall monitor and measure the characteristics of the product to verify that product requirements have been met. This shall be carried out at applicable stages of the product realization process in accordance with the planned and documented arrangements and documented procedures. Evidence of conformity to the acceptance criteria shall be maintained. The identity of the person authorizing release of product shall be recorded (see 4.2.5). As appropriate, records shall identify the test equipment used to perform measurement activities. Product release and service delivery shall not proceed until the planned and documented arrangements have been satisfactorily completed. For implantable medical devices, the organization shall record the identity of personnel performing any inspection or testing. | /
8.3 Control of nonconforming product | /
8.3.1 General The organization shall ensure that product which does not conform to product requirements is identified and controlled to prevent its unintended use or delivery. The organization shall document a procedure to define the controls and related responsibilities and authorities for the identification, documentation, segregation, evaluation and disposition of nonconforming product. The evaluation of nonconformity shall include a determination of the need for an investigation and notification of any external party responsible for the nonconformity. Records of the nature of the nonconformities and any subsequent action taken, including the evaluation, any investigation and the rationale for decisions shall be maintained (see 4.2.5) | /
8.3.2 Actions in response to nonconforming product detected before delivery The organization shall deal with nonconforming product by one or more of the following ways: a) taking action to eliminate the detected nonconformity; b) taking action to preclude its original intended use or application; c) authorizing its use, release or acceptance under concession. The organization shall ensure that nonconforming product is accepted by concession only if the justification is provided, approval is obtained and applicable regulatory requirements are met. Records of the acceptance by concession and the identity of the person authorizing the concession shall be maintained (see 4.2.5). | /
8.3.3 Actions in response to nonconforming product detected after delivery When nonconforming product is detected after delivery or use has started, the organization shall take action appropriate to the effects, or potential effects, of the nonconformity. Records of actions taken shall be maintained (see 4.2.5). The organization shall document procedures for issuing advisory notices in accordance with applicable regulatory requirements. These procedures shall be capable of being put into effect at any time. Records of actions relating to the issuance of advisory notices shall be maintained (see 4.2.5). | 2.4 Field Corrective Action (FCA) and Field Safety Notice (FSN) “Field Corrective Action (FCA)” is an action taken by a manufacturer to reduce a risk of death or serious deterioration in the state of health associated with the use of a medical device. “Field Safety Notice (FSN)” means a communication sent out by a manufacturer or its representative to the device users in relation to a FCA. The establishment shall— (i) establish documented procedures for handling of FCA and FSN (ii) define the responsibilities for planning, conducting and reporting of corrective actions in the documented procedure; (iii) establish in writing a recall or withdrawal procedure in consultation with manufacturer (iv) inform the MDB prior to execution of FCA and FSN (v) inform all customers to whom the medical device was distributed with the appropriate degree of urgency; (vi) inform overseas counterparts on the FCA and FSN if the medical devices are exported; (vii) request that the affected medical devices be removed immediately from usable stock and stored separately in a secure area until they are disposed of accordingly; and (viii) maintain records of all actions taken in connection with the FCA and FSN. The establishment shall establish documented procedure for incident/problem reporting to comply with the regulatory requirements, which include— (i) the identification of the nature of the incident/problem; (ii)the investigation; (iii) the evaluation and analysis; and (iv) the action to be taken. Corrective actions as applicable shall be taken after final report.
2.5 Recall and Return The establishment shall— (i) establish a documented procedure to effectively and promptly recall and return medical device known or suspected to be defective; (ii) ensure that the system comply with the regulatory requirements; (iii) the manufacturer and/or authorized representative shall be informed in the event of a recall; (iv) where a recall is instituted by an entity other than the manufacturer and/or authorized representative, consultation with the manufacturer and/or authorized representative should, where possible, take place before the recall is instituted; (v) recall information shall be reported to the MDB. (vi) All returned medical devices shall be segregated apart from saleable stock to prevent redistribution until a decision has been reached regarding their disposal and treated as nonconforming product. Some criteria for medical devices to be returned to saleable stock are:- (i) The medical devices are in their original unopened containers and in good condition; (ii) It is known that the medical devices have been stored and handled under proper conditions; (iii) The remaining shelf life period is acceptable; and (iv) The medical devices have been examined and assessed by appropriate personnel. This assessment should take into account the nature of the medical device, any special storage conditions required, and the time that has elapsed since it was distributed. Special attention should be given to thermo-labile medical devices. Advice should be sought from the product owner as necessary. The returned medical devices should only be formally released to saleable stock, following a satisfactory quality re-evaluation by a nominated, responsible person. Medical devices returned to saleable stock should be placed in accordance with the stock rotation system established.
2.6 Disposal of Medical Devices The establishment shall establish a documented procedure for the disposal of medical devices, keeping in view the regulatory requirements. If the medical devices have not been immediately sent for disposal, they shall be kept in a clearly segregated area and identified so that they will not be sold inadvertently or contaminate other medical devices. Control should be established to ensure that:- (i) the status is clearly identified; (ii) the products cannot re-enter the distribution system, and; (iii) it is disposed of safely.
8.3.4 Rework The organization shall perform rework in accordance with documented procedures that takes into account the potential adverse effect of the rework on the product. These procedures shall undergo the same review and approval as the original procedure. After the completion of rework, product shall be verified to ensure that it meets applicable acceptance criteria and regulatory requirements. Records of rework shall be maintained (see 4.2.5). | /
8.4 Analysis of data The organization shall document procedures to determine, collect and analyse appropriate data to demonstrate the suitability, adequacy and effectiveness of the quality management system. The procedures shall include determination of appropriate methods, including statistical techniques and the extent of their use. The analysis of data shall include data generated as a result of monitoring and measurement and from other relevant sources and include, at a minimum, input from: a) feedback; b) conformity to product requirements; c) characteristics and trends of processes and product, including opportunities for improvement; d) suppliers; e) audits; f) service reports, as appropriate. If the analysis of data shows that the quality management system is not suitable, adequate or effective, the organization shall use this analysis as input for improvement as required in 8.5. Records of the results of analyses shall be maintained (see 4.2.5). | /
8.5 Improvement | /
8.5.1 General The organization shall identify and implement any changes necessary to ensure and maintain the continued suitability, adequacy and effectiveness of the quality management system as well as medical device safety and performance through the use of the quality policy, quality objectives, audit results, post- market surveillance, analysis of data, corrective actions, preventive actions and management review. | /
8.5.2 Corrective action The organization shall take action to eliminate the cause of nonconformities in order to prevent recurrence. Any necessary corrective actions shall be taken without undue delay. Corrective actions shall be proportionate to the effects of the nonconformities encountered. The organization shall document a procedure to define requirements for: a) reviewing nonconformities (including complaints); b) determining the causes of nonconformities; c) evaluating the need for action to ensure that nonconformities do not recur; d) planning and documenting action needed and implementing such action, including, as appropriate, updating documentation; e) verifying that the corrective action does not adversely affect the ability to meet applicable regulatory requirements or the safety and performance of the medical device; f) reviewing the effectiveness of corrective action taken. Records of the results of any investigation and of action taken shall be maintained (see 4.2.5). | 2.8 Corrective and Preventive Action For corrective actions, The establishment shall— (i) take action to eliminate the cause of nonconformities in order to comply with GDPMD and regulatory requirements; and (ii) establish a documented procedure to define requirements for: - reviewing nonconformities (including customer complaints); - determining the causes of nonconformities; - evaluating the need for action to ensure that nonconformities do not recur; - determining and implementing action needed, including, if appropriate, updating documentation, - recording of the results of any investigation and of action taken; and - reviewing the corrective action taken and its compliance with GDPMD and regulatory requirements. For preventive actions, The establishment shall— (i) determine proactive action to eliminate the causes of potential nonconformities in order to comply with GDPMD and regulatory requirements and preventive actions shall be appropriate to the effects of the potential problems; and (ii) establish a documented procedure to define requirements for— - determining potential nonconformities and their causes; - evaluating the need for action to prevent occurrence of nonconformities; - determining and implementing action needed; - recording of the results of any investigations and of action taken; - and reviewing preventive action taken and its effectiveness.
8.5.3 Preventive action The organization shall determine action to eliminate the causes of potential nonconformities in order to prevent their occurrence. Preventive actions shall be proportionate to the effects of the potential problems. The organization shall document a procedure to describe requirements for: a) determining potential nonconformities and their causes; b) evaluating the need for action to prevent occurrence of nonconformities; c) planning and documenting action needed and implementing such action, including, as appropriate, updating documentation; d) verifying that the action does not adversely affect the ability to meet applicable regulatory requirements or the safety and performance of the medical device; e) reviewing the effectiveness of the preventive action taken, as appropriate. Records of the results of any investigations and of action taken shall be maintained (see 4.2.5).
/ | SCOPE OF GDPMD CERTIFICATION The GDPMD certificate shall specify the following— (i) scope of activities performed by the establishment and category (s) of medical devices dealt by the establishment; (ii) outsourced activities, if applicable; (iii) any special storage and handling conditions, such as chill room or cold room for cold chain management; and (iv) applicable rules of the Medical Device Rules, 2015. (v) name, complete address and contact information of the establishment involved in performing activities. (vi) issuing date and expiry date of the certificate. (vii) particulars of CAB (if applicable) issuing the certificate which include the name and address, logo, registration number where applicable and/or accreditation number issued by Pakistan National Accreditation Council (PNAC) and the name and signature of the certification manager of the CAB; and (viii) number, issuing date and expiry date of the certificate. Scope of activities for establishment to be certified include anyone or combination of the following activities— (i) import; (ii) storage and warehousing; (iii)distribution. (iv)installation, testing & commissioning; (v) servicing and maintenance; and

## ISO 13485 & Saudi Req.
ISO13485: 2016 | Saudi Arabia - Requirements for Inspections and Quality Management System for Medical Devices
1 Scope
This International Standard specifies requirements for a quality management system where an
organization needs to demonstrate its ability to provide medical devices and related services that | The purpose of this document is to specify and clarify the requirements for inspection and Quality Management System for establishments and medical devices, which subject to the medical devices law and its executive regulation. In accordance to Article (33/3) of the regulations, stipulated that, “The SFDA shall issue requirements for inspection and Quality Management System for medical devices, which includes duties, authorities, obligations, and rights of the inspector”.
consistently meet customer and applicable regulatory requirements. Such organizations can be involved
in one or more stages of the life-cycle, including design and development, production, storage and | Establishments wishing to engage in any activities subject to the “Medical Devices Law” shall comply with the law and its regulation, including the requirements for inspection and Quality Management System specified in this document.
distribution, installation, or servicing of a medical device and design and development or provision of
associated activities (e.g. technical support). This International Standard can also be used by suppliers
or external parties that provide product, including quality management system-related services to such
organizations.
Requirements of this International Standard are applicable to organizations regardless of their size
and regardless of their type except where explicitly stated. Wherever requirements are specified as
applying to medical devices, the requirements apply equally to associated services as supplied by the
organization.
The processes required by this International Standard that are applicable to the organization, but
are not performed by the organization, are the responsibility of the organization and are accounted
for in the organization’s quality management system by monitoring, maintaining, and controlling the
processes.
If applicable regulatory requirements permit exclusions of design and development controls, this can
be used as a justification for their exclusion from the quality management system. These regulatory
requirements can provide alternative approaches that are to be addressed in the quality management
system. It is the responsibility of the organization to ensure that claims of conformity to this
International Standard reflect any exclusion of design and development controls.
If any requirement in Clauses 6, 7 or 8 of this International Standard is not applicable due to the
activities undertaken by the organization or the nature of the medical device for which the quality
management system is applied, the organization does not need to include such a requirement in its
quality management system. For any clause that is determined to be not applicable, the organization
records the justification as described in 4.2.2.
2 Normative references
The following documents, in whole or in part, are normatively referenced in this document and are
indispensable for its application. For dated references, only the edition cited applies. For undated
references, the latest edition of the referenced document (including any amendments) applies.
ISO 9000:20151), Quality management systems — Fundamentals and vocabulary
3 Terms and definitions
For the purposes of this document, the terms and definitions given in ISO 9000:2015 and the
following apply.
3.1 advisory notice
notice issued by the organization, subsequent to delivery of the medical device, to provide supplementary information or to advise on action to be taken in the: — use of a medical device, — modification of a medical device, — return of the medical device to the organization that supplied it, or — destruction of a medical device Note 1 to entry: Issuance of an advisory notice can be required to comply with applicable regulatory requirements.
3.2 authorized representative
natural or legal person established within a country or jurisdiction who has received a written mandate
from the manufacturer to act on his behalf for specified tasks with regard to the latter’s obligations
under that country or jurisdiction’s legislation
[SOURCE: GHTF/SG1/N055:2009, 5.2]
3.3 clinical evaluation
assessment and analysis of clinical data pertaining to a medical device to verify the clinical safety and
performance of the device when used as intended by the manufacturer
[SOURCE: GHTF/SG5/N4:2010, Clause 4]
3.4 complaint
written, electronic or oral communication that alleges deficiencies related to the identity, quality,
durability, reliability, usability, safety or performance of a medical device that has been released from
the organization’s control or related to a service that affects the performance of such medical devices
Note 1 to entry: This definition of “complaint” differs from the definition given in ISO 9000:2015.
3.5 distributor
natural or legal person in the supply chain who, on his own behalf, furthers the availability of a medical
device to the end user
Note 1 to entry: More than one distributor may be involved in the supply chain.
Note 2 to entry: Persons in the supply chain involved in activities such as storage and transport on behalf of the
manufacturer, importer or distributor, are not distributors under this definition.
[SOURCE: GHTF/SG1/N055:2009, 5.3]
3.6 implantable medical device
medical device which can only be removed by medical or surgical intervention and which is intended to:
— be totally or partially introduced into the human body or a natural orifice, or
— replace an epithelial surface or the surface of the eye, and
— remain after the procedure for at least 30 days
Note 1 to entry: This definition of implantable medical device includes active implantable medical device
3.7 importer
natural or legal person in the supply chain who is the first in a supply chain to make a medical device,
manufactured in another country or jurisdiction, available in the country or jurisdiction where it is to
be marketed
[SOURCE: GHTF/SG1/N055:2009, 5.4]
3.8 labelling
label, instructions for use, and any other information that is related to identification, technical
description, intended purpose and proper use of the medical device, but excluding shipping documents
[SOURCE: GHTF/SG1/N70:2011, Clause 4]
3.9 life-cycle
all phases in the life of a medical device, from the initial conception to final decommissioning and
disposal
[SOURCE: ISO 14971:2007, 2.7]
3.1
manufacturer
natural or legal person with responsibility for design and/or manufacture of a medical device with the
intention of making the medical device available for use, under his name; whether or not such a medical
device is designed and/or manufactured by that person himself or on his behalf by another person(s)
Note 1 to entry: This “natural or legal person” has ultimate legal responsibility for ensuring compliance with
all applicable regulatory requirements for the medical devices in the countries or jurisdictions where it is
intended to be made available or sold, unless this responsibility is specifically imposed on another person by the
Regulatory Authority (RA) within that jurisdiction.
Note 2 to entry: The manufacturer’s responsibilities are described in other GHTF guidance documents. These
responsibilities include meeting both pre-market requirements and post-market requirements, such as adverse
event reporting and notification of corrective actions.
Note 3 to entry: “Design and/or manufacture”, as referred to in the above definition, may include specification
development, production, fabrication, assembly, processing, packaging, repackaging, labelling, relabelling,
sterilization, installation, or remanufacturing of a medical device; or putting a collection of devices, and possibly
other products, together for a medical purpose.
Note 4 to entry: Any person who assembles or adapts a medical device that has already been supplied by another
person for an individual patient, in accordance with the instructions for use, is not the manufacturer, provided
the assembly or adaptation does not change the intended use of the medical device.
Note 5 to entry: Any person who changes the intended use of, or modifies, a medical device without acting on
behalf of the original manufacturer and who makes it available for use under his own name, should be considered
the manufacturer of the modified medical device.
Note 6 to entry: An authorized representative, distributor or importer who only adds its own address and
contact details to the medical device or the packaging, without covering or changing the existing labelling, is not
considered a manufacturer.
Note 7 to entry: To the extent that an accessory is subject to the regulatory requirements of a medical device, the
person responsible for the design and/or manufacture of that accessory is considered to be a manufacturer.
[SOURCE: GHTF/SG1/N055:2009, 5.1]
3.11
medical device
instrument, apparatus, implement, machine, appliance, implant, reagent for in vitro use, software,
material or other similar or related article, intended by the manufacturer to be used, alone or in
combination, for human beings, for one or more of the specific medical purpose(s) of:
— diagnosis, prevention, monitoring, treatment or alleviation of disease;
— diagnosis, monitoring, treatment, alleviation of or compensation for an injury;
— investigation, replacement, modification, or support of the anatomy or of a physiological process;
— supporting or sustaining life;
— control of conception;
— disinfection of medical devices;
— providing information by means of in vitro examination of specimens derived from the human body;
and does not achieve its primary intended action by pharmacological, immunological or metabolic
means, in or on the human body, but which may be assisted in its intended function by such means
Note 1 to entry: Products which may be considered to be medical devices in some jurisdictions but not in others
include:
— disinfection substances;
— aids for persons with disabilities;
— devices incorporating animal and/or human tissues;
— devices for in vitro fertilization or assisted reproduction technologies.
[SOURCE: GHTF/SG1/N071:2012, 5.1]
3.12
medical device family
group of medical devices manufactured by or for the same organization and having the same basic
design and performance characteristics related to safety, intended use and function
3.13
performance evaluation
assessment and analysis of data to establish or verify the ability of an in vitro diagnostic medical device
to achieve its intended use
3.14
post-market surveillance
systematic process to collect and analyse experience gained from medical devices that have been placed
on the market
3.15
product
result of a process
Note 1 to entry: There are four generic product categories, as follows:
— services (e.g. transport);
— software (e.g. computer program, dictionary);
— hardware (e.g. engine mechanical part);
— processed materials (e.g. lubricant).
Many products comprise elements belonging to different generic product categories. Whether the product is
then called service, software, hardware or processed material depends on the dominant element. For example,
the offered product “automobile” consists of hardware (e.g. tyres), processed materials (e.g. fuel, cooling liquid),
software (e.g. engine control software, driver’s manual), and service (e.g. operating explanations given by the
salesman).
Note 2 to entry: Service is the result of at least one activity necessarily performed at the interface between the
supplier and customer and is generally intangible. Provision of a service can involve, for example, the following:
— an activity performed on a customer-supplied tangible product (e.g. automobile to be repaired);
— an activity performed on a customer-supplied intangible product (e.g. the income statement needed to
prepare a tax return);
— the delivery of an intangible product (e.g. the delivery of information in the context of knowledge
transmission);
— the creation of ambience for the customer (e.g. in hotels and restaurants).
Software consists of information and is generally intangible and can be in the form of approaches, transactions
or procedures.
Hardware is generally tangible and its amount is a countable characteristic. Processed materials are generally
tangible and their amount is a continuous characteristic. Hardware and processed materials often are referred
to as goods.
Note 3 to entry: This definition of “product” differs from the definition given in ISO 9000:2015.
[SOURCE: ISO 9000:20052), 3.4.2, modified]
3.16
purchased product
product provided by a party outside the organization’s quality management system
Note 1 to entry: The provision of product does not necessarily infer a commercial or financial arrangement.
3.17
risk
combination of the probability of occurrence of harm and the severity of that harm
Note 1 to entry: This definition of “risk” differs from the definition given in ISO 9000:2015.
[SOURCE: ISO 14971:2007, 2.16]
3.18
risk management
systematic application of management policies, procedures and practices to the tasks of analysing,
evaluating, controlling and monitoring risk
[SOURCE: ISO 14971:2007, 2.22]
3.19
sterile barrier system
minimum package that prevents ingress of microorganisms and allows aseptic presentation of the
product at the point of use
[SOURCE: ISO 11607-1:2006, 3.22]
2) Superseded by ISO 9000:2015.
3.2
sterile medical device
medical device intended to meet the requirements for sterility
Note 1 to entry: The requirements for sterility of a medical device can be subject to applicable regulatory
requirements or standards.
4 Quality management system
4.1 General requirements
4.1.1 The organization shall document a quality management system and maintain its effectiveness in accordance with the requirements of this International Standard and applicable regulatory requirements. | SFDA has issued this document in accordance to “Medical Devices Law” published by the Royal Decree No. (M/54) dated 6/7/1442H in reference to the following: - Article (22), which stipulates that “Establishments seeking to circulate medical devices in the Kingdom shall adhere to the Quality Management System”.
The organization shall establish, implement and maintain any requirement, procedure, activity or arrangement required to be documented by this International Standard or applicable regulatory requirements.
The organization shall document the role(s) undertaken by the organization under the applicable regulatory requirements.
NOTE Roles undertaken by the organization can include manufacturer, authorized representative, importer or distributor.
4.1.2 The organization shall:
a) determine the processes needed for the quality management system and the application of these processes throughout the organization taking into account the roles undertaken by the organization;
b) apply a risk based approach to the control of the appropriate processes needed for the quality management system;
c) determine the sequence and interaction of these processes.
4.1.3 For each quality management system process, the organization shall:
a) determine criteria and methods needed to ensure that both the operation and control of these processes are effective;
b) ensure the availability of resources and information necessary to support the operation and monitoring of these processes;
c) implement actions necessary to achieve planned results and maintain the effectiveness of these processes;
d) monitor, measure as appropriate, and analyse these processes;
e) establish and maintain records needed to demonstrate conformance to this International Standard and compliance with applicable regulatory requirements (see 4.2.5).
4.1.4 The organization shall manage these quality management system processes in accordance with the requirements of this International Standard and applicable regulatory requirements. Changes to be made to these processes shall be:
a) evaluated for their impact on the quality management system;
b) evaluated for their impact on the medical devices produced under this quality management system;
c) controlled in accordance with the requirements of this International Standard and applicable regulatory requirements.
4.1.5 When the organization chooses to outsource any process that affects product conformity to requirements, it shall monitor and ensure control over such processes. The organization shall retain responsibility of conformity to this International Standard and to customer and applicable regulatory requirements for outsourced processes. The controls shall be proportionate to the risk involved and the ability of the external party to meet the requirements in accordance with 7.4. The controls shall include written quality agreements.
4.1.6 The organization shall document procedures for the validation of the application of computer software used in the quality management system. Such software applications shall be validated prior to initial use and, as appropriate, after changes to such software or its application.
The specific approach and activities associated with software validation and revalidation shall be proportionate to the risk associated with the use of the software.
Records of such activities shall be maintained (see 4.2.5).
4.2 Documentation requirements
4.2.1 General
The quality management system documentation (see 4.2.4) shall include: a) documented statements of a quality policy and quality objectives; b) a quality manual; c) documented procedures and records required by this International Standard; d) documents, including records, determined by the organization to be necessary to ensure the effective planning, operation, and control of its processes; e) other documentation specified by applicable regulatory requirements.
4.2.2 Quality manual
The organization shall document a quality manual that includes:
a) the scope of the quality management system, including details of and justification for any exclusion or non-application;
b) the documented procedures for the quality management system, or reference to them;
c) a description of the interaction between the processes of the quality management system.
The quality manual shall outline the structure of the documentation used in the quality management system.
4.2.3 | B.Particular Requirements First: Medical Devices Manufacturers 1.5. Maintain and provide the following documents upon request: 1.5.1. The identifying information (labeling) in English, as well as in Arabic if the user is a lay person. 1.5.2. Instructions related to the handling, transportation, storage, installation, maintenance and disposal of medical devices in English, as well as in Arabic if the user is a lay person. 1.5.3. Advertising, promotional and marketing information in English, as well as in Arabic if the user is a lay person. 1.5.4. Description of the medical device and its specifications, including differences and accessories. 1.5.5. Design and manufacturing information. 1.5.6. Risk management file. 1.5.7. Verification and validation of the product including clinical trials (studies). 1.5.8. Post-marketing surveillance plan. 1.5.9. Post-marketing surveillance report and periodic safety update report. 1.5.10. Written procedures to take the necessary field safety corrective actions mentioned in the safety notice.
4.2.4 Control of documents
Documents required by the quality management system shall be controlled. Records are a special type of document and shall be controlled according to the requirements given in 4.2.5. A documented procedure shall define the controls needed to: a) review and approve documents for adequacy prior to issue; b) review, update as necessary and re-approve documents; c) ensure that the current revision status of and changes to documents are identified; d) ensure that relevant versions of applicable documents are available at points of use; e) ensure that documents remain legible and readily identifiable; f) ensure that documents of external origin, determined by the organization to be necessary for the planning and operation of the quality management system, are identified and their distribution controlled; g) prevent deterioration or loss of documents; h) prevent the unintended use of obsolete documents and apply suitable identification to them. The organization shall ensure that changes to documents are reviewed and approved either by the original approving function or another designated function that has access to pertinent background information upon which to base its decisions. The organization shall define the period for which at least one copy of obsolete documents shall be retained. This period shall ensure that documents to which medical devices have been manufactured and tested are available for at least the lifetime of the medical device as defined by the organization, but not less than the retention period of any resulting record (see 4.2.5), or as specified by applicable regulatory requirements.
4.2.5 Control of records
Records shall be maintained to provide evidence of conformity to requirements and of the effective operation of the quality management system. The organization shall document procedures to define the controls needed for the identification, storage, security and integrity, retrieval, retention time and disposition of records. The organization shall define and implement methods for protecting confidential health information contained in records in accordance with the applicable regulatory requirements. Records shall remain legible, readily identifiable and retrievable. Changes to a record shall remain identifiable. The organization shall retain the records for at least the lifetime of the medical device as defined by the organization, or as specified by applicable regulatory requirements, but not less than two years from the medical device release by the organization.Records shall be maintained to provide evidence of conformity to requirements and of the effective | A.General Requirements A.15.Maintaining records of medical devices disposal in accordance with the requirements for post-marketing surveillance for medical devices published on the SFDA’s website. A.19 Maintaining records of medical devices disposal in accordance with the requirements for post-marketing surveillance for medical devices published on the SFDA’s website.
5 Management responsibility
5.1 Management commitment | B.Particular Requirements First: Medical Devices Manufacturers 1.3. A pledge that the manufacturer bears full responsibility for the quality of all manufactured batches
Top management shall provide evidence of its commitment to the development and implementation of the quality management system and maintenance of its effectiveness by:
a) communicating to the organization the importance of meeting customer as well as applicable regulatory requirements;
b) establishing the quality policy;
c) ensuring that quality objectives are established;
d) conducting management reviews;
e) ensuring the availability of resources.
5.2 Customer focus
Top management shall ensure that customer requirements and applicable regulatory requirements are determined and met.
5.3 Quality policy
Top management shall ensure that the quality policy:
a) is applicable to the purpose of the organization;
b) includes a commitment to comply with requirements and to maintain the effectiveness of the quality management system;
c) provides a framework for establishing and reviewing quality objectives;
d) is communicated and understood within the organization;
e) is reviewed for continuing suitability.
5.4 Planning
5.4.1 Quality objectives
Top management shall ensure that quality objectives, including those needed to meet applicable regulatory requirements and requirements for product, are established at relevant functions and levels within the organization. The quality objectives shall be measurable and consistent with the quality policy.
5.4.2 Quality management system planning
Top management shall ensure that:
a) the planning of the quality management system is carried out in order to meet the requirements given in 4.1, as well as the quality objectives;
b) the integrity of the quality management system is maintained when changes to the quality management system are planned and implemented.
5.5 Responsibility, authority and communication
5.5.1 Responsibility and authority
Top management shall ensure that responsibilities and authorities are defined, documented and communicated within the organization.
Top management shall document the interrelation of all personnel who manage, perform and verify work affecting quality and shall ensure the independence and authority necessary to perform these tasks.
5.5.2 Management representative
Top management shall appoint a member of management who, irrespective of other responsibilities,has responsibility and authority that includes:
a) ensuring that processes needed for the quality management system are documented;
b) reporting to top management on the effectiveness of the quality management system and any need for improvement;
c) ensuring the promotion of awareness of applicable regulatory requirements and quality management system requirements throughout the organization.
5.5.3 Internal communication
Top management shall ensure that appropriate communication processes are established within the organization and that communication takes place regarding the effectiveness of the quality management system.
5.6 Management review
5.6.1 General
The organization shall document procedures for management review. Top management shall review the organization’s quality management system at documented planned intervals to ensure its continuing suitability, adequacy and effectiveness. The review shall include assessing opportunities for improvement and the need for changes to the quality management system, including the quality policy and quality objectives.
Records from management reviews shall be maintained (see 4.2.5).
5.6.2 Review input
The input to management review shall include, but is not limited to, information arising from:
a) feedback;
b) complaint handling;
c) reporting to regulatory authorities;
d) audits;
e) monitoring and measurement of processes;
f) monitoring and measurement of product;
g) corrective action;
h) preventive action;
i) follow-up actions from previous management reviews;
j) changes that could affect the quality management system;
k) recommendations for improvement;
l) applicable new or revised regulatory requirements.
5.6.3 Review output
The output from management review shall be recorded (see 4.2.5) and include the input reviewed and any decisions and actions related to:
a) improvement needed to maintain the suitability, adequacy, and effectiveness of the quality management system and its processes; b) improvement of product related to customer requirements; c) changes needed to respond to applicable new or revised regulatory requirements; d) resource needs.
6 Resource management
6.1 Provision of resources | A.9.Providing a sufficient and appropriate human resources and other resources shall be provided according to the activity of the establishment and the tasks assigned to it to fulfill all the requirements of the Law and its regulation.
The organization shall determine and provide the resources needed to:
a) implement the quality management system and to maintain its effectiveness;
b) meet applicable regulatory and customer requirements.
6.2 Human resources
Personnel performing work affecting product quality shall be competent on the basis of appropriate education, training, skills and experience. | B.Particular Requirements First: Medical Devices Manufacturers 1.2 The presence of a full-time technical manager and a full-time quality manager who are medical devices engineers/technicians, or qualified in one of the related specialties.
The organization shall document the process(es) for establishing competence, providing needed training, and ensuring awareness of personnel.
The organization shall: a) determine the necessary competence for personnel performing work affecting product quality; b) provide training or take other actions to achieve or maintain the necessary competence; c) evaluate the effectiveness of the actions taken; d) ensure that its personnel are aware of the relevance and importance of their activities and how they contribute to the achievement of the quality objectives; e) maintain appropriate records of education, training, skills and experience (see 4.2.4). NOTE The methodology used to check effectiveness is proportionate to the risk associated with the work for which the training or other action is being provided.
6.3 Infrastructure
The organization shall document the requirements for the infrastructure needed to achieve conformity to product requirements, prevent product mix-up and ensure orderly handling of product.Infrastructure includes, as appropriate: | B.Particular Requirements First: Medical Devices Manufacturers 1.1 The design of the manufacturer’s facility compatible with the nature of the medical device to be manufactured.
a) buildings, workspace and associated utilities;
b) process equipment (both hardware and software);
c) supporting services (such as transport, communication, or information systems).
The organization shall document requirements for the maintenance activities, including the interval of performing the maintenance activities, when such maintenance activities, or lack thereof, can affect product quality. As appropriate, the requirements shall apply to equipment used in production, the control of the work environment and monitoring and measurement. | A. 20. Disclose about any other unlicensed branches or warehouses of the establishment.
Records of such maintenance shall be maintained (see 4.2.5).
6.4 Work environment and contamination control
6.4.1 Work environment
The organization shall document the requirements for the work environment needed to achieve conformity to product requirements.
If the conditions for the work environment can have an adverse effect on product quality, the organization shall document the requirements for the work environment and the procedures to monitor and control the work environment.
The organization shall:
a) document requirements for health, cleanliness and clothing of personnel if contact between such personnel and the product or work environment could affect medical device safety or performance;
b) ensure that all personnel who are required to work temporarily under special environmental conditions within the work environment are competent or supervised by a competent person.
NOTE Further information can be found in ISO 14644 and ISO 14698.
6.4.2 Contamination control
As appropriate, the organization shall plan and document arrangements for the control of contaminated or potentially contaminated product in order to prevent contamination of the work environment, personnel, or product.
For sterile medical devices, the organization shall document requirements for control of contamination with microorganisms or particulate matter and maintain the required cleanliness during assembly or packaging processes.
/
/
/
/
7.2 Customer-related processes | B.Particular Requirements Second: Distributors, Importers and Authorized Representatives Medical devices categorized as high risk may not be dispensed for use outside the facility of the healthcare provider without prescription, in accordance with the list of high-risk medical devices published on the SFDA’s website, while keeping its records for a period of no less than (5) years.
7.2.1 Determination of requirements related to product
The organization shall determine:
a) requirements specified by the customer, including the requirements for delivery and post-delivery activities;
b) requirements not stated by the customer but necessary for specified or intended use, as known;
c) applicable regulatory requirements related to the product;
d) any user training needed to ensure specified performance and safe use of the medical device;
e) any additional requirements determined by the organization.
7.2.2 Review of requirements related to product
The organization shall review the requirements related to product. This review shall be conducted prior to the organization’s commitment to supply product to the customer (e.g. submission of tenders, acceptance of contracts or orders, acceptance of changes to contracts or orders) and shall ensure that: | A. 13. Follow up on the expiry date of medical devices.
a) product requirements are defined and documented;
b) contract or order requirements differing from those previously expressed are resolved;
c) applicable regulatory requirements are met;
d) any user training identified in accordance with 7.2.1 is available or planned to be available;
e) the organization has the ability to meet the defined requirements.
Records of the results of the review and actions arising from the review shall be maintained (see 4.2.5).
When the customer provides no documented statement of requirement, the customer requirements shall be confirmed by the organization before acceptance.
When product requirements are changed, the organization shall ensure that relevant documents are amended and that relevant personnel are made aware of the changed requirements.
7.2.3 Communication
The organization shall plan and document arrangements for communicating with customers in relation to: | B.Particular Requirements Second: Distributors, Importers and Authorized Representatives Individuals who are concerned with marketing and selling medical devices shall have sufficient information about the devices in order to provide the correct information related to its marketing and sale.
a) product information;
b) enquiries, contracts or order handling, including amendments;
c) customer feedback, including complaints;
d) advisory notices.
The organization shall communicate with regulatory authorities in accordance with applicable regulatory requirements.
7.4 Purchasing
7.4.1 Purchasing process
The organization shall document procedures (see 4.2.4) to ensure that purchased product conforms to specified purchasing information.
The organization shall establish criteria for the evaluation and selection of suppliers. The criteria shall be:
a) based on the supplier’s ability to provide product that meets the organization’s requirements;
b) based on the performance of the supplier;
c) based on the effect of the purchased product on the quality of the medical device;
d) proportionate to the risk associated with the medical device.
The organization shall plan the monitoring and re-evaluation of suppliers. Supplier performance in meeting requirements for the purchased product shall be monitored. The results of the monitoring shall provide an input into the supplier re-evaluation process.
Non-fulfilment of purchasing requirements shall be addressed with the supplier proportionate to the risk associated with the purchased product and compliance with applicable regulatory requirements.
Records of the results of evaluation, selection, monitoring and re-evaluation of supplier capability or performance and any necessary actions arising from these activities shall be maintained (see 4.2.5).
7.4.2 Purchasing information
Purchasing information shall describe or reference the product to be purchased, including as appropriate:
a) product specifications;
b) requirements for product acceptance, procedures, processes and equipment;
c) requirements for qualification of supplier personnel;
d) quality management system requirements.
The organization shall ensure the adequacy of specified purchasing requirements prior to their communication to the supplier.
Purchasing information shall include, as applicable, a written agreement that the supplier notify the organization of changes in the purchased product prior to implementation of any changes that affect the ability of the purchased product to meet specified purchase requirements.
To the extent required for traceability given in 7.5.9, the organization shall maintain relevant purchasing information in the form of documents (see 4.2.4) and records (see 4.2.5).
7.4.3 Verification of purchased product
The organization shall establish and implement the inspection or other activities necessary for ensuring that purchased product meets specified purchasing requirements. The extent of verification activities shall be based on the supplier evaluation results and proportionate to the risks associated with the purchased product.
When the organization becomes aware of any changes to the purchased product, the organization shall determine whether these changes affect the product realization process or the medical device.
When the organization or its customer intends to perform verification at the supplier’s premises, the organization shall state the intended verification activities and method of product release in the the organization shall state the intended verification activities and method of product release in the purchasing information.
Records of the verification shall be maintained (see 4.2.5).
7.5 Production and service provision
7.5.3 Installation activities | B.Particular Requirements Second: Distributors, Importers and Authorized Representatives 1.4. If the establishment wishes to provide maintenance services for its medical devices, it must adhere to the manufacturer’s instructions and the requirements for maintenance services of medical devices contained in the requirements for post-marketing surveillance for medical devices published on the SFDA’s website. In the event that the establishment wishes to provide maintenance services for medical devices that are not affiliated with it, license for providing maintenance services for medical devices shall be obtained in accordance with the requirements for licensing medical devices establishments.
The organization shall document requirements for medical device installation and acceptance criteria for verification of installation, as appropriate.
If the agreed customer requirements allow installation of the medical device to be performed by an external party other than the organization or its supplier, the organization shall provide documented requirements for medical device installation and verification of installation.
Records of medical device installation and verification of installation performed by the organization or its supplier shall be maintained (see 4.2.5).
7.5.4 Servicing activities
If servicing of the medical device is a specified requirement, the organization shall document servicing procedures, reference materials, and reference measurements, as necessary, for performing servicing activities and verifying that product requirements are met. | B.Particular Requirements First: Medical Devices Manufacturers 1.7. Provide after-sales services in accordance with Articles (16/1) and (16/2) of the Regulation.
The organization shall analyse records of servicing activities carried out by the organization or its supplier:
a) to determine if the information is to be handled as a complaint;
b) as appropriate, for input to the improvement process.
Records of servicing activities carried out by the organization or its supplier shall be maintained(see 4.2.5).
7.5.8 Identification
The organization shall document procedures for product identification and identify product by suitable means throughout product realization.
The organization shall identify product status with respect to monitoring and measurement requirements throughout product realization. Identification of product status shall be maintained throughout production, storage, installation and servicing of product to ensure that only product that has passed the required inspections and tests or released under an authorized concession is dispatched, used or installed.
If required by applicable regulatory requirements, the organization shall document a system to assign unique device identification to the medical device. | B.Particular Requirements First: Medical Devices Manufacturers Unique Device Identification (UDI) requirements for medical devices
The organization shall document procedures to ensure that medical devices returned to the organization are identified and distinguished from conforming product.
/
7.5.9 Traceability
7.5.9.1 General
The organization shall document procedures for traceability. These procedures shall define the extent | A. 11.Providing and implementing effective and appropriate tracking procedures for documenting, distribution and use of medical devices, which help in tracking the medical device according to the Requirements for Unique Device Identification (UDI) for Medical Devices.
of traceability in accordance with applicable regulatory requirements and the records to be maintained(see 4.2.5).
7.5.9.2 Particular requirements for implantable medical devices | B.Particular Requirements First: Medical Devices Manufacturers 1.8. The implanted medical device card shall contain the following information: 1.8.1. Information that allow identification of the device, including device name, serial number, batch number, UDI data, device model, name, address and website of the manufacturer. 1.8.2. Any warnings, precautions, or measures to be taken by the patient or healthcare provider regarding mutual interference with foreseeable external influences, medical examinations or anticipated environmental conditions. 1.8.3. Any information about the expected lifetime of the device and any necessary follow-up of the device's performance. 1.8.4. Any other information to ensure the safe use of the device.
The records required for traceability shall include records of components, materials, and conditions for the work environment used, if these could cause the medical device not to satisfy its specified safety and performance requirements.
The organization shall require that suppliers of distribution services or distributors maintain records of the distribution of medical devices to allow traceability and that these records are available for inspection.
Records of the name and address of the shipping package consignee shall be maintained (see 4.2.5).
7.5.10 Customer property
The organization shall identify, verify, protect, and safeguard customer property provided for use or incorporation into the product while it is under the organization’s control or being used by the organization. If any customer property is lost, damaged or otherwise found to be unsuitable for use, the organization shall report this to the customer and maintain records (see 4.2.5).
7.5.11 Preservation of product
The organization shall document procedures for preserving the conformity of product to requirements during processing, storage, handling, and distribution. Preservation shall apply to the constituent parts of a medical device. | A. 12. Providing and implementing effective and appropriate procedures for storage and transportation according to the requirements of the manufacturer and the requirements for transportation and storage of medical devices published on the SFDA’s website. B.Particular Requirements Second: Distributors, Importers and Authorized Representatives 1.3. Acquiring a warehouse license or a third-party storage license issued by the SFDA in accordance with the requirements for licensing medical devices establishments. As for retailers, a storage area inside the establishment can be sufficient with adherence to the requirements of transportation and storage for medical devices.
The organization shall protect product from alteration, contamination or damage when exposed to expected conditions and hazards during processing, storage, handling, and distribution by:
a) designing and constructing suitable packaging and shipping containers;
b) documenting requirements for special conditions needed if packaging alone cannot provide preservation.
If special conditions are required, they shall be controlled and recorded (see 4.2.5).
7.6 Control of monitoring and measuring equipment
The organization shall determine the monitoring and measurement to be undertaken and the monitoring and measuring equipment needed to provide evidence of conformity of product to determined requirements.
The organization shall document procedures to ensure that monitoring and measurement can be carried out and are carried out in a manner that is consistent with the monitoring and measurement requirements.
As necessary to ensure valid results, measuring equipment shall:
a) be calibrated or verified, or both, at specified intervals, or prior to use, against measurement standards traceable to international or national measurement standards: when no such standards exist, the basis used for calibration or verification shall be recorded (see 4.2.5);
b) be adjusted or re-adjusted as necessary: such adjustments or re-adjustments shall be recorded(see 4.2.5);
c) have identification in order to determine its calibration status;
d) be safeguarded from adjustments that would invalidate the measurement result;
e) be protected from damage and deterioration during handling, maintenance and storage.
The organization shall perform calibration or verification in accordance with documented procedures.
In addition, the organization shall assess and record the validity of the previous measuring results when the equipment is found not to conform to requirements. The organization shall take appropriate action in regard to the equipment and any product affected.
Records of the results of calibration and verification shall be maintained (see 4.2.5).
The organization shall document procedures for the validation of the application of computer software used for the monitoring and measurement of requirements. Such software applications shall be validated prior to initial use and, as appropriate, after changes to such software or its application.
The specific approach and activities associated with software validation and revalidation shall be proportionate to the risk associated with the use of the software, including the effect on the ability of the product to conform to specifications.
Records of the results and conclusion of validation and necessary actions from the validation shall be maintained (see 4.2.4 and 4.2.5).
NOTE Further information can be found in ISO 10012.
/
8 Measurement, analysis and improvement
8.1 General
The organization shall plan and implement the monitoring, measurement, analysis and improvement processes needed to:
a) demonstrate conformity of product;
b) ensure conformity of the quality management system;
c) maintain the effectiveness of the quality management system.
This shall include determination of appropriate methods, including statistical techniques, and the extent of their use.
8.2 Monitoring and measurement
8.2.1 Feedback
As one of the measurements of the effectiveness of the quality management system, the organization shall gather and monitor information relating to whether the organization has met customer requirements. The methods for obtaining and using this information shall be documented.
The organization shall document procedures for the feedback process. This feedback process shall include provisions to gather data from production as well as post-production activities.
The information gathered in the feedback process shall serve as potential input into risk management for monitoring and maintaining the product requirements as well as the product realization or improvement processes.
If applicable regulatory requirements require the organization to gain specific experience from postproduction activities, the review of this experience shall form part of the feedback process.
8.2.2 Complaint handling
The organization shall document procedures for timely complaint handling in accordance with applicable regulatory requirements.
These procedures shall include at a minimum requirements and responsibilities for: a) receiving and recording information; b) evaluating information to determine if the feedback constitutes a complaint; c) investigating complaints; d) determining the need to report the information to the appropriate regulatory authorities; e) handling of complaint-related product; f) determining the need to initiate corrections or corrective actions.
If any complaint is not investigated, justification shall be documented. Any correction or corrective action resulting from the complaint handling process shall be documented.
If an investigation determines activities outside the organization contributed to the complaint, relevant information shall be exchanged between the organization and the external party involved.
Complaint handling records shall be maintained (see 4.2.5).
8.2.3 Reporting to regulatory authorities
If applicable regulatory requirements require notification of complaints that meet specified reporting criteria of adverse events or issuance of advisory notices, the organization shall document procedures for providing notification to the appropriate regulatory authorities. | B.Particular Requirements Second: Distributors, Importers and Authorized Representatives 2.3. Report to the National Center for Medical Devices Reporting (NCMDR) about accidents of its medical devices.
Records of reporting to regulatory authorities shall be maintained (see 4.2.5).
8.2.4 Internal audit
The organization shall conduct internal audits at planned intervals to determine whether the quality management system:
a) conforms to planned and documented arrangements, requirements of this International Standard, quality management system requirements established by the organization, and applicable regulatory requirements;
b) is effectively implemented and maintained.
The organization shall document a procedure to describe the responsibilities and requirements for planning and conducting audits and recording and reporting audit results.
An audit program shall be planned, taking into consideration the status and importance of the processes and area to be audited, as well as the results of previous audits. The audit criteria, scope, interval and methods shall be defined and recorded (see 4.2.5). The selection of auditors and conduct of audits shall ensure objectivity and impartiality of the audit process. Auditors shall not audit their own work.
Records of the audits and their results, including identification of the processes and areas audited and the conclusions, shall be maintained (see 4.2.5).
The management responsible for the area being audited shall ensure that any necessary corrections and corrective actions are taken without undue delay to eliminate detected nonconformities and their causes. Follow-up activities shall include the verification of the actions taken and the reporting of verification results.
NOTE Further information can be found in ISO 19011.
8.2.5 Monitoring and measurement of processes
The organization shall apply suitable methods for monitoring and, as appropriate, measurement of the quality management system processes. These methods shall demonstrate the ability of the processes to achieve planned results. When planned results are not achieved, correction and corrective action shall be taken, as appropriate.
8.2.6 Monitoring and measurement of product
The organization shall monitor and measure the characteristics of the product to verify that product requirements have been met. This shall be carried out at applicable stages of the product realization process in accordance with the planned and documented arrangements and documented procedures.
Evidence of conformity to the acceptance criteria shall be maintained. The identity of the person authorizing release of product shall be recorded (see 4.2.5). As appropriate, records shall identify the test equipment used to perform measurement activities.
Product release and service delivery shall not proceed until the planned and documented arrangements have been satisfactorily completed.
For implantable medical devices, the organization shall record the identity of personnel performing any inspection or testing.
/
8.3 Control of nonconforming product
8.3.1 General
The organization shall ensure that product which does not conform to product requirements is identified and controlled to prevent its unintended use or delivery. The organization shall document a procedure to define the controls and related responsibilities and authorities for the identification, documentation, segregation, evaluation and disposition of nonconforming product. | A. General Requirements 23. Adhere to the following after the inspection: a. Signing/acknowledgement of receipt the inspection report. b. Conducting a comprehensive investigation if requested by the SFDA in the events of non-conformity mentioned in the inspection report. c. Submit a corrective plan, if requested by the SFDA, to address cases of non-conformity - if any - within the period specified in the request. d. Implementation of the corrective plan after its acceptance by the SFDA within the specified period. e. Notify SFDA of the completion of the implementation of the corrective plan.
The evaluation of nonconformity shall include a determination of the need for an investigation and notification of any external party responsible for the nonconformity.
Records of the nature of the nonconformities and any subsequent action taken, including the evaluation, any investigation and the rationale for decisions shall be maintained (see 4.2.5)
8.3.2 Actions in response to nonconforming product detected before delivery
The organization shall deal with nonconforming product by one or more of the following ways: | B.Particular Requirements First: Medical Devices Manufacturers 6. After the visit, the manufacturer shall comply with the following, in addition to what was mentioned in Paragraph No. (23) of the “General Requirements” section: 6.1. Respond to cases of non-conformity and submit the corrective plan using the form sent with the inspection report, within the specified period, provided that the plan shall include the following: 6.1.1. Respond to all cases of non-conformity detected in the inspection report. 6.1.2. Clarify the root cause of the non-conformities. 6.1.3. Procedures for immediate correction of non-conformities. 6.1.4. Corrective actions for non-conformities. 6.1.5. Preventive procedures for non-conformities. 6.1.6. Date of implementation of the corrective plan.
a) taking action to eliminate the detected nonconformity;
b) taking action to preclude its original intended use or application;
c) authorizing its use, release or acceptance under concession.
The organization shall ensure that nonconforming product is accepted by concession only if the justification is provided, approval is obtained and applicable regulatory requirements are met. Records of the acceptance by concession and the identity of the person authorizing the concession shall be maintained (see 4.2.5).
8.3.3 Actions in response to nonconforming product detected after delivery
When nonconforming product is detected after delivery or use has started, the organization shall take action appropriate to the effects, or potential effects, of the nonconformity. Records of actions taken shall be maintained (see 4.2.5). | A.General Requirements 14. Compliance with the conditions and requirements of destruction, reprocessing, refurbishing, reselling or donating in accordance with the requirements for post-marketing surveillance for medical devices published on the SFDA’s website.
The organization shall document procedures for issuing advisory notices in accordance with applicable regulatory requirements. These procedures shall be capable of being put into effect at any time. Records of actions relating to the issuance of advisory notices shall be maintained (see 4.2.5).
8.4 Analysis of data
The organization shall document procedures to determine, collect and analyse appropriate data to demonstrate the suitability, adequacy and effectiveness of the quality management system. The procedures shall include determination of appropriate methods, including statistical techniques and the extent of their use. | A. General Requirements 17. Process all observations contained in previous inspection reports.
The analysis of data shall include data generated as a result of monitoring and measurement and from other relevant sources and include, at a minimum, input from:
a) feedback;
b) conformity to product requirements;
c) characteristics and trends of processes and product, including opportunities for improvement;
d) suppliers;
e) audits;
f) service reports, as appropriate.
If the analysis of data shows that the quality management system is not suitable, adequate or effective, the organization shall use this analysis as input for improvement as required in 8.5.
Records of the results of analyses shall be maintained (see 4.2.5).
8.5 Improvement
8.5.1 General
The organization shall identify and implement any changes necessary to ensure and maintain the continued suitability, adequacy and effectiveness of the quality management system as well as medical device safety and performance through the use of the quality policy, quality objectives, audit results, postmarket surveillance, analysis of data, corrective actions, preventive actions and management review.
8.5.2 Corrective action
The organization shall take action to eliminate the cause of nonconformities in order to prevent recurrence. Any necessary corrective actions shall be taken without undue delay. Corrective actions shall be proportionate to the effects of the nonconformities encountered. | A. General Requirements 16. In the event that a product was seized in violation of the Law, the seizure of the medical device will continue until the completion of corrective action if the product is capable of being corrected. B.Particular Requirements First: Medical Devices Manufacturers 6.2. The date of implementation of the corrective plan shall be proportional to the nature of non-conformities. 6.3. The SFDA shall evaluates the corrective plan submitted by the manufacturer, and respond to the manufacture with the evaluation result, either acceptance or not, with returning the form to the manufacture to amend the corrective plan. 6.4. Submit the amended corrective plan within the period specified in the request.
The organization shall document a procedure to define requirements for:
a) reviewing nonconformities (including complaints);
b) determining the causes of nonconformities;
c) evaluating the need for action to ensure that nonconformities do not recur;
d) planning and documenting action needed and implementing such action, including, as appropriate,
updating documentation;
e) verifying that the corrective action does not adversely affect the ability to meet applicable
regulatory requirements or the safety and performance of the medical device;
f) reviewing the effectiveness of corrective action taken.
Records of the results of any investigation and of action taken shall be maintained (see 4.2.5).
8.5.3 Preventive action
The organization shall determine action to eliminate the causes of potential nonconformities in order to prevent their occurrence. Preventive actions shall be proportionate to the effects of the potential problems.
The organization shall document a procedure to describe requirements for:
a) determining potential nonconformities and their causes;
b) evaluating the need for action to prevent occurrence of nonconformities;
c) planning and documenting action needed and implementing such action, including, as appropriate,
updating documentation;
d) verifying that the action does not adversely affect the ability to meet applicable regulatory
requirements or the safety and performance of the medical device;
e) reviewing the effectiveness of the preventive action taken, as appropriate.
Records of the results of any investigations and of action taken shall be maintained (see 4.2.5).
/

## ISO 13485 & Indonesia GMP
ISO 13485 QMS Requirements |  |  | Indonesia Regulation MOH No. 20 of 2017 HOW TO MAKE HEALTH DEVICES AND HOUSEHOLD MEDICAL SUPPLIES GOOD （Google Translation) |  |  | Difference
Article | Title | Requirement | Article | Title | Requirements
1 | Scope | This International Standard specifies requirements for a quality management system where an organization needs to demonstrate its ability to provide medical devices and related services that consistently meet customer and applicable regulatory requirements. Such organizations can be involved in one or more stages of the life-cycle, including design and development, production, storage and distribution, installation, or servicing of a medical device and design and development or provision of associated activities (e.g. technical support). This International Standard can also be used by suppliers or external parties that provide product, including quality management system-related services to such organizations. Requirements of this International Standard are applicable to organizations regardless of their size and regardless of their type except where explicitly stated. Wherever requirements are specified as applying to medical devices, the requirements apply equally to associated services as supplied by the organization. The processes required by this International Standard that are applicable to the organization, but are not performed by the organization, are the responsibility of the organization and are accounted for in the organization’s quality management system by monitoring, maintaining, and controlling the processes. If applicable regulatory requirements permit exclusions of design and development controls, this can be used as a justification for their exclusion from the quality management system. These regulatory requirements can provide alternative approaches that are to be addressed in the quality management system. It is the responsibility of the organization to ensure that claims of conformity to this International Standard reflect any exclusion of design and development controls. If any requirement in Clauses 6, 7 or 8 of this International Standard is not applicable due to the activities undertaken by the organization or the nature of the medical device for which the quality management system is applied, the organization does not need to include such a requirement in its quality management system. For any clause that is determined to be not applicable, the organization records the justification as described in 4.2.2. | Article 1&2 | / | (1) How to manufacture medical devices and medical supplies good housekeeping includes: a. how to make medical devices that are good; And b. method making supplies health Household (2) Method Making Health devices called CPAKB are guidelines which are used in suite activity making Health devices and quality control which ensure that health device products which are produced fulfill conditions in accordance with objective use. (3) Method Making Supplies Health Household called CPPKRTB is a guideline which is used in suite activity making supplies health householder and quality control to ensure that product supplies health Household Which ensure that produced goods to fulfill conditions which are set in accordance with the intended use. Article 2 (1) Every company Which produce health device or domestic health supplies carrying out production activities must implement: a. CPAKB For company Which produce health device; And b. CPPKRTB For company Which producedhousehold health supplies. (2) Application CPAKB And CPPKRTB as meant in paragraph (1) is proven by a certificate based on CPAKB or CPPKRTB audit results. (3) CPAKB And CPPKRTB as meant on paragraph (1) includes aspects of: a. quality management system; b. management responsibility; c. resource management; d. product realization; And e. measurement, analysis and improvement. (4) CPAKB And CPPKRTB as meant on paragraph (1) listed in Attachment I And Attachment II Which is part No inseparable from Regulation this minister. | The scope of Indonesia Regulation MOH No. 20 of 2017 applies to both Medical Device and Household products.
|  |  | Appendix I I | Introduction | Provisions of Article 106 paragraph (1 ) of Law Number 36 of 2009 about Health mandate that whole preparations pharmacy and Medical devices can only be distributed after obtaining a distribution permit. For support mandate the, so Ministry Health published the Regulation of the Minister of Health Number 1189/Menkes/Per/VII/2010 concerning Health Device Production and Supplies Health Household In Chapter 6 stated that production tool health and/or Supplies Health Household only can be done by companies that have production certificates, and in Article 9 paragraph (2) that company must can ensure that the product made in accordance with Method Making health device and/or Supplies Health Household Good and No happen decline quality and performance during process storage, use, and transportation. CPAKB must be applied by every means of production tool health to guarantee that all processes in the system which each other related in the manufacture of medical devices has been managed in order to achieve safety, quality and benefits of medical devices produced. Application CPAKB also ensure that tool health made in a manner consistent, fulfill condition which set and in accordance with objective its use. CPAKB is a reference for medical device production facilities that can apply system management quality in making tool health. This CPAKB is also a reference for the Central Government and Government Area in do coaching for means production medical devices. CPAKB set condition means health device production, in order to fulfill system management quality. Scope system CPAKB quality management includes design and development, production, installation, not quite enough answer management, management source Power, realization product, measurement, analysis And remediation as well as all related aspects which are done in framework to produce health device that is safe, quality and beneficial.
3 | Terms and definitions | For the purposes of this document, the terms and definitions given in ISO 9000:2015 and the following apply. 3.1 advisory notice notice issued by the organization, subsequent to delivery of the medical device, to provide supplementary information or to advise on action to be taken in the: — use of a medical device, — modification of a medical device, — return of the medical device to the organization that supplied it, or — destruction of a medical device Note 1 to entry: Issuance of an advisory notice can be required to comply with applicable regulatory requirements. 3.2 authorized representative natural or legal person established within a country or jurisdiction who has received a written mandate from the manufacturer to act on his behalf for specified tasks with regard to the latter’s obligations under that country or jurisdiction’s legislation [SOURCE: GHTF/SG1/N055:2009, 5.2] 3.3 clinical evaluation assessment and analysis of clinical data pertaining to a medical device to verify the clinical safety and performance of the device when used as intended by the manufacturer [SOURCE: GHTF/SG5/N4:2010, Clause 4] 3.4 complaint written, electronic or oral communication that alleges deficiencies related to the identity, quality, durability, reliability, usability, safety or performance of a medical device that has been released from the organization’s control or related to a service that affects the performance of such medical devices Note 1 to entry: This definition of “complaint” differs from the definition given in ISO 9000:2015. 3.5 distributor natural or legal person in the supply chain who, on his own behalf, furthers the availability of a medical device to the end user Note 1 to entry: More than one distributor may be involved in the supply chain. Note 2 to entry: Persons in the supply chain involved in activities such as storage and transport on behalf of the manufacturer, importer or distributor, are not distributors under this definition. [SOURCE: GHTF/SG1/N055:2009, 5.3] 3.6 implantable medical device medical device which can only be removed by medical or surgical intervention and which is intended to: — be totally or partially introduced into the human body or a natural orifice, or — replace an epithelial surface or the surface of the eye, and — remain after the procedure for at least 30 days Note 1 to entry: This definition of implantable medical device includes active implantable medical device 3.7 importer natural or legal person in the supply chain who is the first in a supply chain to make a medical device, manufactured in another country or jurisdiction, available in the country or jurisdiction where it is to be marketed [SOURCE: GHTF/SG1/N055:2009, 5.4] 3.8 labelling label, instructions for use, and any other information that is related to identification, technical description, intended purpose and proper use of the medical device, but excluding shipping documents [SOURCE: GHTF/SG1/N70:2011, Clause 4] 3.9 life-cycle all phases in the life of a medical device, from the initial conception to final decommissioning and disposal [SOURCE: ISO 14971:2007, 2.7] 3.10 manufacturer natural or legal person with responsibility for design and/or manufacture of a medical device with the intention of making the medical device available for use, under his name; whether or not such a medical device is designed and/or manufactured by that person himself or on his behalf by another person(s) Note 1 to entry: This “natural or legal person” has ultimate legal responsibility for ensuring compliance with all applicable regulatory requirements for the medical devices in the countries or jurisdictions where it is intended to be made available or sold, unless this responsibility is specifically imposed on another person by the Regulatory Authority (RA) within that jurisdiction. Note 2 to entry: The manufacturer’s responsibilities are described in other GHTF guidance documents. These responsibilities include meeting both pre-market requirements and post-market requirements, such as adverse event reporting and notification of corrective actions. Note 3 to entry: “Design and/or manufacture”, as referred to in the above definition, may include specification development, production, fabrication, assembly, processing, packaging, repackaging, labelling, relabelling, sterilization, installation, or remanufacturing of a medical device; or putting a collection of devices, and possibly other products, together for a medical purpose. Note 4 to entry: Any person who assembles or adapts a medical device that has already been supplied by another person for an individual patient, in accordance with the instructions for use, is not the manufacturer, provided the assembly or adaptation does not change the intended use of the medical device. Note 5 to entry: Any person who changes the intended use of, or modifies, a medical device without acting on behalf of the original manufacturer and who makes it available for use under his own name, should be considered the manufacturer of the modified medical device. Note 6 to entry: An authorized representative, distributor or importer who only adds its own address and contact details to the medical device or the packaging, without covering or changing the existing labelling, is not considered a manufacturer. Note 7 to entry: To the extent that an accessory is subject to the regulatory requirements of a medical device, the person responsible for the design and/or manufacture of that accessory is considered to be a manufacturer. [SOURCE: GHTF/SG1/N055:2009, 5.1] 3.11 medical device instrument, apparatus, implement, machine, appliance, implant, reagent for in vitro use, software, material or other similar or related article, intended by the manufacturer to be used, alone or in combination, for human beings, for one or more of the specific medical purpose(s) of: — diagnosis, prevention, monitoring, treatment or alleviation of disease; — diagnosis, monitoring, treatment, alleviation of or compensation for an injury; — investigation, replacement, modification, or support of the anatomy or of a physiological process; — supporting or sustaining life; — control of conception; — disinfection of medical devices; — providing information by means of in vitro examination of specimens derived from the human body; and does not achieve its primary intended action by pharmacological, immunological or metabolic means, in or on the human body, but which may be assisted in its intended function by such means Note 1 to entry: Products which may be considered to be medical devices in some jurisdictions but not in others include: — disinfection substances; — aids for persons with disabilities; — devices incorporating animal and/or human tissues; — devices for in vitro fertilization or assisted reproduction technologies. [SOURCE: GHTF/SG1/N071:2012, 5.1] 3.12 medical device family group of medical devices manufactured by or for the same organization and having the same basic design and performance characteristics related to safety, intended use and function 3.13 performance evaluation assessment and analysis of data to establish or verify the ability of an in vitro diagnostic medical device to achieve its intended use 3.14 post-market surveillance systematic process to collect and analyse experience gained from medical devices that have been placed on the market 3.15 product result of a process Note 1 to entry: There are four generic product categories, as follows: — services (e.g. transport); — software (e.g. computer program, dictionary); — hardware (e.g. engine mechanical part); — processed materials (e.g. lubricant) Many products comprise elements belonging to different generic product categories. Whether the product is then called service, software, hardware or processed material depends on the dominant element. For example, the offered product “automobile” consists of hardware (e.g. tyres), processed materials (e.g. fuel, cooling liquid), software (e.g. engine control software, driver’s manual), and service (e.g. operating explanations given by the salesman). Note 2 to entry: Service is the result of at least one activity necessarily performed at the interface between the supplier and customer and is generally intangible. Provision of a service can involve, for example, the following: — an activity performed on a customer-supplied tangible product (e.g. automobile to be repaired); — an activity performed on a customer-supplied intangible product (e.g. the income statement needed to prepare a tax return); — the delivery of an intangible product (e.g. the delivery of information in the context of knowledge transmission); — the creation of ambience for the customer (e.g. in hotels and restaurants). Software consists of information and is generally intangible and can be in the form of approaches, transactions or procedures. Hardware is generally tangible and its amount is a countable characteristic. Processed materials are generally tangible and their amount is a continuous characteristic. Hardware and processed materials often are referred to as goods. Note 3 to entry: This definition of “product” differs from the definition given in ISO 9000:2015. [SOURCE: ISO 9000:20052), 3.4.2, modified] 3.16 purchased product product provided by a party outside the organization’s quality management system Note 1 to entry: The provision of product does not necessarily infer a commercial or financial arrangement. 3.17 risk combination of the probability of occurrence of harm and the severity of that harm Note 1 to entry: This definition of “risk” differs from the definition given in ISO 9000:2015. [SOURCE: ISO 14971:2007, 2.16] 3.18 risk management systematic application of management policies, procedures and practices to the tasks of analysing, evaluating, controlling and monitoring risk [SOURCE: ISO 14971:2007, 2.22] 3.19 sterile barrier system minimum package that prevents ingress of microorganisms and allows aseptic presentation of the product at the point of use [SOURCE: ISO 11607‑1:2006, 3.22] 3.20 sterile medical device medical device intended to meet the requirements for sterility Note 1 to entry: The requirements for sterility of a medical device can be subject to applicable regulatory requirements or standards. | / | / | /
4 | Quality Management System |  | Appendix I III A | Aspects of Good Medical Device Making Methods Quality Management System
4.1 | General Requirements |  | Appendix I III A1 | General requirement
4.1.1 | (4.1 General Requirements) | 4.1.1 The organization shall document a quality management system and maintain its effectiveness in accordance with the requirements of this International Standard and applicable regulatory requirements. The organization shall establish, implement and maintain any requirement, procedure, activity or arrangement required to be documented by this International Standard or applicable regulatory requirements. The organization shall document the role(s) undertaken by the organization under the applicable regulatory requirements. NOTE Roles undertaken by the organization can include manufacturer, authorized representative, importer or distributor. |  |  | a. Company must have production certificate which still applies according to the category of medical devices produced. b. Companies must: 7) Determine the parts or functions of the company which have their own task and not quite enough answers which are defined clearly and unequivocally. Companies must establish, document, implement and maintain the CPAKB and maintain its effectiveness in accordance with the condition CPAKB. | Indonesia GMP specifies the CPAKB certificate requirement
4.1.2 | (4.1 General Requirements) | 4.1.2 The organization shall: a) determine the processes needed for the quality management system and the application of these processes throughout the organization taking into account the roles undertaken by the organization; b) apply a risk based approach to the control of the appropriate processes needed for the quality management system; c) determine the sequence and interaction of these processes. |  |  | b. Companies must: 1) Identify and set processes which are needed for CPAKB; 2) Determine the order and interaction of the above processes; | Indonesia GMP does not mention risk based approach
4.1.3 | (4.1 General Requirements) | 4.1.3 For each quality management system process, the organization shall: a) determine criteria and methods needed to ensure that both the operation and control of these processes are effective; b) ensure the availability of resources and information necessary to support the operation and monitoring of these processes; c) implement actions necessary to achieve planned results and maintain the effectiveness of these processes; d) monitor, measure as appropriate, and analyse these processes; e) establish and maintain records needed to demonstrate conformance to this International Standard and compliance with applicable regulatory requirements (see 4.2.5). |  |  | 3) Determine criteria and method which needed to ensure the effectiveness of the implementation and control of the process; 4) Ensure availability source power and information which needed to support the implementation and monitoring of this process; 5) Monitor, measure and analyze these processes; 6) Implement action which needed to reach results which planned and maintain the effectiveness of this process; | No signficant difference
4.1.4 | (4.1 General Requirements) | 4.1.4 The organization shall manage these quality management system processes in accordance with the requirements of this International Standard and applicable regulatory requirements. Changes to be made to these processes shall be: a) evaluated for their impact on the quality management system; b) evaluated for their impact on the medical devices produced under this quality management system; c) controlled in accordance with the requirements of this International Standard and applicable regulatory requirements. | / | / | / | No requirement defined in Indonesia GMP for QMS process changes
4.1.5 | (4.1 General Requirements) | 4.1.5 When the organization chooses to outsource any process that affects product conformity to requirements, it shall monitor and ensure control over such processes. The organization shall retain responsibility of conformity to this International Standard and to customer and applicable regulatory requirements for outsourced processes. The controls shall be proportionate to the risk involved and the ability of the external party to meet the requirements in accordance with 7.4. The controls shall include written quality agreements. | Appendix I III A1 | General requirement | c. If the company chooses to submit to the other part process which influences the suitability product with condition, the company must ensure the entire control process in question. The control process submitted to parties must be identified in CPAKB. | No signficant difference
4.1.6 | (4.1 General Requirements) | 4.1.6 The organization shall document procedures for the validation of the application of computer software used in the quality management system. Such software applications shall be validated prior to initial use and, as appropriate, after changes to such software or its application. The specific approach and activities associated with software validation and revalidation shall be proportionate to the risk associated with the use of the software. Records of such activities shall be maintained (see 4.2.5). | / | / | / | No requirement defined in Indonesia GMP for software validation
4.2 | Documentation Requirements |  | Appendix I III A2 | Documentation and Control
4.2.1 | General | 4.2.1 General The quality management system documentation (see 4.2.4) shall include: a) documented statements of a quality policy and quality objectives; b) a quality manual; c) documented procedures and records required by this International Standard; d) documents, including records, determined by the organization to be necessary to ensure the effective planning, operation, and control of its processes; e) other documentation specified by applicable regulatory requirements. | Appendix I III A2 | Documentation and Control | CPAKB documentation should include: a. Quality Guidelines/Quality Manual Quality Guidelines is a guidelines which containing statement and commitment company about CPAKB application, contains: 1) Vision and mission; 2) Quality policy and objectives Quality Policy is a document about a statement of company commitment regarding the implementation of CPAKB. Example of quality policy: We PT. ALKES continues to increase quality products and service to better meet the needs of our customers. Quality objective is the target which is achieved by the company. Examples of quality objectives: Management: IDR 500,000,000 profit in 2015 Production Department: 5% production reject rate Marketing Department: increase in product sales by 30% from 2014 Maintenance Department: production delays due maximum machine damage 15 minutes / day Quality Inspection Section: 95% product release Part Source Power Man: level presence appropriate employee time 95% 3) Organizational structure; 4) Business process is the governance of the production process; b. Mandatory procedures and records required by CPAKB 1) Mandatory procedures required in CPAKB, include: a) Procedures for purchasing raw materials/packaging materials; b) Procedures for handling raw materials/packaging materials; c) Design and development procedures; *(If there are) d) Procedures related to the production process; e) Packaging and labeling procedures; f) Production equipment maintenance procedures g) Document control procedures; h) Records control procedures; i) Procedure for taking samples of finished products; j) Procedure testing/examination product ( in process And finished product); k) Inappropriate product control procedures; l) Procedure for reprocessing a batch/lot of product; * (If there are) m) Sterile process validation procedures; * (If there are) n) Procedure validation application device soft computer; * (If There is) o) traceability procedures; p) Product storage procedure; q) Procedure control product with age To use limited or need condition storage special required; * (If there are) r) After-sales service procedures; s) Procedure calibration And verification tool inspection And measurement; t) Status identification procedures including returned products; u) Procedure training personnel/HR about education, training, skills and experience; v) Management review procedures w) Continuous improvement procedures x) Internal audit procedures; y) Procedure control problem related product quality ; z) Data analysis procedures related to the quality management system; a a) Procedures related to corrective and preventive actions; bb) Waste handling procedures; And cc) Procedures related to sanitation and hygiene. The above procedure must include: (1) Procedure title; (2) Valid number and date; (3) Revision number and date; (4) Fill, in the form of: objective, room scope, definition, document related, responsibilities, procedures, exceptions, attachments, forms; And (5) The position of the personnel carrying out the procedure, check and approve procedures. 2) The records required in the CPAKB include: a) recording related design And development product;*(if There is) b) records related to material purchasing information raw/packaging and supplier evaluation; c) records of inspection of raw materials/packaging; d) recording of production processes and production results; e) production process validation record; f) records of audit results and follow-up plans; g) management review records; h) recording personnel/HR about education, training, ability and experience; i) sterilization process validation record (if any); j) recording validation use device soft computer (If any); k) product quality testing records l) records for each batch of medical device production m) recording processing repeat something batch/lot product (If any) n) production equipment maintenance records; o) recording of calibration results of measuring instruments; p) room cleaning records; q) pest control records; r) records of installation activities and their verification;*(if any) s) records of after-sales activities; t) records of the distribution of medical devices; u) recording product Which No in accordance And action Which taken; v) recording results investigation And action corrective And preventive; And w) records regarding customer properties (if any). d. Documents related to applicable regulations: 1) Medical Device Production Certificate. 2) Medical Device product standard documents. 3) Product distribution license. 4) Medical Device Distributor Permit (if any). 5) Export-import certificate (if any). 6) Copy Constitution Number 36 Year 2009 about Health. 7) Copy of Regulation of the Minister of Health Number 1189/MENKES/PER/VIII/2010 concerning Tool Production Health and Household Health Supplies. 8) Copy of Regulation of the Minister of Health Number 1190/MENKES/PER/VIII/2010 about Permission Circ health device and Household Health Supplies. 9) Copy of Regulation of the Minister of Health Number 1191/MENKES/PER/VIII/2010 about Distribution health device. | Indonesia GMP provides detailed example of Quality Policy and Quality Objective Indonesia GMP provides comprehensive procedure/record list required by CPAKB
4.2.2 | Quality Manual | 4.2.2 Quality manual The organization shall document a quality manual that includes: a) the scope of the quality management system, including details of and justification for any exclusion or non-application; b) the documented procedures for the quality management system, or reference to them; c) a description of the interaction between the processes of the quality management system. The quality manual shall outline the structure of the documentation used in the quality management system. |  |  | 5) Scope of quality management system, including details and justification for every exception activities/activities that are not carried out by medical device companies, for example: design, etc; 6) documented procedure list which set for CPAKB; 7) job description; 8) description of the interaction between the process of implementing CPAKB; 9) channel process is channel production process health device starting from concept and design, production, packaging labeling, sales, use up to disposal; And 10) management representative appointment letter. Hierarchy of quality management system documentation is established | Indonesia GMP includes JD and clear hierarchy of documentation used in quality management system
4.2.3 | Medical device file | For each medical device type or medical device family, the organization shall establish and maintain one or more files either containing or referencing documents generated to demonstrate conformity to the requirement of this International Standard and compliance with applicable regulatory requirements. The content of the file(s) shall include, but is not limited to: a) general description of the medical device, intended use/purpose, and labelling, including any instructions for use; b) specifications for product; c) specifications or procedures for manufacturing, packaging, storage, handling and distribution; d) procedures for measuring and monitoring; e) as appropriate, requirements for installation; f) as appropriate, procedures for servicing. |  |  | c. Technical documents related to manufactured products The medical device technical documents produced include: 1) raw material specifications. 2) finished product specifications. 3) Process flow for production and packaging. 4) packaging material specifications. 5) bulk product specifications, if any. 6) formula, composition of constituent components, if any. 7) product design sketch drawings ( wiring diagrams ), If any 8) list of equipment used in production. 9) sampling procedure for testing. 10) product quality testing. 11) storage method. 12) useful life and/or expiration date. 13) how to install and maintain it, if any. | Indonesia develops more detailed document list for Medical Device File.
4.2.4 | Control of documents | 4.2.4 Control of documents Documents required by the quality management system shall be controlled. Records are a special type of document and shall be controlled according to the requirements given in 4.2.5. A documented procedure shall define the controls needed to: a) review and approve documents for adequacy prior to issue; b) review, update as necessary and re-approve documents; c) ensure that the current revision status of and changes to documents are identified; d) ensure that relevant versions of applicable documents are available at points of use; e) ensure that documents remain legible and readily identifiable; f) ensure that documents of external origin, determined by the organization to be necessary for the planning and operation of the quality management system, are identified and their distribution controlled; g) prevent deterioration or loss of documents; h) prevent the unintended use of obsolete documents and apply suitable identification to them. The organization shall ensure that changes to documents are reviewed and approved either by the original approving function or another designated function that has access to pertinent background information upon which to base its decisions. The organization shall define the period for which at least one copy of obsolete documents shall be retained. This period shall ensure that documents to which medical devices have been manufactured and tested are available for at least the lifetime of the medical device as defined by the organization, but not less than the retention period of any resulting record (see 4.2.5), or as specified by applicable regulatory requirements. |  |  | e. Control of documents and records Document control aims to: 1) Ensure that document which needed available for required personnel. 2) Manage the arrangement, distribution and how to maintain document. 3) Arrange withdrawal and extermination document which related to CPAKB. Companies must ensure that change to document reviewed and proven alignement with orignal function or with other designed functions it has access to matter which relates with information background behind even come to the basis of the decision. Companies must determine the retention period of obsolete documents for document control shall be retained. The retention period will ensure that documents from health devices which have been produced and tested available for at least the lifetime of health devices like which have been determined by organization, but not less than 2 (two) years. | No detailed document control requirements specified in Indonesia GMP
4.2.5 | Control of records | 4.2.5 Control of records Records shall be maintained to provide evidence of conformity to requirements and of the effective operation of the quality management system. The organization shall document procedures to define the controls needed for the identification, storage, security and integrity, retrieval, retention time and disposition of records. The organization shall define and implement methods for protecting confidential health information contained in records in accordance with the applicable regulatory requirements. Records shall remain legible, readily identifiable and retrievable. Changes to a record shall remain identifiable. The organization shall retain the records for at least the lifetime of the medical device as defined by the organization, or as specified by applicable regulatory requirements, but not less than two years from the medical device release by the organization. |  |  | f. Record controller _ Recording must be set and maintained to give evidence of conformity to requirements and effectiveness of operation system management quality. Recording must still easily be read, can identified easily, and can obtain return. Procedure control recording must be made for objective identification, storage, protection, retrievable, time needed and disposal recording. Companies must look after recording for a period of time at least the lifetime of the health device as specified by the organization, but not less than 2 (two) years from the date the product is marketed by the organization. | No significant difference
5 | Management Responsibility |  | Appendix I III B | Management Responsibilities
5.1 | Management Commitment | 5.1 Management commitment Top management shall provide evidence of its commitment to the development and implementation of the quality management system and maintenance of its effectiveness by: a) communicating to the organization the importance of meeting customer as well as applicable regulatory requirements; b) establishing the quality policy; c) ensuring that quality objectives are established; d) conducting management reviews; e) ensuring the availability of resources. | Appendix I III B1 | Management commitment | 1. Management commitment Leaders must prove their commitment to development and implementation CPAKB aAnd maintain effectiveness by: a. communicate with internal organizations; b. understand customers; c. comply with statutory provisions in the field medical devices; d. establish quality policy; e. ensure the quality objectives set; f. conduct a management review; and g. guarantee the availability of resources. | Similar
5.2 | Customer focus | 5.2 Customer focus Top management shall ensure that customer requirements and applicable regulatory requirements are determined and met. | Appendix I III B2 | Customer focus | 2. Customer focus Leaders guarantee that the needs of the requirements of the customer specified can be met. | Similar
5.3 | Quality Policy | 5.3 Quality policy Top management shall ensure that the quality policy: a) is applicable to the purpose of the organization; b) includes a commitment to comply with requirements and to maintain the effectiveness of the quality management system; c) provides a framework for establishing and reviewing quality objectives; d) is communicated and understood within the organization; e) is reviewed for continuing suitability. | Appendix I III B3 | Quality policy | 3. Quality policy Management must ensure that the quality policy: a. in accordance with the goals of the organization, namely in line with the vision and company mission; b. includes a written commitment to comply with requirements and maintain the effectiveness of the quality management system; c. provide a framework for defining and reviewing quality objectives; d. communicated and understandable in organization, It means must socialized to whole employee start from management level up to the executive level; e. reviewed for continuing suitability. | Similar
5.4 | Planning
5.4.1 | Quality Objectives | 5.4.1 Quality objectives Top management shall ensure that quality objectives, including those needed to meet applicable regulatory requirements and requirements for product, are established at relevant functions and levels within the organization. The quality objectives shall be measurable and consistent with the quality policy. | Appendix I III B4 | Quality Goals | 4. Quality Goals Leader must ensure that the quality target is made to fulfill the condition for the product and set on function and level according to the organization. Quality objectives must be appropriate and consistent with quality ploicy. Quality target is goals to be achieved by the organization. Quality policy to be referenced in set target quality to ensure suitability and consistency of quality objectives with the quality policy. | Similar
5.4.2 | Quality management system planning | 5.4.2 Quality management system planning Top management shall ensure that: a) the planning of the quality management system is carried out in order to meet the requirements given in 4.1, as well as the quality objectives; b) the integrity of the quality management system is maintained when changes to the quality management system are planned and implemented. | Appendix I III B5 | Quality management system planning | 5. Quality management system planning Leader must ensure that planning system management quality held in framework fulfill condition CPAKB and to achieve quality objectives. The quality management system plan contains information on: a. quality target which will be achieved, can written per part or function; b. plan activity related with achievement quality target; And c. personnel which will carry out and responsible for these activities. | Indonesia GMP does not mention the planning & implementation of changes to QMS
5.5 | Responsibility, authority and communication
5.5.1 | Responsibility and authority | 5.5.1 Responsibility and authority Top management shall ensure that responsibilities and authorities are defined, documented and communicated within the organization. Top management shall document the interrelation of all personnel who manage, perform and verify work affecting quality and shall ensure the independence and authority necessary to perform these tasks. | Appendix I III B6 | Leadership responsibility and authority | 6. Leadership responsibility and authority Leaders ensure that responsiblity and authority are defined, documented and communicated in the organization. Leaders set connections from all personnel, who organize, execute and check the work influence quality, and must ensure independence and authority required to carry out this task. | Similar
5.5.2 | Management representative | 5.5.2 Management representative Top management shall appoint a member of management who, irrespective of other responsibilities, has responsibility and authority that includes: a) ensuring that processes needed for the quality management system are documented; b) reporting to top management on the effectiveness of the quality management system and any need for improvement; c) ensuring the promotion of awareness of applicable regulatory requirements and quality management system requirements throughout the organization. | Appendix I III B7 | Management representative | 7. Management representative Leader must pointing one member management to become representative management which own responsibility and authority in: a. Ensure that process which needed for quality management system is established, implemented and maintained; b. Report to management about performance from quality management system and any required for enhancement; c. Push awareness to fulfillment of regulation and customer requirements; And d. Connection with external party about problem related to the quality management system. | Similar
5.5.3 | Internal communication | 5.5.3 Internal communication Top management shall ensure that appropriate communication processes are established within the organization and that communication takes place regarding the effectiveness of the quality management system. | Appendix I III B8 | Internal communication | 8. Internal communication Leaders must ensure that process communication is appropriately implemented in the company and the communication held regarding effective system management quality. | Similar
5.6 | Management review
5.6.1 | General | 5.6.1 General The organization shall document procedures for management review. Top management shall review the organization’s quality management system at documented planned intervals to ensure its continuing suitability, adequacy and effectiveness. The review shall include assessing opportunities for improvement and the need for changes to the quality management system, including the quality policy and quality objectives. Records from management reviews shall be maintained (see 4.2.5). | Appendix I III B9 | Management Review | Leader must perform management review to review quality management system organization, at planned intervals, to ensure continuing suitability, adequacy and effectiveness. The implementation time of management review depends on determination by the organization, but at least once a year. Management review is carried out by way of meetings shared between parts or functions to describe achievement and constraints which are faced. Overview management must cover evaluation chances for remediation and need to improve the quality management system, quality policy and quality objectives. | Indonesia GMP defines annual management review
5.6.2 | Review Input | 5.6.2 Review input The input to management review shall include, but is not limited to, information arising from: a) feedback; b) complaint handling; c) reporting to regulatory authorities; d) audits; e) monitoring and measurement of processes; f) monitoring and measurement of product; g) corrective action; h) preventive action; i) follow-up actions from previous management reviews; j) changes that could affect the quality management system; k) recommendations for improvement; l) applicable new or revised regulatory requirements. |  | Input management review | a. Input management review Input to management review should include information about: 1) Audit results, both internal and external audits; 2) Bait come back customer, containing about various complaint customer Which come in a manner direct nor from results survey satisfaction customer, as well as step prevention And improvement of customer complaints; 3) Performance process and product suitability, each part or function displays achievement target quality Which has determined, as well as step prevention And improvement for quality objectives that cannot be achieved; 4) Status of preventive and corrective actions; 5) Actions from previous management reviews; 6) Change which can influence quality management system, for example: changes in procedures, quality manuals, processes, organizational structure; 7) Recommendations for development; And 8) New or revised regulatory requirements. | Similar
5.6.3 | Review Output | 5.6.3 Review output The output from management review shall be recorded (see 4.2.5) and include the input reviewed and any decisions and actions related to: a) improvement needed to maintain the suitability, adequacy, and effectiveness of the quality management system and its processes; b) improvement of product related to customer requirements; c) changes needed to respond to applicable new or revised regulatory requirements; d) resource needs. |  | Review output | b. Review output Output from review repeat management must covers every decisions and actions related to: 1) Remediation needed to maintain the effectiveness of the quality management system and its processes; 2) Product improvements related to customer requirements; And 3) Required resources. | Indonesia GMP does not mention the changes needed to respond to applicable new or revised regulatory requirements;
6 | Resource management |  | Appendix I III C | Resource Management
6.1 | Provision of resources | 6.1 Provision of resources The organization shall determine and provide the resources needed to: a) implement the quality management system and to maintain its effectiveness; b) meet applicable regulatory and customer requirements. |  |  | Companies must determine and provide resources that are needed to implement CPAKB and maintain its effectiveness. | No significant difference
6.2 | Human resources | 6.2 Human resources Personnel performing work affecting product quality shall be competent on the basis of appropriate education, training, skills and experience. The organization shall document the process(es) for establishing competence, providing needed training, and ensuring awareness of personnel. The organization shall: a) determine the necessary competence for personnel performing work affecting product quality; b) provide training or take other actions to achieve or maintain the necessary competence; c) evaluate the effectiveness of the actions taken; d) ensure that its personnel are aware of the relevance and importance of their activities and how they contribute to the achievement of the quality objectives; e) maintain appropriate records of education, training, skills and experience (see 4.2.5). NOTE The methodology used to check effectiveness is proportionate to the risk associated with the work for which the training or other action is being provided. | Appendix I III C1 | Personnel | 1. Personnel a. Assure technical own competence, ability and/or experience that is appropriate with the medical device it produced. b. Assure product quality supervision responsible for product safety, quality and benefits. c. Every personnel responsible to evaluate effectiveness to action which has been done as well as aware of the relation and meaning of the actions taken, and their contribution in achieving quality objectives. d. Available personnel which look after recording which in accordance with education, training, abilities and experience. | Similar
6.3 | Infrastructure | 6.3 Infrastructure The organization shall document the requirements for the infrastructure needed to achieve conformity to product requirements, prevent product mix-up and ensure orderly handling of product. Infrastructure includes, as appropriate: a) buildings, workspace and associated utilities; b) process equipment (both hardware and software); c) supporting services (such as transport, communication, or information systems). The organization shall document requirements for the maintenance activities, including the interval of performing the maintenance activities, when such maintenance activities, or lack thereof, can affect product quality. As appropriate, the requirements shall apply to equipment used in production, the control of the work environment and monitoring and measurement. Records of such maintenance shall be maintained (see 4.2.5). | Appendix I III C2 | Buildings and facilities | 2. Buildings and facilities Companies must determine, provide and look after the building and facility which is needed to reach compliance with product requirements. Rooms in production facility buildings include: a. Raw material/packaging/component receiving area; b. Storage room for raw materials/packaging/components; c. Production room; d. Room inspection quality material raw/packaging/ component and finished products; e. Area quarantine For material raw Which Not yet tested, product recalled ( recall ) and rejected/returned products; f. Area for storage product between and product which untested; g. Finished product storage room; h. Room for changing clothes, nursing room, dining room and toilet; And i. Health service room. Must-have facilities include: 1) production equipment; 2) equipment for testing the quality of raw materials and finished products; 3) equipment for maintenance; And 4) equipment for installation, if any. General provisions of buildings and facilities: a) Own design and construction which adequate and ensure K3 (Occupational Safety and Health). b) The layout of the building is designed in such a way as to avoid from the contaminated environment around .like polluted air, land and water. c) Designed, equipped and maintained to obtain protection maximum to influence weather, flood, seepage from land as well as enter and nest insects, birds, rodents, or other animals. d) Must treated with careful, cleaned and when need disinfected. e) Whole building and facility including room production, testing quality, area storage and environment around the building must be maintained in a clean and tidy condition. f) Power electricity, light lighting, temperature, humidity, and ventilation must be maintained so as not to cause an impact which harms the products in a manner directly nor indirectly to the product. g) Precautions must be taken for personnel who are not interested in entering. h) Design and system location room must ensure suitability with activity production other which Possible done in same or adjoining and appropriate production space with the production process flow. i) Design and system location room must ensure that room production is not utilized as track then cross general for personnel and material/product, or as place storage of materials or products other than those being processed. | Indonesia GMP provides detailed requirements for facilities, infrastructure of supporting services is not covered.
6.4 | Work environment and contamination control
6.4.1 | Work environment | 6.4.1 Work environment The organization shall document the requirements for the work environment needed to achieve conformity to product requirements. If the conditions for the work environment can have an adverse effect on product quality, the organization shall document the requirements for the work environment and the procedures to monitor and control the work environment. The organization shall: a) document requirements for health, cleanliness and clothing of personnel if contact between such personnel and the product or work environment could affect medical device safety or performance; b) ensure that all personnel who are required to work temporarily under special environmental conditions within the work environment are competent or supervised by a competent person. NOTE Further information can be found in ISO 14644 and ISO 14698. | Appendix I III C3 | Work environment | 3. Work environment Companies must determine and arrange environmental work which is needed to reach suitability with condition products. a. Companies must set and implement mandatory requirements for health, hygiene and safety personnel, if contact between personnel and product or work environment can affect the quality product and personnel health and safety. b. If condition environment work can result in an effect which is not expected for product quality, the company must establish mandatory requirements for working environmental conditions and procedure must or work instruction to monitor and control the working environment. c. Companies must ensure that all personnel which require to work in a manner temporarily on condition special in environment work are trained and regularly supervised by trained personnel. | Similar
6.4.2 | Contamination control | 6.4.2 Contamination control As appropriate, the organization shall plan and document arrangements for the control of contaminated or potentially contaminated product in order to prevent contamination of the work environment, personnel, or product. For sterile medical devices, the organization shall document requirements for control of contamination with microorganisms or particulate matter and maintain the required cleanliness during assembly or packaging processes. |  |  | d. If appropriate, a special order should be established and documented for product control contaminated or which in a manner potentially contaminated In order to prevent contamination of other products, work environment or personnel. e. The company must have a waste handling procedure. | No significant difference
7 | Product realization |  | Appendix I III D | Product Realization
7.1 | Planning of product realization | The organization shall plan and develop the processes needed for product realization. Planning of product realization shall be consistent with the requirements of the other processes of the quality management system. The organization shall document one or more processes for risk management in product realization. Records of risk management activities shall be maintained (see 4.2.5). In planning product realization, the organization shall determine the following, as appropriate: a) quality objectives and requirements for the product; b) the need to establish processes and documents (see 4.2.4) and to provide resources specific to the product, including infrastructure and work environment; c) required verification, validation, monitoring, measurement, inspection and test, handling, storage, distribution and traceability activities specific to the product together with the criteria for product acceptance; d) records needed to provide evidence that the realization processes and resulting product meet requirements (see 4.2.5). The output of this planning shall be documented in a form suitable for the organization’s method of operations. NOTE Further information can be found in ISO 14971. | Appendix I III D 1 | Product realization planning | 1. Product realization planning Companies must plan and develop the process which is needed for the realization of the product. Planning realization of products must align with the other process requirements of CPAKB. In product realization planning, companies must determine: a. Quality objectives and product quality requirements; b. The need to define processes, documents, and availability of resources; c. Activity verification which required, validation, monitoring, inspection and testing quality to product and criteria product acceptability; And d. Recording which needed to give proof of process realization and product which was generated to fulfill the condition. Production planning includes: 1) Production planning based on customer demand. 2) Planning for procurement of raw materials based on planning production and stock. 3) Raw material testing planning. 4) Planning timetable production with consider remainder stock. 5) Plan storage raw material in accordance capacity (raw material which new accepted, raw material which quarantine, and raw materials that have passed the test). 6) Planning supply label identity for every raw material. 7) Planning for release or rejection of label supplies all stored raw materials. | Indonesia GMP does not mention risk management in product realization Indonesia GMP provides detailed production planning requirement
7.2 | Customer-related processes |  | Appendix I III D2 | Process related to customers
7.2.1 | Determination of requirements related to product | 7.2.1 Determination of requirements related to product The organization shall determine: a) requirements specified by the customer, including the requirements for delivery and post-delivery activities; b) requirements not stated by the customer but necessary for specified or intended use, as known; c) applicable regulatory requirements related to the product; d) any user training needed to ensure specified performance and safe use of the medical device; e) any additional requirements determined by the organization. | Appendix I III D2 a | Company determine condition related with product | a. Company determine condition related with product includes: 1) Condition which set by customer, including distribution and after-sale requirements; 2) Condition which is not set by customer but necessary related to the use of medical devices; 3) Regulation legislation and regulation which hook with products, such as distribution permits for medical devices; And 4) Additional requirements determined by the company. | Similar
7.2.2 | Review of requirements related to product | 7.2.2 Review of requirements related to product The organization shall review the requirements related to product. This review shall be conducted prior to the organization’s commitment to supply product to the customer (e.g. submission of tenders, acceptance of contracts or orders, acceptance of changes to contracts or orders) and shall ensure that: a) product requirements are defined and documented; b) contract or order requirements differing from those previously expressed are resolved; c) applicable regulatory requirements are met; d) any user training identified in accordance with 7.2.1 is available or planned to be available; e) the organization has the ability to meet the defined requirements. Records of the results of the review and actions arising from the review shall be maintained (see 4.2.5). When the customer provides no documented statement of requirement, the customer requirements shall be confirmed by the organization before acceptance. When product requirements are changed, the organization shall ensure that relevant documents are amended and that relevant personnel are made aware of the changed requirements. | Appendix I III D2 b | Reviews the requirements related to the product | b. The company reviews the requirements related to the product which was held before committed company supply products to consumers. Review includes: 1) Product requirements are defined; 2) Condition contract or request which is different from previously stated; And 3) The company has the ability to fulfill stipulated requirements. If a customer does not give a statement written about product requirements, customer requirements must be confirmed by the company before acceptance work. If the condition of the product is changed, the company must ensure that document related changes and personnel concerned are reminded of the change in requirements. | Indonesia GMP does not mention the applicable regulatory requirements to product
7.2.3 | Communication | 7.2.3 Communication The organization shall plan and document arrangements for communicating with customers in relation to: a) product information; b) enquiries, contracts or order handling, including amendments; c) customer feedback, including complaints; d) advisory notices. The organization shall communicate with regulatory authorities in accordance with applicable regulatory requirements. | Appendix I III D2 c | Customer communication | c. Customer communication Company must communicate with related customer: 1) Product information; 2) Inquiries, handling of contracts or requests, including changes; 3) Bait come back customer, including complaint customer and customer satisfaction; and 4) Warning note. | Similar
7.3 | Design and development
7.3.1 | General | 7.3.1 General The organization shall document procedures for design and development.
7.3.2 | Design and development planning | 7.3.2 Design and development planning The organization shall plan and control the design and development of product. As appropriate, design and development planning documents shall be maintained and updated as the design and development progresses. During design and development planning, the organization shall document: a) the design and development stages; b) the review(s) needed at each design and development stage; c) the verification, validation, and design transfer activities that are appropriate at each design and development stage; d) the responsibilities and authorities for design and development; e) the methods to ensure traceability of design and development outputs to design and development inputs; f) the resources needed, including necessary competence of personnel. | Appendix I III D3 | Product design development planning | 3. Product design development and planning (if any) , Company must set procedure must be related with planning and control design and development. During design and planning development, company must specify: a. Design and development conditions; b. Review, verification, validation and activity transfer design which in accordance on every condition design and development; And c. Responsible for design and development. Product development design and planning activities include: Planning output . Output planning ( planning output ) must be documented, and renewed if allowed, including design and development process Design transfer activities during the design process and development guarantees that the design and output Development is already in accordance with made before it becomes a final product specification. | Simplified requirements defined in Indonesia GMP
7.3.3 | Design and development inputs | 7.3.3 Design and development inputs Inputs relating to product requirements shall be determined and records maintained (see 4.2.5). These inputs shall include: a) functional, performance, usability and safety requirements, according to the intended use; b) applicable regulatory requirements and standards; c) applicable output(s) of risk management; d) as appropriate, information derived from previous similar designs; e) other requirements essential for design and development of the product and processes. These inputs shall be reviewed for adequacy and approved. Requirements shall be complete, unambiguous, able to be verified or validated, and not in conflict with each other. NOTE Further information can be found in IEC 62366–1. | Appendix I III D4 | Design and development input | 4. Inputs related to product requirements should be specified and records are maintained. This input should include: a) Functional, performance and condition safety, in accordance with intended use; b) Regulatory requirements and statutory regulations can be applied; c) Where applicable, the information derived from the design is the same previously; And d) Other key requirements for design and development. Input must be reviewed repeatedly for adequacy and endorsement. Conditions must be complete, not ambiguous and no conflict with any other terms. | No risk management requirement defined in Indonesia GMP
7.3.4 | Design and development outputs | 7.3.4 Design and development outputs Design and development outputs shall: a) meet the input requirements for design and development; b) provide appropriate information for purchasing, production and service provision; c) contain or reference product acceptance criteria; d) specify the characteristics of the product that are essential for its safe and proper use. The outputs of design and development shall be in a form suitable for verification against the design and development inputs and shall be approved prior to release. Records of the design and development outputs shall be maintained (see 4.2.5). | Appendix I III D5 | Design and development output | 5. Design and development output Output of design and development must be available in form which make it easy verification to input design and development and must be approved before release to the market. Design and development outputs should: a. Meet the input requirements for design and development, b. Give information which in accordance for purchase, production and for terms of service, c. Contains or refers to product acceptance criteria, and d. Define product characteristics that are important to safety and correct use of the product. Recording output design and development must be maintained. Recording output design and development can include specification, procedure making, picture engineering, and logbooks engineering and research. | Similar
7.3.5 | Design and development review | 7.3.5 Design and development review At suitable stages, systematic reviews of design and development shall be performed in accordance with planned and documented arrangements to: a) evaluate the ability of the results of design and development to meet requirements; b) identify and propose necessary actions. Participants in such reviews shall include representatives of functions concerned with the design and development stage being reviewed, as well as other specialist personnel. Records of the results of the reviews and any necessary actions shall be maintained and include the identification of the design under review, the participants involved and the date of the review (see 4.2.5). | Appendix I III D6 | Design and development review | 6. Design and development review On condition which in accordance, review repeat which systematic to design and development must be carried out according to the order planned to evaluate ability results design and development to meet requirements, and to identify every problem and plan action which needed. Participants associated with the review should include representatives from function which responsible answer with condition design and development which currently reviewed repeat, also personnel special other. Recording results review repeat and every action which is required must be maintained. | Similar
7.3.6 | Design and development verification | 7.3.6 Design and development verification Design and development verification shall be performed in accordance with planned and documented arrangements to ensure that the design and development outputs have met the design and development input requirements. The organization shall document verification plans that include methods, acceptance criteria and, as appropriate, statistical techniques with rationale for sample size. If the intended use requires that the medical device be connected to, or have an interface with, other medical device(s), verification shall include confirmation that the design outputs meet design inputs when so connected or interfaced. Records of the results and conclusions of the verification and necessary actions shall be maintained (see 4.2.4 and 4.2.5). | Appendix I III D7 | Design and development verification | 7. Design and development verification Verification must be held in accordance with the arrangement which is planned to ensure that output design and development has fulfilled condition input design and development. Records of verification results and any actions taken required must be maintained | Simplified requirements defined in Indonesia GMP
7.3.7 | Design and development validation | 7.3.7 Design and development validation Design and development validation shall be performed in accordance with planned and documented arrangements to ensure that the resulting product is capable of meeting the requirements for the specified application or intended use. The organization shall document validation plans that include methods, acceptance criteria and, as appropriate, statistical techniques with rationale for sample size. Design validation shall be conducted on representative product. Representative product includes initial production units, batches or their equivalents. The rationale for the choice of product used for validation shall be recorded (see 4.2.5). As part of design and development validation, the organization shall perform clinical evaluations or performance evaluations of the medical device in accordance with applicable regulatory requirements. A medical device used for clinical evaluation or performance evaluation is not considered to be released for use to the customer. If the intended use requires that the medical device be connected to, or have an interface with, other medical device(s), validation shall include confirmation that the requirements for the specified application or intended use have been met when so connected or interfaced. Validation shall be completed prior to release for use of the product to the customer. Records of the results and conclusion of validation and necessary actions shall be maintained (see 4.2.4 and 4.2.5). | Appendix I III D8 | Design and development validation | 8. Design and development validation Validation design and development must be held in accordance with order which planned to ensure that results product capable fulfill conditions for application specific or meaning use. Validation must be completed before product delivery or implementation. Recording results validation and every action which is needed must be maintained. As part of validation design and development, organizations must do clinical and/or performance evaluation of medical devices, as required by national or regional regulations. If certain health devices can only be validated from the following assembly and installation on point use, delivery, not considered complete until product is submitted to the official customer. Provision about health devices for objective evaluation clinical and/or performance evaluation is not considered for submission. | Simplified requirements defined in Indonesia GMP
7.3.8 | Design and development transfer | 7.3.8 Design and development transfer The organization shall document procedures for transfer of design and development outputs to manufacturing. These procedures shall ensure that design and development outputs are verified as suitable for manufacturing before becoming final production specifications and that production capability can meet product requirements. Results and conclusions of the transfer shall be recorded (see 4.2.5).
7.3.9 | Control of design and development changes | 7.3.9 Control of design and development changes The organization shall document procedures to control design and development changes. The organization shall determine the significance of the change to function, performance, usability, safety and applicable regulatory requirements for the medical device and its intended use. Design and development changes shall be identified. Before implementation, the changes shall be: a) reviewed; b) verified; c) validated, as appropriate; d) approved. The review of design and development changes shall include evaluation of the effect of the changes on constituent parts and product in process or already delivered, inputs or outputs of risk management and product realization processes. Records of changes, their review and any necessary actions shall be maintained (see 4.2.5). | Appendix I III D9 | Design and development change control | 9. Design and development change control Change design and development must be identified and recording must be maintained. Change must be reviewed repeatedly, verified and validated if in accordance, and approved before implementation. Review design changes and development must cover evaluation effect change on the main part of products and products that have been marketed. Recording results of the review of changes and any actions required must be maintained. | Similar
7.3.10 | Design and development files | 7.3.10 Design and development files The organization shall maintain a design and development file for each medical device type or medical device family. This file shall include or reference records generated to demonstrate conformity to the requirements for design and development and records for design and development changes. | / | / | / | No specific DHF requirement defined in Indonesia GMP
7.4 | Purchasing |  | Appendix I III D10 | Purchase
7.4.1 | Purchasing process | 7.4.1 Purchasing process The organization shall document procedures (see 4.2.4) to ensure that purchased product conforms to specified purchasing information. The organization shall establish criteria for the evaluation and selection of suppliers. The criteria shall be: a) based on the supplier’s ability to provide product that meets the organization’s requirements; b) based on the performance of the supplier; c) based on the effect of the purchased product on the quality of the medical device; d) proportionate to the risk associated with the medical device. The organization shall plan the monitoring and re-evaluation of suppliers. Supplier performance in meeting requirements for the purchased product shall be monitored. The results of the monitoring shall provide an input into the supplier re-evaluation process. Non-fulfilment of purchasing requirements shall be addressed with the supplier proportionate to the risk associated with the purchased product and compliance with applicable regulatory requirements. Records of the results of evaluation, selection, monitoring and re-evaluation of supplier capability or performance and any necessary actions arising from these activities shall be maintained (see 4.2.5). | Appendix I III D10 a | Purchase process | a. Purchase process The company must establish mandatory procedures for ensuring that purchase material raw and/or component and/or material packer fulfill condition purchase which has been set. Type and stage of control applied to suppliers and product purchases should depend on effect from purchase product on deviation realization product or the final product. Companies must evaluate and choose suppliers based on ability to supply products in accordance with the condition of the company. Criteria selection, evaluation and re-evaluation must be set. Recording results evaluation and any required action arising from evaluation must be maintained. | Simplified requirements defined in Indonesia GMP Lack of risk control element in Indonesia GMP
7.4.2 | Purchasing information | 7.4.2 Purchasing information Purchasing information shall describe or reference the product to be purchased, including as appropriate: a) product specifications; b) requirements for product acceptance, procedures, processes and equipment; c) requirements for qualification of supplier personnel; d) quality management system requirements. The organization shall ensure the adequacy of specified purchasing requirements prior to their communication to the supplier. Purchasing information shall include, as applicable, a written agreement that the supplier notify the organization of changes in the purchased product prior to implementation of any changes that affect the ability of the purchased product to meet specified purchase requirements. To the extent required for traceability given in 7.5.9, the organization shall maintain relevant purchasing information in the form of documents (see 4.2.4) and records (see 4.2.5). | Appendix I III D10 b | Purchase information | b. Purchase information Information purchase must explain product which bought, where appropriate, including: 1) Condition for product receiving, procedure, process and equipment; 2) Requirements for personnel qualifications; And 3) Quality management system requirements. The company must guarantee the adequacy of the requirements purchase which is set before communication to the supplier. The stages required for traceability, the organization must look after information purchase which related, create document and recording. Information purchase stated in a manner written. Things which are included on information purchase can cover specification, picture, catalog, number code, date and place shipping, packaging, labels, certificates of analysis and test results. | Similar
7.4.3 | Verification of purchased product | 7.4.3 Verification of purchased product The organization shall establish and implement the inspection or other activities necessary for ensuring that purchased product meets specified purchasing requirements. The extent of verification activities shall be based on the supplier evaluation results and proportionate to the risks associated with the purchased product. When the organization becomes aware of any changes to the purchased product, the organization shall determine whether these changes affect the product realization process or the medical device. When the organization or its customer intends to perform verification at the supplier’s premises, the organization shall state the intended verification activities and method of product release in the purchasing information. Records of the verification shall be maintained (see 4.2.5). | Appendix I III D10 c | Purchased product verification | c. Purchased product verification Company must set and implement inspection or action other which needed to ensure that the product which bought a fulfill condition purchase set. If the company means to verify the supplier, the company must tell aspect- aspect which will verify and method release product in information purchase. Certificate analysis or results test must be available as supporters in activity verification. Verification records must be maintained. | The requirement of changes to the purchased product is not defined in Indonesia GMP
7.5 | Production and service provision |  | Appendix I III D11 | Production Process and Service Provision
7.5.1 | Control of production and service provision | 7.5.1 Control of production and service provision Production and service provision shall be planned, carried out, monitored and controlled to ensure that product conforms to specification. As appropriate, production controls shall include but are not limited to: a) documentation of procedures and methods for the control of production (see 4.2.4); b) qualification of infrastructure; c) implementation of monitoring and measurement of process parameters and product characteristics; d) availability and use of monitoring and measuring equipment; e) implementation of defined operations for labelling and packaging; f) implementation of product release, delivery and post-delivery activities. The organization shall establish and maintain a record (see 4.2.5) for each medical device or batch of medical devices that provides traceability to the extent specified in 7.5.9 and identifies the amount manufactured and amount approved for distribution. The record shall be verified and approved. | Appendix I III D11 a | Production control and after-sales service provision | a. Production control and after-sales service provision Company must plan and carry out production and provision service full sell in condition which under control, which includes: 1) Availability information which explain product characteristics; 2) Availability of documented procedures,documented requirements, work instructions, and material reference and measurement procedure reference; 3) Availability and use tool monitor and tool measuring; 4) All equipment which used in production must be checked before use; 5) All materials which used in in production must be checked before use; 6) All activity production must be held in accordance written procedures/work instructions; 7) Implementation release, activity distribution and activity after-sales; 8) Implementation of labeling and packaging procedures; And 9) All supervision during the process which is required must be accurately recorded at the time of execution. Company must set and look after recording For every tool batch which provides traceability to stage which has set and identifies the amount which produced and amount which approved for distribution. Batch records are mandatory, verified and approved. Completed batch/lot verification records include: 1) Identify employees who carry out and watch each step according to the date performed. 2) Identify the main equipment used. 3) Search batch/lot from whole material raw and material packer. 4) The amount of material used. 5) Packing notes during the process. 6) Record of irregularities (if any). 7) Track readiness records including cleanup records. 8) Example labels. | More comprehensive requirements are defined in Indonesia GMP, including requirment for batch record
7.5.2 | Cleanliness of product | 7.5.2 Cleanliness of product The organization shall document requirements for cleanliness of product or contamination control of product if: a) product is cleaned by the organization prior to sterilization or its use; b) product is supplied non-sterile and is to be subjected to a cleaning process prior to sterilization or its use; c) product cannot be cleaned prior to sterilization or its use, and its cleanliness is of significance in use; d) product is supplied to be used non-sterile, and its cleanliness is of significance in use; e) process agents are to be removed from product during manufacture. If product is cleaned in accordance with a) or b) above, the requirements contained in 6.4.1 do not apply prior to the cleaning process. | Appendix I III D11 b | Product hygiene and contamination control | b. Product hygiene and contamination control Condition environment in production room must be monitored and controlled because it can influence the product which will be produced. Size of work area and material storage area or product Which currently in process let adequate and possible placement equipment and material in a manner orderly and in accordance with the flow of the production process. Company must set condition for product hygiene if: 1) Product cleaned alone before sterilization and/or before use; 2) Non-sterile products supplied undergo processing cleaning before sterilization and/or use; 3) Non-sterile products which supplied for used and cleanliness is of utmost importance use; or 4) Substance which is used during making but must be removed from the finished product. | Similar
7.5.3 | Installation activities | 7.5.3 Installation activities The organization shall document requirements for medical device installation and acceptance criteria for verification of installation, as appropriate. If the agreed customer requirements allow installation of the medical device to be performed by an external party other than the organization or its supplier, the organization shall provide documented requirements for medical device installation and verification of installation. Records of medical device installation and verification of installation performed by the organization or its supplier shall be maintained (see 4.2.5). | Appendix I III D11 c | Installation activity | c. Installation activity, (if any) Companies must set a condition which contains acceptance criteria for installation and verification of equipment installation health. If the agreed customer requirements allow installation of the health device to be performed by external party besides from company or party appointed, the organization shall provide mandatory requirements for installation and verification. Recording installation and verification which is held by the company or by the party which has been appointed must be maintained. | Similar
7.5.4 | Servicing activities | 7.5.4 Servicing activities If servicing of the medical device is a specified requirement, the organization shall document servicing procedures, reference materials, and reference measurements, as necessary, for performing servicing activities and verifying that product requirements are met. The organization shall analyse records of servicing activities carried out by the organization or its supplier: a) to determine if the information is to be handled as a complaint; b) as appropriate, for input to the improvement process. Records of servicing activities carried out by the organization or its supplier shall be maintained (see 4.2.5). | Appendix I III D11 d | After-sales service activities | d. After-sales service activities, (if any) Companies must set necessary procedure, work instructions and material, reference procedure measurement, If necessary, carry out service activity and verification so that they fulfill the condition which set. Report activity service which are held by the company must be maintained. Service can include repair and maintenance. | No servicing activities analysis (complaint, improvement input) is defined in Indonesia GMP.
7.5.5 | Particular requirements for sterile medical devices | 7.5.5 Particular requirements for sterile medical devices The organization shall maintain records of the sterilization process parameters used for each sterilization batch (see 4.2.5). Sterilization records shall be traceable to each production batch of medical devices. | Appendix I III D11 e | Special Condition for health device non-sterile electromedical | e. Special Condition for health device non-sterile electromedical Companies must look after recording parameter process sterilization which has been used for every batch batch process sterilization. Recording sterilization must be traced For every batch of production batches of medical devices . | Similar
7.5.6 | Validation of processes for production and service provision | 7.5.6 Validation of processes for production and service provision The organization shall validate any processes for production and service provision where the resulting output cannot be or is not verified by subsequent monitoring or measurement and, as a consequence, deficiencies become apparent only after the product is in use or the service has been delivered. Validation shall demonstrate the ability of these processes to achieve planned results consistently. The organization shall document procedures for validation of processes, including: a) defined criteria for review and approval of the processes; b) equipment qualification and qualification of personnel; c) use of specific methods, procedures and acceptance criteria; d) as appropriate, statistical techniques with rationale for sample sizes; e) requirements for records (see 4.2.5); f) revalidation, including criteria for revalidation; g) approval of changes to the processes. The organization shall document procedures for the validation of the application of computer software used in production and service provision. Such software applications shall be validated prior to initial use and, as appropriate, after changes to such software or its application. The specific approach and activities associated with software validation and revalidation shall be proportionate to the risk associated with the use of the software, including the effect on the ability of the product to conform to specifications. Records of the results and conclusion of validation and necessary actions from the validation shall be maintained (see 4.2.4 and 4.2.5). | Appendix I III D11 f | Validation of production processes and service provision | f. Validation of production processes and service provision Companies must validate every process production and provision service if output can not be verified with monitoring and measurement sequentially. This includes certain processes, if the drawbacks are only seen after product used or service has been submitted. Validation must prove the ability process to achieve the planned results. Company must set order process, this including, if applicable: a) Criteria which defined for review repeat and process approval, b) Approval of tools and qualifications of personnel, c) Use of specific methods and procedures, d) Requirements for recording, and e) Revalidate Companies must set procedures for validation application device software (and software change and/or use) for provision production and service full sell which influences product performance. For fulfilling conditions, using this software must be validated before use. Recording validation must be maintained. | Simplified requirement for software validation in Indonesia GMP
7.5.7 | Particular requirements for validation of processes for sterilization and sterile barrier systems | 7.5.7 Particular requirements for validation of processes for sterilization and sterile barrier systems The organization shall document procedures (see 4.2.4) for the validation of processes for sterilization and sterile barrier systems. Processes for sterilization and sterile barrier systems shall be validated prior to implementation and following product or process changes, as appropriate. Records of the results and, conclusion of validation and necessary actions from the validation shall be maintained (see 4.2.4 and 4.2.5). NOTE Further information can be found in ISO 11607-1 and ISO 11607-2. | Appendix I III D11 g | Special requirements for sterile medical devices | g. Special requirements for sterile medical devices Company must set a procedure for the sterile process validation. Sterilization process must be validated before use and records validating each sterilization process shall be maintained .
7.5.8 | Identification | 7.5.8 Identification The organization shall document procedures for product identification and identify product by suitable means throughout product realization. The organization shall identify product status with respect to monitoring and measurement requirements throughout product realization. Identification of product status shall be maintained throughout production, storage, installation and servicing of product to ensure that only product that has passed the required inspections and tests or released under an authorized concession is dispatched, used or installed. If required by applicable regulatory requirements, the organization shall document a system to assign unique device identification to the medical device. The organization shall document procedures to ensure that medical devices returned to the organization are identified and distinguished from conforming product. | Appendix I III D11 h | Identification and traceability | 1) Identification Companies must identify products in a way that is appropriate for all products that have been realized, and must set procedures for identifying products which mean. Companies must set procedures documented to ensure that health devices which returned to the company have been identified and distinguished from products that meet standards 4) Status identification Companies must identify product status with apply condition monitoring and measurement. Identification status product must be maintained starting from production, storage, installation and service of products to ensure that only products which passed condition inspection and testing which were sent, used or installed (or released after going through the concession stage). | No significant difference
7.5.9 | Traceability
7.5.9.1 | General | 7.5.9.1 General The organization shall document procedures for traceability. These procedures shall define the extent of traceability in accordance with applicable regulatory requirements and the records to be maintained (see 4.2.5). | Appendix I III D11 h | Identification and traceability | 2) Traceability Companies must set procedures for traceability. The procedure in question must define the traceability scope of products and records which are required. If traceability is something required, the company must control and record identification of unique products. Management configuration is one means how identification and traceability can be maintained. | Similar
7.5.9.2 | Particular requirements for implantable medical devices | 7.5.9.2 Particular requirements for implantable medical devices The records required for traceability shall include records of components, materials, and conditions for the work environment used, if these could cause the medical device not to satisfy its specified safety and performance requirements. The organization shall require that suppliers of distribution services or distributors maintain records of the distribution of medical devices to allow traceability and that these records are available for inspection. Records of the name and address of the shipping package consignee shall be maintained (see 4.2.5). | Appendix I III D11 h | Identification and traceability | 3) Special condition for health device implant active and implant medical devices In set condition recording for capable traceability, the company must enclose recording all components, materials and condition environment work. If this can cause health devices not meeting the requirements that have been set. Company must require so that distributor look after recording distribution health device for easy traceability and recording is available for inspection. Name record and return address must be preserved. | Similar
7.5.10 | Customer Property | 7.5.10 Customer property The organization shall identify, verify, protect, and safeguard customer property provided for use or incorporation into the product while it is under the organization’s control or being used by the organization. If any customer property is lost, damaged or otherwise found to be unsuitable for use, the organization shall report this to the customer and maintain records (see 4.2.5). | Appendix I III D11 i | Customer Assets | i. Customer Assets Companies must look after goods owned by customers while in control of a company or used by organizations. The company must identify, verify, protect and protect owned by customers which were provided for use or put together in the product. If owned by a customer is lost, damaged or found useless, it must be reported to the customer and records are maintained. Customer assets can include money, intellectual or health information which must be kept secret. | Similar
7.5.11 | Preservation of product | 7.5.11 Preservation of product The organization shall document procedures for preserving the conformity of product to requirements during processing, storage, handling, and distribution. Preservation shall apply to the constituent parts of a medical device. The organization shall protect product from alteration, contamination or damage when exposed to expected conditions and hazards during processing, storage, handling, and distribution by: a) designing and constructing suitable packaging and shipping containers; b) documenting requirements for special conditions needed if packaging alone cannot provide preservation. If special conditions are required, they shall be controlled and recorded (see 4.2.5). | Appendix I III D11 j | Product security | j. Product security Companies must set procedures or work instructions to maintain product conformity during internal processes and delivery to the specified place. The safeguarding of this product shall include identification, handling, packaging, storage and protection. Product security must also apply for part components of the product. Companies must set procedure or work instructions for product control with age to use limited or special condition storage which is required. Condition storage is special which means it must be controlled and recorded. | Similar
7.6 | Control of monitoring and measuring equipment | 7.6 Control of monitoring and measuring equipment The organization shall determine the monitoring and measurement to be undertaken and the monitoring and measuring equipment needed to provide evidence of conformity of product to determined requirements. The organization shall document procedures to ensure that monitoring and measurement can be carried out and are carried out in a manner that is consistent with the monitoring and measurement requirements. As necessary to ensure valid results, measuring equipment shall: a) be calibrated or verified, or both, at specified intervals, or prior to use, against measurement standards traceable to international or national measurement standards: when no such standards exist, the basis used for calibration or verification shall be recorded (see 4.2.5); b) be adjusted or re-adjusted as necessary: such adjustments or re-adjustments shall be recorded (see 4.2.5); c) have identification in order to determine its calibration status; d) be safeguarded from adjustments that would invalidate the measurement result; e) be protected from damage and deterioration during handling, maintenance and storage. The organization shall perform calibration or verification in accordance with documented procedures. In addition, the organization shall assess and record the validity of the previous measuring results when the equipment is found not to conform to requirements. The organization shall take appropriate action in regard to the equipment and any product affected. Records of the results of calibration and verification shall be maintained (see 4.2.5). The organization shall document procedures for the validation of the application of computer software used for the monitoring and measurement of requirements. Such software applications shall be validated prior to initial use and, as appropriate, after changes to such software or its application. The specific approach and activities associated with software validation and revalidation shall be proportionate to the risk associated with the use of the software, including the effect on the ability of the product to conform to specifications. Records of the results and conclusion of validation and necessary actions from the validation shall be maintained (see 4.2.4 and 4.2.5). NOTE Further information can be found in ISO 10012. | Appendix I III D11 k | Control of quality testing tools and measuring instruments | k. Control of quality testing tools and measuring instruments Company must set monitoring and measurement which is done and tool monitoring and tool measuring which needed to prove suitability product to predetermined requirements. The company must establish mandatory procedures for ensuring that monitoring and measurement can be performed and performed consistently with requirements monitoring and measurement. If needed to ensure results which valid, tool measuring must: 1) Calibrated or verified on hose time certain, or before used, to standard measurement which traced to standard measurement international or national, If standard is not available, base reference which used for calibration or verification must be recorded; 2) Can be set or rearranged if needed; 3) Identified to make it easy in matter determine calibration status; 4) Protected from arrangement which can resulted invalid measurement results; 5) Protected from damage and disabled during handling, maintenance and storage. In addition, companies must evaluate and record validity results measurement previously if the tool is known does not meet the requirements. The company must take appropriate action on the equipment and any product involved. Recording and results calibration and verification must be maintained. If a software is used in monitoring and measurement so the ability software this must be ensured. | Similar
8 | Measurement, analysis and improvement |  | Appendix I III E | Measurement, Analysis and Improvement
8.1 | General | 8.1 General The organization shall plan and implement the monitoring, measurement, analysis and improvement processes needed to: a) demonstrate conformity of product; b) ensure conformity of the quality management system; c) maintain the effectiveness of the quality management system. This shall include determination of appropriate methods, including statistical techniques, and the extent of their use. |  |  | Company must plan and implement process measurement, analysis and improvement needed for: 1. Prove product conformity; 2. Ensure compliance with CPAKB; And 3. Maintain the effectiveness of CPAKB. | Similar
8.2 | Monitoring and measurement
8.2.1 | Feedback | 8.2.1 Feedback As one of the measurements of the effectiveness of the quality management system, the organization shall gather and monitor information relating to whether the organization has met customer requirements. The methods for obtaining and using this information shall be documented. The organization shall document procedures for the feedback process. This feedback process shall include provisions to gather data from production as well as post-production activities. The information gathered in the feedback process shall serve as potential input into risk management for monitoring and maintaining the product requirements as well as the product realization or improvement processes. If applicable regulatory requirements require the organization to gain specific experience from postproduction activities, the review of this experience shall form part of the feedback process. | Appendix I III E a | Measurement, control of nonconforming products and analysis | 1) Review of feedback Companies must monitor information related to fulfilling customer requirements. Companies must specify documented procedures for customer feedback systems, as a warning of early problem quality and input for the process of corrective and preventive action. There is an incident report that unexpected (KTD) on use of health devices to proceed for system feedback to watch out for . | No significant difference
8.2.2 | Complaint handling | 8.2.2 Complaint handling The organization shall document procedures for timely complaint handling in accordance with applicable regulatory requirements. These procedures shall include at a minimum requirements and responsibilities for: a) receiving and recording information; b) evaluating information to determine if the feedback constitutes a complaint; c) investigating complaints; d) determining the need to report the information to the appropriate regulatory authorities; e) handling of complaint-related product; f) determining the need to initiate corrections or corrective actions. If any complaint is not investigated, justification shall be documented. Any correction or corrective action resulting from the complaint handling process shall be documented. If an investigation determines activities outside the organization contributed to the complaint, relevant information shall be exchanged between the organization and the external party involved. Complaint handling records shall be maintained (see 4.2.5). | Appendix I III E d | Remediation | Companies must establish mandatory procedures for handling problems and warning notes implementation. Procedure must be implemented at any time. Recording all investigations of customer complaints must be maintained. If investigation determines that activity outside the company owes to customer complaints, related information must be exchanged between the companies involved. If every customer complaint does not follow with corrective and/or preventive action , the reason must be clear and recorded. | No significant difference
8.2.3 | Reporting to regulatory authorities | 8.2.3 Reporting to regulatory authorities If applicable regulatory requirements require notification of complaints that meet specified reporting criteria of adverse events or issuance of advisory notices, the organization shall document procedures for providing notification to the appropriate regulatory authorities. Records of reporting to regulatory authorities shall be maintained (see 4.2.5). |  |  | If national or regional regulations require notification to an incident which is not expected ( adverse event ) which has fulfilled defined reporting criteria, the company shall establish mandatory notification procedures intended for regulatory authority. | Similar
8.2.4 | Internal audit | 8.2.4 Internal audit The organization shall conduct internal audits at planned intervals to determine whether the quality management system: a) conforms to planned and documented arrangements, requirements of this International Standard, quality management system requirements established by the organization, and applicable regulatory requirements; b) is effectively implemented and maintained. The organization shall document a procedure to describe the responsibilities and requirements for planning and conducting audits and recording and reporting audit results. An audit program shall be planned, taking into consideration the status and importance of the processes and area to be audited, as well as the results of previous audits. The audit criteria, scope, interval and methods shall be defined and recorded (see 4.2.5). The selection of auditors and conduct of audits shall ensure objectivity and impartiality of the audit process. Auditors shall not audit their own work. Records of the audits and their results, including identification of the processes and areas audited and the conclusions, shall be maintained (see 4.2.5). The management responsible for the area being audited shall ensure that any necessary corrections and corrective actions are taken without undue delay to eliminate detected nonconformities and their causes. Follow-up activities shall include the verification of the actions taken and the reporting of verification results. NOTE Further information can be found in ISO 19011. | Appendix I III E a | Measurement, control of nonconforming products and analysis | 2) Internal audit Companies must do internal audits minimum once a year to determine if companies already apply CPAKB in accordance with the process which has been planned in an effective manner. Companies must appoint team auditors from each part to do audits internally. Program audits must be planned, with take into account status and important process and scope which audited, also results of previous audits. Audit criteria, scope, frequency and method must be set. Auditor selection and the conduct of the audit must ensure objectivity and impartiality process audits. Auditors cannot audit their own work. Responsibility and conditions for planning and implementation audit, and results report and recording which should be defined in mandatory procedures. Responsibility management for scope which is currently audited must ensure that action which is taken. Don’t slow down in a manner. Reasonable way to eliminate discrepancy which is found and the cause. Next activity must cover verification action which is taken and reporting of verification results. | No significant difference
8.2.5 | Monitoring and measurement of processes | 8.2.5 Monitoring and measurement of processes The organization shall apply suitable methods for monitoring and, as appropriate, measurement of the quality management system processes. These methods shall demonstrate the ability of the processes to achieve planned results. When planned results are not achieved, correction and corrective action shall be taken, as appropriate. | Appendix I III E a | Measurement, control of nonconforming products and analysis | /
8.2.6 | Monitoring and measurement of product | 8.2.6 Monitoring and measurement of product The organization shall monitor and measure the characteristics of the product to verify that product requirements have been met. This shall be carried out at applicable stages of the product realization process in accordance with the planned and documented arrangements and documented procedures. Evidence of conformity to the acceptance criteria shall be maintained. The identity of the person authorizing release of product shall be recorded (see 4.2.5). As appropriate, records shall identify the test equipment used to perform measurement activities. Product release and service delivery shall not proceed until the planned and documented arrangements have been satisfactorily completed. For implantable medical devices, the organization shall record the identity of personnel performing any inspection or testing. | Appendix I III E a | Measurement, control of nonconforming products and analysis | 4) Product monitoring and measurement Company must monitor and measure product characteristics to verify that the condition of the product has been fulfilled. It must be executed under the appropriate conditions of the product realization process in accordance with stages planning and mandatory procedures. Proof suitability with criteria which can be accepted must be maintained. Recording must identify the person who is responsible for releasing the product. Release product and free service is not necessarily to be processed until the planning stage is complete. For special conditions for active implantable and implantable health devices，companies must record identity personnel performing each inspection or test. | Similar
8.3 | Control of nonconforming product
8.3.1 | General | 8.3.1 General The organization shall ensure that product which does not conform to product requirements is identified and controlled to prevent its unintended use or delivery. The organization shall document a procedure to define the controls and related responsibilities and authorities for the identification, documentation, segregation, evaluation and disposition of nonconforming product. The evaluation of nonconformity shall include a determination of the need for an investigation and notification of any external party responsible for the nonconformity. Records of the nature of the nonconformities and any subsequent action taken, including the evaluation, any investigation and the rationale for decisions shall be maintained (see 4.2.5) | Appendix I III E b | Non-conforming product control | Companies must ensure that products which do not fulfill product conditions are identified and controlled, to prevent use or delivery which is not in accordance with intended. Control and resposibility and authority associated with non-conforming products should be defined in mandatory procedure. | Simplified requirement in Indonesia GMP
8.3.2 | Actions in response to nonconforming product detected before delivery | 8.3.2 Actions in response to nonconforming product detected before delivery The organization shall deal with nonconforming product by one or more of the following ways: a) taking action to eliminate the detected nonconformity; b) taking action to preclude its original intended use or application; c) authorizing its use, release or acceptance under concession. The organization shall ensure that nonconforming product is accepted by concession only if the justification is provided, approval is obtained and applicable regulatory requirements are met. Records of the acceptance by concession and the identity of the person authorizing the concession shall be maintained (see 4.2.5). |  |  | Company must handle products which are not in accordance with one or more of the following ways: (1) Take action to eliminate detected discrepancy; (2) With authorized use, release or accept products in accordance with applicable provisions; (3) Take action to prevent objective use or product application; (4) Companies must ensure that products which are not accepted with concession only if regulation conditions are fulfilled. Records of the identity of personnel authorized to permit must be maintained. Authenticity record, discrepancy and every action which is taken next, including concessions obtained, must be maintained. | Similar
8.3.3 | Actions in response to nonconforming product detected after delivery | 8.3.3 Actions in response to nonconforming product detected after delivery When nonconforming product is detected after delivery or use has started, the organization shall take action appropriate to the effects, or potential effects, of the nonconformity. Records of actions taken shall be maintained (see 4.2.5). The organization shall document procedures for issuing advisory notices in accordance with applicable regulatory requirements. These procedures shall be capable of being put into effect at any time. Records of actions relating to the issuance of advisory notices shall be maintained (see 4.2.5). |  |  | If non-conforming products are detected after delivery or after use, the company must take action which is appropriate to impact, or impact potential which is caused by nonconforming products. | No detailed documentation requirements are defined in Indonesia GMP
8.3.4 | Rework | 8.3.4 Rework The organization shall perform rework in accordance with documented procedures that takes into account the potential adverse effect of the rework on the product. These procedures shall undergo the same review and approval as the original procedure. After the completion of rework, product shall be verified to ensure that it meets applicable acceptance criteria and regulatory requirements. Records of rework shall be maintained (see 4.2.5). |  |  | If non-conforming products are corrected, the product must undergo verification to prove suitability with requirements. If products need done repeatedly (One time or more), the organization must document the rework process in the work instructions on their own right which is the same and approved procedure like original work instructions. Before authorizing and agreeing to the work instructions, determining any adverse effects resulting from work reproduction of non-conforming products must be made and documented. | No significant difference
8.4 | Analysis of data | 8.4 Analysis of data The organization shall document procedures to determine, collect and analyse appropriate data to demonstrate the suitability, adequacy and effectiveness of the quality management system. The procedures shall include determination of appropriate methods, including statistical techniques and the extent of their use. ﻿ 24 © ISO 2016 – All rights reserved Licensed to Roche Diagnostics GmbH - GLOBALNORM ILNAS eShop 2016 00756 / Max. Networking : 1 / downloaded : 2016-03-01 NOT FOR COMMERCIAL USE OR REPRODUCTION ﻿ ISO 13485:2016(E) The analysis of data shall include data generated as a result of monitoring and measurement and from other relevant sources and include, at a minimum, input from: a) feedback; b) conformity to product requirements; c) characteristics and trends of processes and product, including opportunities for improvement; d) suppliers; e) audits; f) service reports, as appropriate. If the analysis of data shows that the quality management system is not suitable, adequate or effective, the organization shall use this analysis as input for improvement as required in 8.5. Records of the results of analyses shall be maintained (see 4.2.5). | Appendix I III E c | Analysis | c. Analysis The company shall establish mandatory procedures to determine, collect and analyze data in accordance to prove the suitability and effectiveness of the quality management system. And to evaluate if the effectiveness of the quality management system can be improved. This includes the development of consequence data from monitoring and measurement and from other related sources. The analysis should provide information related to: a) Feedback; b) Conformity to product requirements; c) Characteristics and trend process and product including opportunities for preventive action; And d) Supplier. Records of data analysis results must be maintained. | No significant difference Only "Audit" is not included as data input in Indonesia GMP
8.5 | Improvement
8.5.1 | General | 8.5.1 General The organization shall identify and implement any changes necessary to ensure and maintain the continued suitability, adequacy and effectiveness of the quality management system as well as medical device safety and performance through the use of the quality policy, quality objectives, audit results, postmarket surveillance, analysis of data, corrective actions, preventive actions and management review. |  |  | Companies must identify and implement every change which is needed to ensure and maintain suitability and effectiveness of a sustainable quality management system through use of quality policy, quality objective, audit results, data analysis, corrective and preventive action and management review. | Similar
8.5.2 | Corrective action | 8.5.2 Corrective action The organization shall take action to eliminate the cause of nonconformities in order to prevent recurrence. Any necessary corrective actions shall be taken without undue delay. Corrective actions shall be proportionate to the effects of the nonconformities encountered. The organization shall document a procedure to define requirements for: a) reviewing nonconformities (including complaints); b) determining the causes of nonconformities; c) evaluating the need for action to ensure that nonconformities do not recur; d) planning and documenting action needed and implementing such action, including, as appropriate, updating documentation; e) verifying that the corrective action does not adversely affect the ability to meet applicable regulatory requirements or the safety and performance of the medical device; f) reviewing the effectiveness of corrective action taken. Records of the results of any investigation and of action taken shall be maintained (see 4.2.5). | Appendix I III E d | Remediation | 1) Corrective action Companies must take action to eliminate the reason for the discrepancy in the framework to prevent repeated discrepancy, corrective action which must match with impact caused by discrepancies found. Procedure must be set to define requirements for: a) Review non-conformities (including customer complaints); b) Determine the cause of the non-conformity; c) Evaluate need for action which will be taken to guarantee that non-conformities will not repeat; d) Determine and implement action which required, including, where appropriate, updates document; e) Record every investigation results and action which taken; And f) Review corrective actions taken and its effectiveness. | Similar
8.5.3 | Preventive action | 8.5.3 Preventive action The organization shall determine action to eliminate the causes of potential nonconformities in order to prevent their occurrence. Preventive actions shall be proportionate to the effects of the potential problems. The organization shall document a procedure to describe requirements for: a) determining potential nonconformities and their causes; b) evaluating the need for action to prevent occurrence of nonconformities; c) planning and documenting action needed and implementing such action, including, as appropriate, updating documentation; d) verifying that the action does not adversely affect the ability to meet applicable regulatory requirements or the safety and performance of the medical device; e) reviewing the effectiveness of the preventive action taken, as appropriate. Records of the results of any investigations and of action taken shall be maintained (see 4.2.5). | Appendix I III E d | Remediation | 2) Preventive action Companies must take action to eliminate the reason for discrepancy in the framework to prevent repeated discrepancy. Corrective action which is taken must be in accordance with the impact which is caused by potential problems. Procedure must be set to define requirements for: a) Determine potential nonconformities and their causes; b) Evaluate need for action to prevent the emergence of non-conformities; c) Determine and implement action which needed; d) Record results for every investigation and action which taken; And e) Review preventive action which was taken and its effectiveness. | No significant difference

## ISO 13485& CHINA
ISO 13485: 2016 | CHINA
ISO 13485: 2016 | 《医疗器械生产质量管理规范》 | GAP
本标准为需要证实自身有能力提供持续满足顾客要求和适用的法规要求的医疗器械和相关服务的组织规定了质量管理体系要求。这类组织能涉及医疗器械生命周期的一个或多个阶段，包括医疗器械的设计和开发、生产、贮存和流通、安装或服务，以及相关活动（例如技术支持）的设计和开发或提供。本标准也能用于向这类组织提供产品（包括与质量管理体系相关的服务）的供方或外部方。 | 第一条为保障医疗器械安全、有效，规范医疗器械生产质量管理，根据《医疗器械监督管理条例》（国务院令第650号）、《医疗器械生产监督管理办法》（国家食品药品监督管理总局令第7号），制定本规范。
除非明确规定，本标准的要求适用于各种规模和类型的组织。本标准中应用于医疗器械的要求同样适用于组织提供的相关服务。对于本标准所要求的适用于组织但不是由组织实施的过程，在质量管理体系中组织通过监视、维护和控制这些过程对其负有责任。 | 第二条医疗器械生产企业（以下简称企业）在医疗器械设计开发、生产、销售和售后服务等过程中应当遵守本规范的要求。
如果适用的法规要求允许对设计和开发控制进行删减，则能作为在质量管理体系中将其删减的理由。若这些法规要求能提供其他方法，这些方法要在质量管理体系中予以说明。组织有责任确保在符合本标准的声明中明确对设计和开发控制的任何删减。 | 第八十一条企业可根据所生产医疗器械的特点，确定不适用本规范的条款，并说明不适用的合理性。
本标准第6、7或8章中的任何要求，如果因组织开展的活动或质量管理体系所涉及的医疗器械的特点而不适用时，组织不需要在其质量管理体系中包含这样的要求。对于经确定不适用的任何条款，组织按照4.2.2的要求记录其理由。 | 同上
4 质量管理体系 | 第三条企业应当按照本规范的要求，结合产品特点，建立健全与所生产医疗器械相适应的质量管理体系，并保证其有效运行。
4.1 总要求
4.1.1 组织应按照本标准的要求和适用的法规要求将质量管理体系形成文件并保持其有效性。
组织应按照本标准或者适用的法规要求建立、实施和保持需要形成的文件的所有要求、程序、活动或安排。
组织应将其在适用的法规要求下所承担的一个或多个角色
形成文件。
注：组织所承担的角色可能包括制造商、授权代表、进口商或经销商。
4.1.2 组织应： | 第四条企业应当将风险管理贯穿于设计开发、生产、销售和售后服务等全过程，所采取的措施应当与产品存在的风险相适应。
a） 考虑组织承担的角色来确定质量管理体系所需的过程及这些过程在整个组织中的应用；
b） 应用基于风险的方法控制质量管理体系所需的适当过程
c） 确定这些过程的顺序和相互作用。
4.1.3 对于每个质量管理体系过程，组织应： | /
a）确定所需的准则和方法，以确保这些过程的有效运行和控制。
b）确保可获得必要的资源和信息，以支持这些过程的运行和对这些过程的监视
c）实施必要的措施，以实现这些过程策划的结果并保持这些过程的有效性
d）监视、测量（适当时）和分析这些过程
e）建立和保持所需的记录以证实符合本标准并满足适用的法规要求（见4.2.5）
4.1.4 组织应按照本标准要求和适用的法规要求管理这些质量管理体系过程。更改这些过程应： | /
a）评价过程更改对质量管理体系的影响
b）评价过程更改对该质量管理体系中所生产的医疗器械的影响
c）按照本标准的要求和适用的法规要求进行控制
4.1.5 若组织选择将影响产品符合要求的任何过程外包，组织应监视这类过程并确保对其控制。组织应保留外包过程符合本标准要求、顾客要求和适用的法规要求的责任。控制应与所涉及的风险和外部方满足7.4中要求的能力相适应。控制应包括书面质量协议。 | /
4.1.6 组织应将用于质量管理体系的计算机软件应用的确认程序形成文件。在软件首次使用前应对软件应用进行确认，适当时，软件或其应用更改后也应对软件应用进行确认。 | /
与软件确认和再确认有关的特定方法和活动应与软件使用有关的风险相适应。
应保留这些活动的记录（见4.2.5）。
4.2 文件要求 | 第二十四条企业应当建立健全质量管理体系文件，包括质量方针和质量目标、质量手册、程序文件、技术文件和记录,以及法规要求的其他文件。
4.2.1 总则
质量管理体系文件（见4.2.4）应包括：
a） 形成文件的质量方针和质量目标；
b） 质量手册；
c） 本标准所要求的形成文件的程序和记录；
d） 组织确定的为确保其过程有效策划、运行和控制所需的文件，包括记录；
e） 适用的法规要求规定的其他文件。
4.2.2 质量手册 | 第二十四条质量手册应当对质量管理体系作出规定。程序文件应当根据产品生产和质量管理过程中需要建立的各种工作程序而制定，包含本规范所规定的各项程序。
组织应编制质量手册，质量手册包括：
a） 质量管理体系的范围，包括任何删减不适用的详细说明和理由；
b） 质量管理体系的形成文件的程序或对其引用；
c） 质量管理体系过程之间的相互作用的表述。
质量手册应概述质量管理体系的文件结构。
4.2.3 医疗器械文档 | 第二十四条技术文件应当包括产品技术要求及相关标准、生产工艺规程、作业指导书、检验和试验操作规程、安装和服务操作规程等相关文件。
组织应为每个医疗器械类型或医疗器械族建立并保持一个或多个文档，文档包含或引用形成的文件以证明符合本标准要求和适用的法规要求。
文档的内容应包括但不限于：
a） 医疗器械的概述、预期用途/预期目的和标记，包括所有使用说明；
b） 产品规范；
c） 制造、包装、贮存、处置和流通的规范或程序；
d） 测量和监视程序；
e） 适当时，安装要求；
f） 适当时，服务程序。
4.2.4 文件控制 | 第二十五条企业应当建立文件控制程序，系统地设计、制定、审核、批准和发放质量管理体系文件，至少应当符合以下要求：（一）文件的起草、修订、审核、批准、替换或者撤销、复制、保管和销毁等应当按照控制程序管理，并有相应的文件分发、替换或者撤销、复制和销毁记录；（二）文件更新或者修订时，应当按规定评审和批准，能够识别文件的更改和修订状态；（三）分发和使用的文件应当为适宜的文本，已撤销或者作废的文件应当进行标识，防止误用。
质量管理体系所要求的文件应予控制。记录是一种特殊类型的文件，应依据4.2.5的要求进行控制。 | 第二十六条企业应当确定作废的技术文件等必要的质量管理体系文件的保存期限，以满足产品维修和产品质量责任追溯等需要。
形成文件的程序应规定以下方面所需的控制：
a） 为使文件充分和适宜，文件发布前得到评审和批准；
b） 必要时对文件进行评审与更新，并再次批准；
c） 确保文件的现行修订状态和更改得到识别；
d） 确保在使用处可获得适用文件的有关版本；
e） 确保文件保持清晰、易于识别；
f） 确保组织所确定的策划和运行质量管理体系所需的外来文件得到识别，并控制其分发；
g） 防止文件的损坏或丢失；
h） 防止作废文件的非预期使用，对这些文件进行适当的标识。
组织应确保文件的更改得到原审批部门或指定的其他审批部门的评审和批准，被指定的审批部门应能获取用于做出决定的相关背景资料。
对于至少应保存一份的作废文件，组织应规定其保存期限。此期限应确保至少在组织所规定的医疗器械寿命期内，可得到这些医疗器械的制造和试验的文件，而且还应不少于记录（见4.2.5）或适用的法规要求所规定的保存期限。
4.2.5 记录控制 | 第二十七条企业应当建立记录控制程序，包括记录的标识、保管、检索、保存期限和处置要求等，并满足以下要求：（一）记录应当保证产品生产、质量控制等活动的可追溯性；（二）记录应当清晰、完整，易于识别和检索，防止破损和丢失；（三）记录不得随意涂改或者销毁，更改记录应当签注姓名和日期，并使原有信息仍清晰可辨，必要时，应当说明更改的理由；（四）记录的保存期限应当至少相当于企业所规定的医疗器械的寿命期，但从放行产品的日期起不少于2年，或者符合相关法规要求，并可追溯。
应保持记录以提供符合要求和质量管理体系有效运行的证据。
组织应建立程序并形成文件，以规定记录的标识、存储、安全和完整性、检索、保留时间和处置所需的控制。
按照适用的法规要求，组织应对记录中包含的保密健康信息规定并实施保护方法。
记录应保持清晰、易于识别和检索。记录的更改应保持可识别。
组织应保持记录的期限至少为组织所规定的或适用的法规要求所规定的医疗器械的寿命期，而且还应从组织放行医疗器械起不少于两年。
5 管理职责 | 第六条企业负责人是医疗器械产品质量的主要责任人，应当履行以下职责：（一）组织制定企业的质量方针和质量目标；（二）确保质量管理体系有效运行所需的人力资源、基础设施和工作环境等；（三）组织实施管理评审，定期对质量管理体系运行情况进行评估，并持续改进；（四）按照法律、法规和规章的要求组织生产。
5.1 管理承诺
最高管理者应通过以下活动，对其建立、实施质量管理体系并保持其有效性的承诺提供证据；
a） 向组织传达满足顾客要求以及适用的法规要求的重要性；
b） 指定质量方针；
c） 确保制定质量目标；
d） 进行管理评审；
e） 确保资源的可获得性。
5.2 以顾客为关注焦点 | /
最高管理者应确保顾客要求和适用的法规要求得到确定和满足。
5.3 质量方针 | /
最高管管理者应确保质量方针：
a) 适应组织的宗旨；
b) 包括对满足要求和保持质量管理体系有效性的承诺；
c) 为制定和评审质量目标提供框架；
d) 在组织内得到沟通和理解；
e) 在持续适宜性方面得到评审。
5.4 策划 | /
5.4.1 质量目标
最高管理者应确保在组织的相关职能和层次上建立质量目标，质量目标包括满足适用的法规要求和产品要求所需的内容。质量目标应是可测量的，并与质量方针保持一致。
5.4.2 质量管理体系策划 | /
最高管理者应确保：
a) 对质量管理体系进行策划，以满足4.1的要求以及质量目标；
b) 在策划和实施质量管理体系变更时保持其完整性。
5.5 职责、权限与沟通 | 第五条企业应当建立与医疗器械生产相适应的管理机构，并有组织机构图，明确各部门的职责和权限，明确质量管理职能。生产管理部门和质量管理部门负责人不得互相兼任。
5.5.1 职责和权限
最高管理者应确保职责和权限得到规定、形成文件并在组织内沟通
最高管理者应将所有从事对质量有影响的管理、执行和验证工作的人员的相互关系形成文件，并应确保其完成这些任务所必要的独立性和权限。
5.5.2 管理者代表 | 第七条企业负责人应当确定一名管理者代表。管理者代表负责建立、实施并保持质量管理体系，报告质量管理体系的运行情况和改进需求，提高员工满足法规、规章和顾客要求的意识。
最高管理者应在管理层中指定一名成员，无论该成员在其他方面的职责如何，应使其具有以下方面的职责和权限：
a) 确保将质量管理体系所需的过程形成文件；
b) 向最高管理者报告质量管理体系的有效性和任何改进的需求；
c) 确保在整个组织内提高满足适用的法规要求和质量管理体系要求的意识。
5.5.3 内部沟通 | /
最高管理者应确保在组织内建立适当的沟通过程，并确保对质量管理体系的有效性进行沟通。
5.6 管理评审 | 第七十八条企业应当定期开展管理评审，对质量管理体系进行评价和审核，以确保其持续的适宜性、充分性和有效性。
5.6.1 总则
组织应将管理评审程序形成文件。最高管理者应按照形成文件的策划的时间间隔对组织的质量管理体系进行评审，以确保其持续的适宜性、充分性和有效性。评审应包括评价改进的机会和质量管理体系变更的需求，包括质量方针和质量目标变更的需求。
应保留管理评审的记录（见4.2.5）
5.6.2 评审输入 | /
管理评审的输入应包括但不限于由以下方面产生的信息：
a) 反馈；
b) 投诉处置；
c) 给监管机构的报告；
d) 审核；
e) 过程的监视和测量；
f) 产品的监视和测量；
g) 纠正措施；
h) 预防措施
i) 以往管理评审的跟踪措施；
j) 可能影响质量管理体系的变更；
k) 改进的建议；
l) 适用的新的或修订的法规要求。
5.6.3 评审输出 | /
管理评审的输出应予以记录（见4.2.5）并包括经评审的输入和以下方面有关的任何决定和措施：
a) 保持质量管理体系及其过程适宜性、充分性和有效性所
需的改进；
b) 与顾客要求有关的产品的改进；
c) 响应适用的新的或修订的法规要求所需的变更；
d) 资源需求。
6 资源管理 | /
6.1 资源提供
组织应确定并提供所需的资源，以：
a) 实施质量管理体系并保持其有效性；
b) 满足适用的法规要求和顾客要求。
6.2 人力资源 | 第八条技术、生产和质量管理部门的负责人应当熟悉医疗器械相关法律法规，具有质量管理的实践经验，有能力对生产管理和质量管理中的实际问题作出正确的判断和处理。
基于适当的教育、培训、技能和经验，从事影响产品质量工作的人员应是胜任的。 | 第九条企业应当配备与生产产品相适应的专业技术人员、管理人员和操作人员，具有相应的质量检验机构或者专职检验人员。
组织应将确立能力、提供所需的培训和确保人员的意识等一个或多个过程形成文件。 | 第十条从事影响产品质量工作的人员，应当经过与其岗位要求相适应的培训，具有相关理论知识和实际操作技能。
组织应：
a) 确定从事影响产品质量工作的人员所需具备的能力；
b) 提供培训或采取其他措施以获得或保持所需的能力；
c) 评价所采取措施的有效性；
d) 确保组织的人员知晓所从事活动的关联性和重要性，以及如何为实现质量目标做出贡献；
e) 保留教育、培训、技能和经验的适当记录（见4.2.5）
注：对于提供培训或采取其他措施的有效性的检查方法应与工作相关的风险相适应。
6.3 基础设施 | 第十二条厂房与设施应当符合生产要求，生产、行政和辅助区的总体布局应当合理，不得互相妨碍。
为达到符合产品要求、防止产品混淆和确保产品有序处置，组织应将所需的基础设施的要求形成文件。适当时，基础设施包括： | 第十四条厂房应当确保生产和贮存产品质量以及相关设备性能不会直接或者间接受到影响，厂房应当有适当的照明、温度、湿度和通风控制条件。
a) 建筑物、工作场所和相关设施； | 第十六条
b) 过程设备（硬件和软件）； | 生产区应当有足够的空间，并与其产品
c) 支持性服务（如运输、通讯或信息系统）。 | 生产规模、品种相适应。
若维护活动或缺少维护活动可能影响产品质量，组织应将此 | 第十七条
类维护活动的要求 包括执行维护活动的时间间隔形成文件。 | 仓储区应当能够满足原材料、包装材料、
适当时，要求应适用于生产设备、工作环境控制设备和监视 | 中间品、产品等的贮存条件和要求，按照待验、合
测量设备。 | 格、不合格、退货或者召回等情形进行分区存放，
应保留此类维护的记录（见4.2.5）。 | 便于检查和监控。
第十八条
企业应当配备与产品生产规模、品种、检
验要求相适应的检验场所和设施。
第十九条
企业应当配备与所生产产品和规模相匹
配的生产设备、工艺装备等，并确保有效运行。
第二十条
生产设备的设计、选型、安装、维修和维
护必须符合预定用途，便于操作、清洁和维护。生产设备应当有明显的状态标识，防止非预期使用。企业应当建立生产设备使用、清洁、维护和维修的操作规程，并保存相应的操作记录。
第二十一条
企业应当配备与产品检验要求相适应的检验仪器和设备，主要检验仪器和设备应当具有明确的操作规程。
第二十二条
企业应当建立检验仪器和设备的使用记录，记录内容包括使用、校准、维护和维修等情况。
第二十三条
企业应当配备适当的计量器具。计量器具的量程和精度应当满足使用要求，标明其校准有效期，并保存相应记录。
6.4 工作环境和污染控制 | 第十一条从事影响产品质量工作的人员，企业应当对其健康进行管理，并建立健康档案。
6.4.1 工作环境 | 第十三条厂房与设施应当根据所生产产品的特性、工艺流程及相应的洁净级别要求合理设计、布局和使用。生产环境应当整洁、符合产品质量需要及相关技术标准的要求。产品有特殊要求的，应当确保厂房的外部环境不能对产品质量产生影响，必要时应当进行验证。
组织应将为达到符合产品要求所需工作环境的要求形成文件。 | 第十五条
如果工作环境条件可能对产品质量有不良的影响，组织应将工作环境要求以及监视和控制工作环境的程序形成文件。 | 厂房与设施的设计和安装应当根据产品
组织应： | 特性采取必要的措施，有效防止昆虫或者其他动物进入。对厂房与设施的维护和维修不得影响产品质量。
a）将特定人员的健康、清洁和着装要求形成文件，此类人员与产品或工作环境的接触可能影响医疗器械的安全或性能；
b）确保需要在工作环境内的特殊环境条件下临时工作的所有人员是胜任的或在胜任人员监督下工作。
注：更多信息见ISO14644和ISO14698。
6.4.2 污染控制 | /
适当时，为了防止对工作环境、人员或产品的污染，组织对受污染或易受污染产品的控制应进行策划并将安排形成文件。
对于无菌医疗器械，组织应将控制微生物或微粒物污染的要求形成文件，在组装或包装过程中保持所要求的洁净度。
7 产品实现 | 第三十八条企业应当在包括设计和开发在内的产品实现全过程中，制定风险管理的要求并形成文件，保持相关记录。
7.1 产品实现的策划
组织应策划和开发产品实现所需的过程。产品实现的策划应与质量管理体系其他过程的要求相一致。
组织应在产品实现过程中，将风险管理的一个或多个过程形成文件。应保留风险管理活动的记录（见4.2.5）。
在策划产品实现的过程中，适当时，组织应确定以下方面的内容：
a） 产品的质量目标和要求；
b） 针对产品建立过程、文件（见4.2.4）和提供资源的需求，包括基础设施和工作环境；
c） 针对产品所要求的验证、确认、监视、测量、检验和试验、处置、贮存、流通和可追溯性活动以及产品接收准则；
d） 为实现过程及其产品满足要求提供证据所需的记录（见4.2.5）。
此策划的输出应以适合于组织运行方式的形式形成文件。
注：更多信息见ISO14971.
7.2 与顾客相关的过程 | /
7.2.1 产品要求的确定
组织应确定：
a） 顾客规定的要求，包括对交付及交付后活动的要求；
b） 顾客虽然没有明示，但规定的用途或已知的预期用途所必需的要求；
c） 与产品有关的适用的法规要求；
d） 确保医疗器械的特定性能和安全使用所需的任何用户培训；
e） 组织确定的任何附加要求。
7.2.2 产品要求的评审 | 第六十二条 企业应当建立产品销售记录，并满足可追溯的要求。销售记录至少包括医疗器械的名称、规格、型号、数量；生产批号、有效期、销售日期、购货单位名称、地址、联系方式等内容。
组织应评审与产品有关的要求。评审应在组织向顾客做出提供产品的承诺（如：提交投标书、接受合同或订单以及接受合同或订单的更改）前进行并应确保； | 第六十三条直接销售自产产品或者选择医疗器械经营企业，应当符合医疗器械相关法规和规范要求。发现医疗器械经营企业存在违法违规经营行为时，应当及时向当地食品药品监督管理部门报告。
a) 产品要求已得到规定并形成文件；
b) 与以前表述不一致的合同或订单要求已得到解决；
c) 满足适用的法规要求；
d) 依照7.2.1识别的任何用户培训是可获得的或按计划可获得的；
e) 组织有能力满足规定的要求。
应保留评审结果及评审所引起的措施的记录（见4.2.5）。
若顾客没有提供形成文件的要求，组织在接受顾客要求前应对顾客要求进行确认。
若产品要求发生变更，组织应确保相关文件得到修改，并确保相关人员知道已更改的要求。
7.2.3 沟通 | 第六十六条 企业应当建立顾客反馈处理程序，对顾客反馈信息进行跟踪分析。
组织应就以下方面与顾客的沟通进行策划并将安排形成文件
a） 产品信息；
b） 处理问询、合同或订单，包括更改；
c） 顾客反馈，包括投诉；
d） 忠告性通知。
组织应按照适用的法规要求与监督机构沟通。
7.3 设计和开发 | 第二十八条企业应当建立设计控制程序并形成文件，对医疗器械的设计和开发过程实施策划和控制。
7.3.1 总则
组织应将设计和开发程序形成文件。
7.3.2 设计和开发策划 | 第二十九条在进行设计和开发策划时，应当确定设计和开发的阶段及对各阶段的评审、验证、确认和设计转换等活动，应当识别和确定各个部门设计和开发的活动和接口，明确职责和分工。
组织应对产品的设计和开发进行策划和控制。适当时，随着设计和开发的进展，应保持并更新设计和开发策划文件。
在设计和开发策划期间，组织应将以下方面形成文件：
a） 设计可开发的各个阶段；
b） 每个设计和开发阶段所需的一个或多个评审；
c） 适合于每个设计和开发阶段的验证、确认和设计转换活动；
d） 设计和开发的职责和权限；
e） 确保设计和开发输出到设计和开发输入的可追溯的方法；
f） 所需的资源，包括必要的人员能力。
7.3.3 设计和开发输入 | 第三十条设计和开发输入应当包括预期用途规定的功能、性能和安全要求、法规要求、风险管理控制措施和其他要求。对设计和开发输入应当进行评审并得到批准，保持相关记录。
应确定与产品要求有关的输入，并保留记录（见4.2.5），这些输入应包括：
a） 根据预期用途所确定的功能、性能、可用性和安全要求；
b） 适用的法规要求和标准；
c） 适用的风险管理的一个或多个输出；
d） 适当时，来源于以前类似设计的信息；
e） 产品和过程的设计和开发所必需的其他要求。
应对这些输入进行评审，以确保输入是充分和适宜的，并经批准。
这些要求应完整、清楚，能够被验证或确认，并且不能相互矛盾。
注：更多信息见IEC 62366-1
7.3.4 设计和开发输出 | 第三十一条设计和开发输出应当满足输入要求，包括采购、生产和服务所需的相关信息、产品技术要求等。设计和开发输出应当得到批准，保持相关记录。
设计和开发输出应：
a) 满足设计和开发输入的要求；
b) 给出采购、生产和服务提供的适当信息；
c) 包括或引用产品接收准则；
d) 规定产品特性，该特性对于产品的安全和正确使用是必需的。
设计和开发输出的方式应适合于对照设计和开发输入进行验证，设计和开发输出应在发布前得到批准。
应保留设计和开发输出的记录（见4.2.5）
7.3.5 设计和开发评审 | 第三十三条企业应当在设计和开发的适宜阶段安排评审，保持评审结果及任何必要措施的记录。
应依据所策划并形成文件的安排，在适宜的阶段对设计和开发进行系统评审，以：
a) 评价设计和开发的结果满足要求的能力；
b) 识别并提议必要的措施。
评审的参加者包括与所评审的设计和开发阶段有关的职能的代表以及其他专业人员。应保留评审结果和任何必要措施的记录，包括所评审的设计、涉及的参加者和评审日期（见4.2.5）。
7.3.6 设计和开发验证 | 第三十四条企业应当对设计和开发进行验证，以确保设计和开发输出满足输入的要求，并保持验证结果和任何必要措施的记录。
为确保设计和开发输出满足设计和开发输入的要求，应依据所策划并形成文件的安排对设计和开发进行验证。
组织应将验证计划形成文件，验证计划包括方法、接收准则，适当时包括包含样本量原理的统计技术。
如果预期用途要求医疗器械连接至或通过接口连接至其他
一个或多个医疗器械，验证应包括证实当这样连接或通过接
口连接时设计输出满足设计输入的要求。
应保留验证结果和结论及必要措施的记录（见
4.2.4和 4.2.5）。
7.3.7 设计和开发确认 | 第三十五条企业应当对设计和开发进行确认，以确保产品满足规定的使用要求或者预期用途的要求，并保持确认结果和任何必要措施的记录。
为确保产品能够满足规定的应用要求或预期用途要求，应依据策划并形成文件的安排对设计和开发进行确认。 | 第三十六条确认可采用临床评价或者性能评价。进行临床试验时应当符合医疗器械临床试验法规的要求。
组织应将确认计划形成文件，确认计划包括方法、接收准则，适当时包括包含样本量原理的统计技术。
设计确认应选择有代表性产品进行。有代表性产品包括最初的生产单元、批次或其等同品。应记录用于确认的产品选择是理由说明（见4.2.5）。
作为设计和开发确认的一部分，组织应按照适用的法规要求进行医疗器械临床评价或性能评价。用于临床评价或性能评价的医疗器械不视为放行给顾客使用。
如果预期用途要求医疗器械连接至或通过接口连接至其他一个或多个医疗器械，确认应包括证实当这样连接或通过接口连接时已满足规定的应用要求或预期用途。
确认应在向顾客放行产品使用前完成。
应保留确认应结果和结论及必要措施的记录（见4.2.4和4.2.5）。
7.3.8 设计和开发转换 | 第三十二条企业应当在设计和开发过程中开展设计和开发到生产的转换活动，以使设计和开发的输出在成为最终产品规范前得以验证，确保设计和开发输出适用于生产。
组织应将设计和开发输出向制造转换的程序形成文件。这些程序应确保设计和开发输出在成为最终生产规范前经验证适合于制造并确保生产能力能满足产品要求。
应记录转换的结果和结论（见4.2.5）。
7.3.9 设计和开发更改的控制 | 第三十七条企业应当对设计和开发的更改进行识别并保持记录。必要时，应当对设计和开发更改进行评审、验证和确认，并在实施前得到批准。
组织应将控制设计和开发更改的程序形成文件。组织应确定更改对于医疗器械功能、性能、可用性、安全、适用的法规要求及其预期用途等的重要程度。 | 选用的材料、零件或者产品功能的改变可能影响到医疗器械产品安全性、有效性时，应当评价因改动可能带来的风险，必要时采取措施将风险降低到可接受水平，同时应当符合相关法规的要求。
应识别设计和开发的更改。更改在实施前应经：
a) 评审；
b) 验证；
c) 适当时，确认；
d) 批准。
设计和开发更改的评审应包括评价更改对在制的或已交付
的组成部件和产品的影响，以及对风险管理的输入或输出和
产品实现过程的影响。
应保留更改及其评审和任何必要的措施的记录（见
4.2.5）。
7.3.10 设计和开发文档 | /
组织应为每个医疗器械类型或医疗器械族保留设计和开发文档。该文档应包含或引用形成的记录以证明符合设计和开发要求，该文档还应包含设计和开发更改的记录。
7.4 采购 | 第三十九条企业应当建立采购控制程序，确保采购物品符合规定的要求，且不低于法律法规的相关规定和国家强制性标准的相关要求。
7.4.1 采购过程 | 第四十条企业应当根据采购物品对产品的影响，确定对采购物品实行控制的方式和程度。
组织应将确保采购的产品符合规定的采购信息的程序形成文件（见4.2.4）。 | 第四十一条企业应当建立供应商审核制度，并应当对供应商进行审核评价。必要时，应当进行现场审核。
组织应建立评价和选择供方的准则。准则应：
a) 基于供方提供满足组织要求的产品的能力；
b) 基于供方的绩效；
c) 基于采购产品对医疗器械质量的影响；
d) 与医疗器械相关风险相适应。
组织应对供方的监视和再评价进行策划。应监视供方满足采购产品的要求的绩效。监视结果应为供方再评价过程提供输入。
对未实现采购要求的供方的处置应与所采购产品有关的风险相适应，并符合使用的法规要求。
应保留供方能力或绩效的评价、选择、监视和再评价的结果及由这些活动所引起的任何必要措施的记录（见4.2.5）。
7.4.2 采购信息 | 第四十二条企业应当与主要原材料供应商签订质量协议，明确双方所承担的质量责任。
拟采购产品的采购信息应表述或引用，适当时包括： | 第四十三条采购时应当明确采购信息，清晰表述采购要求，包括采购物品类别、验收准则、规格型号、规程、图样等内容。应当建立采购记录，包括采购合同、原材料清单、供应商资质证明文件、质量标准、检验报告及验收标准等。采购记录应当满足可追溯要求。
a) 产品规范；
b) 产品接收、程序、过程和设备的要求；
c) 供方人员资格要求；
d) 质量管理体系要求。
组织应确保在与供方沟通前所规定的采购要求是充分和适宜的。
适用时，采购信息应包括书面协议，该协议明确了在影响采购产品满足规定的采购要求的能力的任何更改实施前，供方应将采购产品方面的更改通知组织。
按照7.5.9规定的可追溯性要求的范围和程度，组织应以文件（见4.2.4）和记录（见4.2.5）的形式保持相关采购信息。
7.4.3 采购产品的验证 | 第四十四条企业应当对采购物品进行检验或者验证，确保满足生产要求。
组织应建立并实施检验或其他必要的活动，以确保采购的产品满足规定的采购要求。验证活动的范围应基于供方评价结果，并与采购产品有关的风险相适应。
若组织觉察到采购产品的任何更改，组织应确定这些更改是否影响产品实现过程或医疗器械最终产品。
若组织或其顾客拟在供方的现场实施验证，组织应在采购信息中对拟验证的活动和产品放行方法做出规定。
应保留验证记录（见4.2.5）。
7.5 生产和服务提供 | 第四十五条企业应当按照建立的质量管理体系进行生产，以保证产品符合强制性标准和经注册或者备案的产品技术要求。
7.5.1 生产和服务提供的控制 | 第四十六条企业应当编制生产工艺规程、作业指导书等，明确关键工序和特殊过程。
生产和服务提供应予策划、实施、监视和控制以确保产品符合规范。适当时，生产控制应包括但不限于： | 第五十条每批（台）产品均应当有生产记录，并满足可追溯的要求。生产记录包括产品名称、规格型号、原材料批号、生产批号或者产品编号、生产日期、数量、主要设备、工艺参数、操作人员等内容。
a） 编制生产控制程序和控制方法的文件（见4.2.4）；
b） 基础设施鉴定；
c） 实施过程参数和产品特性的监视和测量；
d） 获得和使用监视和测量设备；
e） 对标记和包装实施规定的操作；
f） 实施产品放行、交付和交付后活动。
组织应为每一台或每一批医疗器械建立并保留记录（见4.2.5），该记录提供了7.5.9规定的可追溯性范围和程度的信息并标明制造数量和批准销售数量。应核验和批准该记录。
7.5.2 产品的清洁 | 第四十七条在生产过程中需要对原材料、中间品等进行清洁处理的，应当明确清洁方法和要求，并对清洁效果进行验证。
在下列情况下，组织应将产品的清洁或污染控制要求形成文件；
a） 产品在灭菌或使用前有组织进行清洁；
b） 产品是以非无菌形式提供且需在灭菌或使用前进行清洁处理；
c） 产品在灭菌或使用前不能进行清洁处理，使用时其清洁是至关重要的；
d） 提供产品为非无菌使用，使用时其清洁时至关重要的；
e） 在制造过程中从产品中除去过程添加物。
如果产品是按照上述a)或b）的要求进行清洁，则6.4.1中包含的要求不适用于清洁处理前的过程。
7.5.3 安装活动 | 第六十五条需要由企业安装的医疗器械，应当确定安装要求和安装验证的接收标准，建立安装和验收记录。由使用单位或者其他企业进行安装、维修的，应当提供安装要求、标准和维修零部件、资料、密码等，并进行指导。
适当时，组织应将医疗器械安装要求和安装验证接受准则形成文件。
如果经同意的顾客要求允许除组织或其供方以外的外部方安装医疗器械，则组织应提供医疗器械安装和安装验证的形成文件的要求。
应保留由组织或其供方完成的医疗器械安装和安装验证的记录（见4.2.5）。
7.5.4 服务活动 | 第六十四条企业应当具备与所生产产品相适应的售后服务能力，建立健全售后服务制度。应当规定售后服务的要求并建立售后服务记录，并满足可追溯的要求。
如果对医疗器械服务有规定的要求，必要时，组织应将服务程序、所涉及的材料和所涉及的测量形成文件，用于实施服
务活动并验证产品要求得到满足。
组织应分析有组织或其供方实施的服务活动记录：
a） 以确定该信息是否作为投诉进行处置；
b） 适当时，为改进过程形成输入。
应保留由组织或其供方实施的服务活动的记录（见4.2.5）。
7.5.5 无菌医疗器械的专用要求 | 无菌附录
组织应保留每一灭菌批的灭菌过程参数的记录（见4.2.5）。灭菌记录应可追溯到医疗器械的每一生产批。
7.5.6 生产和服务提供过程的确认 | 第四十九条企业应当对生产的特殊过程进行确认，并保存记录，包括确认方案、确认方法、操作人员、结果评价、再确认等内容。生产过程中采用的计算机软件对产品质量有影响的，应当进行验证或者确认。
当生产和服务提供过程的输出不能或不是由后续的监视或测量加以验证，并因此使问题仅在产品使用后或服务交付后才显现时，组织应对任何这样的过程进行确认。
确认应证实这些过程具有稳定地实现所策划的结果的能力。
组织应将过程确认程序形成文件，过程确认包括：
a) 为过程的评审和批准所规定的准则；
b) 设备鉴定和人员资格鉴定；
c) 使用特定的方法、程序和接收准则；
d) 适当时包括包含样本量原理的统计技术；
e) 记录的要求（见4.2.5）；
f) 再确认，包括再确认的准则；
g) 对过程更改的批准。
组织应将用于生产和服务提供的计算机软件应用的确认程序形成文件。此类软件的应用在首次使用前应予确认，适当时，此类软件或其应用更改后也应予确认。有关软件确认和再确认的特定方法和活动应与软件使用有关的风险（包括对产品符合规范的能力的影响）相适应。
应保留确认的结果和结论以及确认所采取的必要措施的记录（见4.2.4和4.2.5）。
7.5.7 灭菌过程和无菌屏障系统确认的专用要求 | 无菌附录
组织应将灭菌过程和无菌屏障系统的确认程序形成文件（见4.2.4）。
灭菌过程和无菌屏障系统应在实施前得到确认，适当时，还应在后续的产品或过程更改实施前得到确认。
应保留确认的结果和结论以及确认所采取的必要措施的记录（见4.2.4和4.2.5）。
注：更多信息见ISO11607-1和ISO11607-2。
7.5.8 标识 | 第五十一条企业应当建立产品标识控制程序，用适宜的方法对产品进行标识，以便识别，防止混用和错用。
组织应将产品标识程序形成文件，并在产品实现的整个过程中使用适当的方法识别产品。
组织应在产品实现的整个过程中按照监视和测量要求识别产品的状态。在产品的生产、贮存、安装和服务的全过程中应保持产品状态的标识，以确保只有通过所要求的检验和试验或经授权让步放行的产品才能被发送、使用或安装。
如果有适用的法规要求，组织应将为医疗器械指定唯一器械标识的系统形成文件。
组织应建立程序并形成文件以确保返回组织的医疗器械能被识别且能与合格的产品区分开。
7.5.9 可追溯性 | 第五十三条企业应当建立产品的可追溯性程序，规定产品追溯范围、程度、标识和必要的记录。
7.5.9.1 总则
组织应将可追溯性程序形成文件。这些程序应依据适用的法规要求规定可追溯性的范围和程序以及拟保留的记录（见4.2.5）。
7.5.9.2 植入性医疗器械的专用要求 | 植入附录
如果所使用的组件、材料和工作环境条件等因素可能导致医疗器械不满足其规定的安全和性能要求，可追溯性所要求的记录应包括这些相关因素的记录。
组织应要求提供流通服务的供方或经销商保留医疗器械的流通记录以便追溯，若检查需要，可获得这些记录。
应保留货运包装收件人的名字和地址的记录（见 4.2.5）。
7.5.10 顾客财产 | /
若顾客财产在组织控制下或由组织使用，对组织使用的或构成产品一部分的顾客财产，组织应予以识别、验证、保护和防护。若任何顾客财产发生丢失、损坏或发现不适用情况，组织应向顾客报告并保留记录（见4.2.5）。
7.5.11 产品保护 | 第五十五条企业应当建立产品防护程序，规定产品及其组成部分的防护要求，包括污染防护、静电防护、粉尘防护、腐蚀防护、运输防护等要求。防护应当包括标识、搬运、包装、贮存和保护等。
在产品的加工、贮存、处置和流通期间，组织应将为产品要求提供防护的程序形成文件。防护应适用于医疗器械的组成部分。
在产品的加工、贮存、处置和流通期间，当其暴露于预期条件和危险（源）时，组织应通过以下法师放置产品发生变化、污染或损坏：
a）设计和制作适当的包装和货运包装箱；
b）如果包装本身不能提供防护，将所需的特殊条件要求形成文件。
如要求特殊条件，其应受控并予以记录（见4.2.5）。
7.6 监视和测量设备的控制 | 第五十七条检验仪器和设备的管理使用应当符合以下要求：
组织应确定需实施的监视和测量以及所需的监视和测量设备，为产品符合确定的要求提供证据。 | （一）定期对检验仪器和设备进行校准或者检定，并予以标识；（二）规定检验仪器和设备在搬运、维护、贮存期间的防护要求，防止检验结果失准；（三）发现检验仪器和设备不符合要求时，应当对以往检验结果进行评价，并保存验证记录；（四）对用于检验的计算机软件，应当确认。
组织应建立程序并形成文件，以确保监视和测量活动可行并以与监视和测量要求相一致的方式实施。
为确保结果有效，必要时，测量设备应：
a ) 对照能溯源到国际或国家标准的测量标准，按照规定的时间间隔或在使用前进行校准和（或）检定，当不存在上述标准时，应记录校准或验证依据（见4.2.5）：
b ) 必要时得到调整或再调整；应记录这种调整或再调整（见4.2.5）；
c ) 具有标识，以确定其校准状态；
d ) 予以防护，防止由于调整使测量结果失效；
e ) 予以保护，防止处置、维护和贮存期间的损坏和衰减。
组织应按照形成文件的程序执行校准或检定。
此外，当发现设备不符合要求时，组织应对以往测量结果的有效性进行评定和记录。组织应对该设备和任何受影响的产品采取适当的措施。
应保留校准和检定（验证）结果的记录（见4.2.5）。
组织应将用于监视和测量要求的计算机软件应用的确认程序形成文件。此类软件的应用在首次使用前应予确认，适当时，此类软件或其应用更改后也应予确认。有关软件确认和再确认的特定方法和活动应与软件使用有关的风险（包括对产品符合规范的能力的影响）相适应。
应保留确认的结果和结论以及确认所采取的必要措施的记录（见4.2.4和4.2.5）。注：更多信息见ISO 10012
8 测量、分析和改进 | /
8.1 总则
组织应策划并实施所需的监视、测量、分析和改进过程以：
a ) 证实产品的符合性；
b ) 确保质量管理体系的符合性；
C ) 保持质量管理体系的有效性。
这应包括对统计技术在内的适当方法及其使用程度的确定。
8.2 监视和测量 | 第六十六条企业应当建立顾客反馈处理程序，对顾客反馈信息进行跟踪分析。
8.2.1 反馈
作为对质量管理体系有效性的一种测量，组织应收集和监视组织是否满足顾客要求的相关信息，并应将获取和利用这种信息的方法形成文件。
组织应将反馈过程程序形成文件。该反馈过程应包括从生产和生产后活动中收集数据的规定。
从反馈过程中收集的信息应用作监视和保护产品要求的风险管理的潜在输入以及产品实现或改进过程的潜在输入。
如果适用的法规要求要求组织从生产后活动获取特定经验，则对该经验的评审应构成反馈过程的一部分。
8.2.2 投诉处置 | 第七十一条企业应当指定相关部门负责接收、调查、评价和处理顾客投诉，并保持相关记录。
组织应按照适用的法规要求将及时处置投诉的程序形成文件。
这些程序应包括对以下方面的最低要求和职责：
a)接收和记录信息；
b)评价信息以确定反馈是否构成投诉；
c)调查投诉；
d)确定是否需要向适当的监管机构报告信息；
e)处置与投诉有关的产品；
f)确定是否需要启动纠正或纠正措施。
如果有任何投诉未经调查，应记录理由。应记录由投诉处置过程形成的任何纠正或纠正措施。
如果一项调查确定是组织外的活动导致了投诉，则应在组织和所涉及的外部方之间交换相关信息。应保留投诉处置记录（见4.2.5）。
8.2.3 向监管机构报告 | 第七十二条企业应当按照有关法规的要求建立医疗器械不良事件监测制度，开展不良事件监测和再评价工作，并保持相关记录。
如果适用的法规有要求，将符合不良事件报告准则的投诉或发布的忠告性通知报告，则组织应将向有关的监管机构报告的程序形成文件。 | 第七十六条企业应当建立产品信息告知程序，及时将产品变动、使用等补充信息通知使用单位、相关企业或者消费者。
应保留向监管机构报告的记录（见4.2.5）。
8.2.4 内部审核 | 第七十七条企业应当建立质量管理体系内部审核程序，规定审核的准则、范围、频次、参加人员、方法、记录要求、纠正预防措施有效性的评定等内
组织应按策划的时间间隔进行内部审核以确定质量管理体系是否： | 容，以确保质量管理体系符合本规范的要求。
a）符合策划并形成文件的安排、本标准的要求以及组织所
确定的质量管理体系要求和适用的法规要求；
b）得到有效实施与保持。
组织应建立程序并形成文件以说明策划和实施审核以及记
录和报告审核结果的职责和要求。
组织应策划审核方案，策划时应考虑拟审核的过程和区域的
状况和重要性以及以往的审核结果。应规定并记录审核的准
则、范围、时间间隔和方法（见 4.2.5）。审核员的选择和审
核的实施应确保审核过程客观公正。审核员不应审核自己的
工作。
应保留审核和审核结果的记录（见 4.2.5），包括过程、受
审核区域和结论。
负责受审核区域的管理者应确保采取任何必要的纠正和
纠正措施，应无不当拖延，以消除所发现的不合格及其原因。
后续活动应包括验证所采取的措施并报告验证结果。
注：更多信息见 ISO 19011。
8.2.5 过程的见识和测量 | /
组织应采用适宜的方法对质量管理体系过程进行监视，并在适当时进行测量。这些方法应证实过程实施策划结果的能力。若未能实施策划结果，适当时，应采取纠正和纠正措施。
8.2.6 产品的监视和测量 | 第五十六条 企业应当建立质量控制程序，规定产品检验部门、人员、操作等要求，并规定检验仪器和设备的使用、校准等要求，以及产品放行的程序。
组织应对产品的特性进行监视和测量，以验证产品要求已得到满足。这种件事和测量应依据策划并形成文件的安排和形成文件的程序，在产品实现过程的适当阶段进行。 | 第五十八条 企业应当根据强制性标准以及经注册或者备案的产品技术要求制定产品的检验规程，并出具相应的检验报告或者证书。需要常规控制的进货检验、过程检验和成品检验项目原则上不得进行委托检验。对于检验条件和设备要求较高，确需委托检验的项目，可委托具有资质的机构进行检验，以证明产品符合强制性标准和经注册或者备案的产品技术要求。
应保持符合接收准则的证据。应记录有权放行产品的人员的身份（见4.2.5）。适当时，记录应识别用于执行测量活动的检测设备。 | 第五十九条 每批（ 台）产品均应当有检验记录，并满足可追溯的要求。检验记录应当包括进货检验、过程检验和成品检验的检验记录、检验报告或者证书等。
在策划并形成文件的安排已圆满完成前不应放行产品和交付服务。 | 第六十条 企业应当规定产品放行程序、条件和放行批准要求。放行的产品应当附有合格证明。
对于植入性医疗器械，组织应记录进行任何检验或试验的人员的身份。 | 第六十一条 企业应当根据产品和工艺特点制定留样管理规定，按规定进行留样，并保持留样观察记录。
8.3 不合格品控制 | 第六十七条 企业应当建立不合格品控制程序，规定不合格品控制的部门和人员的职责与权限。
8.3.1 总则 | 第六十八条 企业应当对不合格品进行标识、记录、隔离、评审，根据评审结果，对不合格品采取相应的处置措施。
组织应确保对不符合产品要求的产品进行识别和控制，以防止非预期的使用或交付。组织应建立程序并形成文件以规定不合格品控制以及与不合格品识别、记录、隔离、评价和处置有关的职责和权限。不合格的评价因包括确定是否需要调查和通知对不合格负责的所有外部方。 | 第七十五条 对于存在安全隐患的医疗器械，企业应当按照有关法规要求采取召回等措施，并按规定向有关部门报告。
应保留不合格的性质以及随后所采取的任何措施的记录，包括评价、任何调查和决策的理由说明（4.2.5）。
8.3.2 交付前发现不合格品的响应措施 | 第七十条不合格品可以返工的，企业应当编制返工控制文件。返工控制文件包括作业指导书、重新检验和重新验证等内容。不能返工的，应当建立相关处置制度。
组织应通过下列一种或几种途径处置不合格品：
a）采取措施以消除已发现的不合格；
b）采取措施以防止其原预期的使用或应用；
c）授权让步使用、放行或接收。
组织应确保不合格品仅在提供理由、或得到批准和满足适用的法规要求的情况下才能让步接收。应保留让步接收和授权让步人员身份的记录（4.2.5.）
8.3.3 交付后发现不合格品的响应措施 | 第六十九条在产品销售后发现产品不合格时，企业应当及时采取相应措施，如召回、销毁等。
当交付后或开始使用后发现不合格品时，组织应采取与不合格的影响或潜在影响的程度相适应的措施。应保留所采取措施的记录（4.2.5）。
组织应按照适用的法规要求将忠告性通知的发布程序形成文件。这些程序应能随时付诸实施。应保留与发布忠告性通知相关的措施的记录（见4.2.5）。
8.4 数据分析 | 第七十三条企业应当建立数据分析程序，收集分析与产品质量、不良事件、顾客反馈和质量管理体系运行有关的数据，验证产品安全性和有效性，并保持相关记录。
组织应将确定、收集和分析适当数据的程序形成文件以证实质量管理体系的适宜性、充分性和有效性。这些程序应包括对统计技术在内的适当方法及其使用程度的确定。
数据分析应包括来自监视和测量的结果以及其他有关来源的数据，并至少包括以下方面的输入：
a)反馈；
b)产品要求的符合性；
c)过程和产品的特性及趋势，包括改进的机会；
d)供方；
e)审核；
f)适当时，服务报告。
如果数据分析表明质量管理体系不是适宜的、充分的或有效的，组织应按照8.5的要求将此分析结果用作改进的输入。
应保留分析结果的记录（见4.2.5）。
8.5 改进 | /
8.5.1 总则
组织应利用质量方针、质量目标、审核结果、上市后监督、数据分析、纠正措施、预防措施和管理评审来识别和实验任何必要的更改，以确保和保持质量管理体系的持续适宜性、充分性和有效性以及医疗器械的安全和性能。
8.5.2 纠正措施 | 第七十四条企业应当建立纠正措施程序，确定产生问题的原因，采取有效措施，防止相关问题再次发生。
组织应采取措施消除不合格的原因以防止不合格的再发生。组织应采取任何必要的纠正措施，应无不当拖延。纠正措施应与不合格的影响程度相适应。
组织应将规定以下方面要求的程序形成文件：
a)评审不合格（包括投诉）;
b)确定不合格的原因；
c)评价确保不合格不再发生的措施的需求；
d)对所需的措施进行策划、形成文件并实施，适当时，包括更新文件；
e)验证纠正措施对满足适用的法规要求的能力和对医疗器械的安全和性能无不良影响；
f)评审所采取的纠正措施的有效性。
应保留任何调查的结果和所采取措施的记录（见4.2.5）。
8.5.3 预防措施 | 第七十四条应当建立预防措施程序，确定潜在问题的原因，采取有效措施，防止问题发生。
组织应确定措施消除潜在不合格的原因以防止不合格的发生。预防措施应与潜在问题的影响程度相适应。
组织应将说明以下发面要求的程序形成文件：
a)确定潜在不合格及其原因；
b)评价防止不合格发生的措施的需求；
c)对所需的措施进行策划、形成文件并实施，适当时，包括更新文件；
d)验证预防措施对满足适用的法规要求的能力和对医疗器械的安全和性能无不良影响；
e)适当时，评审所采取的预防措施的有效性。
应保留任何调查的结果和所采取措施的记录（见4.2.5)。