## Slide 1: STED and CSDT for IVDs – Technical Documentation for the Regulatory Review under GHTF and AHWP
Dr. Petra Kaars-Wiele
Director International Regulatory Affairs/ Affiliate Compliance, Abbott, Germany
15th AHWP Meeting, November 28, 2010, Riyadh/ Kingdom of Saudi Arabia

## Slide 2: STED- What is it?
STED = Summary Technical Documentation
  The STED does not represent the full technical documentation which is controlled under the Quality System
  STED should be in a language acceptable to the reviewing organization
  Depth and detail may be dependant on classification and risk, whether it is novel technology or already marketed
  The STED represents a subset of documents which provides sufficient evidence for a regulatory body to confirm conformity to the essential principles
  The STED reflects the status of the device at the time of submission
  STED has been developed for IVD Medical Devices and all other Medical Devices (separate documents)

## Slide 3: STED = Summary Technical Documentation
  Technical documentation is to be prepared for all IVDs according SG1(PD)/N063
  STED submission and review by the authorities not required for Class A and B, but if in justified cases may be submitted on request.
  STED review by the authorities for Class C and D, based on risk of the products and experience of the manufacturer

## Slide 4: Why was it developed?
  Today, pre-market submission content and formats vary widely from country to country and from manufacturer to manufacturer
  This approach can be time consuming for both the manufacturer and the regulator
  The STED document represents the best practice for a regulator to assess the conformity to the Essential Principles

## Slide 5: Content of STED
  General description and list of specified features
    Description of the product and intended use
    Reference to previous generations
  Set of labels, instruction of use in a language acceptabe to the Regulatory Authority or Conformity Assessment Body
  Summary of technical documentation concerning design and manufacturing
    Sufficient to allow the reviewer to obtain a general understanding of the manufacturing processes
    Identify the sites, where the activities are performed

## Slide 6: Content of STED – cont‘d
  Essential principles of safety and performance check list
    Used to better understand how the manufacturer demonstrates conformity to essential principles for his device
    May reference specific standards, test reports, study reports, validation documents
    Sample table contained within the STED guidance
  Summary of risk analysis and mode of control
    Based on recognized standards (ISO 14971)

## Slide 7: Content of STED – cont‘d
  Summary of verification and validation studies – Performance Evaluation
    Level of detail will vary depending upon Class and complexity of the device
    Analytical Studies:
      Specimen types
      Accuracy
      Traceability of calibrators and control material
      Analytical sensitivity
      Analytical specificity
      Measuring range of the assay
      Validation of assay cut-off
    Stability (excluding specimen stability)
      Claimed shelf life
      In use stability
      Shipping stability

## Slide 8: Content of STED – cont‘d
Software verification and validation
Clinical evidence:
    Clinical Utility
      For established assays literature to support the use of the assay is sufficient
      For new markers (first in the market) and no literature available, a study may be required
    Clinical Performance (or: performance evaluation)

## Slide 9: Declaration of Conformity
  A declaration made by the manufacturer
  Name and address of the responsible manufacturer
  Identification of the device (product name, product number, GMDN code)
  Statement that the device complies with
    The Essential Principles of Safety and Performance
    The Classification and Conformity Assessment Procedure
  List standards used in the Conformity Assessment Procedure
  Dated and signed by the manufacturer (appropriate senior personnel)
  May be not part of the STED, but can be annexed to it

## Slide 10: STED – When is it used?
Premarket:
  STED for Class C and D devices to be prepared and submitted to Regulatory Authority (RA) or Conformity Assessment Body (CAB)
  Manufacturer should always retain a copy for any STED submitted for future reference

## Slide 11: STED – When is it used?
Post-Market:
  Submitted upon request of a Regulatory Authority (RA) or Conformity Assessment Body (CAB)
    To demonstrate continued conformity of a Class C or D device
    To investigate conformity of a Class A or B device
BUT
  The STED is not usually used to assist in any post market investigation of medical events

## Slide 12: CSDT = Common Submission Dossier Template
  AHWP WG1 developed the draft document in their last version from 2006
  Draft document can be found on www.ahwp.info
  No final document published yet
  Currently discussed within the AHWP Technical Committee

## Slide 13: CSDT and STED: Two of a Kind? Yes and No
  Substantially both documents are the same
  CSDT describes one document for all Medical Devices (some language is not related to IVD Medical Devices)
  CSDT request an executive summary, while STED request a description of the device
  STED has a risk based conformity assessment frame work (established guidance on respective requirements on e.g. QMS, safety and performance evaluation, PMS for each class of device)
  Difference in requirement on the detail of manufacturing processes (interpretation or fact?)
  the concept of "pre- and post-market use of STED" does not have its counterpart in CSDT
  CSDT does not give focus on performance evaluation for IVDs
  CSDT does not elaborate on the use of standards

## Slide 14: One Summary Technical Documentation – Potential Benefits
Goal of STED (Regulator perspective)
  To provide improved access to new technologies
  To reduce the burden on the regulators by promoting a review of documents presented in a consistent manner
  To provide direction to those markets developing device regulations for the first time

## Slide 15: One Summary Technical Documentation – Potential Benefits
Goal of STED (Manufacturer perspective)
  To reduce the burden on manufacturers of conflicting pre-market submission formats and unique content
  To permit the creation of one core dossier for all submissions
  To minimize regulatory barriers
  To facilitate trade

## Slide 16: One Summary Technical Documentation – Win-Win for all Stakeholders
The implementation of one set of Summary Technical Documentation for IVD Medical Devices is a win-win opportunity for regulators and manufacturers around the world

## Slide 17: Pictures are from Microsoft : www.office.microsoft.com, except cover page (copyright is with Abbott)
Thank you for your attention!
Questions?